## HazardAnalysis
# | Harm | Hazard | Potential Cause | Target | Initial Risk Evaluation |  |  | Risk Controls | Residual Risk Evaluation
[5 empty] | Severity | Probability | Risk Level |  | Severity | Probability | Risk Level
1 | mild hypothermia (32-35°C core body temperatures) | IV fluid or blood product administered at temperatures below 35°C | operator sets water bath temperature set point too low | patient | S2 | P3 | R2 | guidence on water bath set point included in user manual water bath temperature always displayed on heater and temperature monitor | S2 | P1 | R1
2 |  |  | water bath drops below set point during operation |  |  | P2 | R2 | active water bath temperature control active water bath temperature monitoring with under temperature alarm |  | P1 | R1
3 |  |  | operator immerses too short of IV tube length in water bath |  |  | P3 | R2 | guidance on IV tube immersion length included in user manual a IV tube length measurement guide is included at device a switch alerts the operator if the IV tubing is not at the water bath |  | P1 | R1
4 |  |  | operator uses a IV fluid flow rate that is too high |  |  | P3 | R2 | operating instructions state that the operator must be present and actively monitor IV fluid flow rate for the entire duration of the procedure guidance on the maximum IV fluid flow rate that the system is capable of at various conditions is included in the user manual operating instructions state that operator must be professionally trained on how to control and measure IV fluid flow rates before they should be authorized to operate this device |  | P1 | R1
5 |  |  | operator uses a IV bag fluid temperature that is too low |  |  | P3 | R2 | operating instructions state that the operator must be present and actively monitor IV bag fluid temperature for the entire duration of the procedure guidance on the minimum allowable IV bag fluid temperature that the system is capable of at various conditions is included in the user manual operating instructions state that operator must be professionally trained on how to measure and/or determine the IV bag fluid temperature at any point during the procedure to be authorized to operate the device |  | P1 | R1
6 | cardiac hyperkalemia | blood product administered at temperatures above 46°C (occurance of hemolysis) | operator sets water bath temperature set point too high | patient | S4 | P3 | R3 | guidence on water bath set point included in user manual water bath temperature always displayed on heater and temperature monitor | S4 | P1 | R2
7 |  |  | water bath warms above set point during operation |  |  | P2 | R2 | active water bath temperature control active water bath temperature monitoring with over temperature alarm and heater cut off |  | P1 | R2
8 |  |  | operator immerses too long of IV tube length in water bath |  |  | P3 | R3 | guidance on IV tube immersion length is included in user manual (guidance includes safety margin to ensure significant additional length of tubing required before hazard occurs) a IV tube length measurement guide is included at device |  | P1 | R2
9 |  |  | operator uses a IV fluid flow rate that is too low |  |  | P3 | R3 | operating instructions state that the operator must be present and actively monitor IV fluid flow rate for the entire duration of the procedure guidance on the minimum allowable flow rate that the system is capable of at various conditions is included in the user manual (guidance includes safety margin to ensure significant decrease in flow rate before hazard occurs) operating instructions state that operator must be professionally trained on how to control and measure IV fluid flow rates before they should be authorized to operate the device |  | P1 | R2
10 |  |  | operator uses a IV bag fluid temperature that is too high |  |  | P3 | R3 | operating instructions state that the operator must be present and actively monitor IV bag fluid temperature for the entire duration of the procedure guidance on the maximum allowable IV bag fluid temperature that the system is capable of at various conditions is included in the user manual (guidance includes safety margin to ensure significant increase in IV bag fluid temperature before hazard occurs) operating instructions state that operator must be professionally trained on how to measure and/or determine the IV bag fluid temperature at any point during the procedure before they should be authorized to operate the device |  | P1 | R2
11 | third degree skin burns | direct skin exposure to water at temperatures above 68°C for durations longer than 1 second | hot water bath is knocked off mounting table and hot water is spilled | patient, operator | S3 | P3 | R3 | nonslip base is used between the bottom of the water bath and the mounting table guidance on proper device orientation and mounting is included in the user manual water bath is covered with a lid during operation | S3 | P1 | R2
12 |  |  | operator reaches into the hot water bath to retrieve the IV tubing | operator |  | P3 | R3 | guidance on how to remove the IV tubing from the hot water bath without reaching into the water is included in the user manual |  | P1 | R2
13 |  | skin exposure to sparks or open flame | dielectric breakdown at electrical conductor | patient, operator |  | P2 | R2 | operating instructions state that the operator must inspect all signal wires, electrical power wires, and water seals for damage or wear before each use external electrically conductive components are sufficiently grounded or isolated |  | P1 | R2
14 | cardiac arrhythmia and arrest | exposure to hazardous voltages | dielectric breakdown at electrical conductor | patient, operator | S5 | P2 | R3 | operating instructions state that the operator must inspect all signal wires, electrical power wires, and water seals for damage or wear before each use external electrically conductive components are sufficiently grounded or isolated | S5 | P1 | R2
16 |  | exposure to leakage current | dielectric breakdown at electrical conductor |  |  | P2 | R3 | operating instructions state that the operator must inspect all signal wires, electrical power wires, and water seals for damage or wear before each use external electrically conductive components are sufficiently grounded or isolated |  | P1 | R2

## Severity
Rating |  | Severity
S5 | catastrophic | results in death
S4 | critical | results in permanent impairment or life-threatening injury
S3 | serios | results in injury or impairment reuquiring professional medical intervention
S2 | minor | results in temporary injury or impairment not requiring professional medical intervention
S1 | negligible | inconvenience or temporary discomfort
*reference EN ISO 14971:2012 Table D.3

## Occurance
Rating |  | Probability of Occurance
P5 | Frequent | >= 10^-3 | Occurs "every time"
P4 | Probable | < 10^-3 and >= 10^-4 | Good chance to occur, considerable certainty to occur, or "reasonable probablility that use will cause"
P3 | Occasional | < 10^-4 and >= 10^-5 | Expected to occur from time to time (no clear trend)
P2 | Remote | <10^-5 and >= 10^-6 | Rarely occuring or "remote probability that use will cause"
P1 | Improbable | <10^-6 | Not expecting to occur, "not likely that use will cause any adverse events"
*reference EN ISO 14971:2012 Table D.4

## RiskEvalMatrix
Severity
|  |  | S1 | S2 | S3 | S4 | S5
|  |  | negligible | minor | serios | critical | catastrophic
Probability of Occurance | P5 | Frequent | R3 | R3 | R3 | R3 | R3
| P4 | Probable | R3 | R3 | R3 | R3 | R3
| P3 | Occasional | R2 | R3 | R3 | R3 | R3
| P2 | Remote | R1 | R2 | R2 | R3 | R3
| P1 | Improbable | R1 | R1 | R2 | R2 | R2
Rating | Risk Evaluation
R3 | Unacceptable risk
R2 | Investigate further risk reduction
R1 | Acceptable risk; insignificant risk
*reference EN ISO 14971:2012 Figure D.7