## OperatingProcedure
Step | Description | Feedback
1 | Review the operating parameters table of the user manual to determine water bath set point temperature and length of IV tubing to immerse in the water bath required for the diameter of IV tubing, IV fluid flow rate, and temperature of the IV fluid in the IV bag | none
2 | Plug the unit into an AC power source | safety companion receives power, sous vide power push button red light ring flashes, set point monitor push button blue light ring flashes
3 | Press the sous vide power push button on the safety companion | sous vide cooker recieves power, sous vide power push button red light ring goes solid
4 | Set the water bath set point temperature on the sous vice cooker | sous vide cooker display shows temperature set point
5 | Press the start button on the sous vide cooker | sous vide cooker display shows that the device is operating, sous vide cooker starts warming the water bath, water circulation can be seen and/or heard in the water bath from the sous vide cooker water agitator
6 | Wait while the sous vide cooker brings the water bath temperature up to the set point | sous vide cooker will indicate the set point temperature and measured water bath temperature are the same
7 | Push the safety companion set point monitor push button once the sous vide cooker has stabilized the water bath temperature at the set point temperature | set point monitor push button blue light ring goes solid
8 | Remove the water bath lid | visible evidence that the lid is no longer on the water bath
9 | Immerse the required length of IV tubing in the water bath | visible evidence that the IV tubing is fully immersed in the water bath
10 | Place the patient side of the IV tubing through the iv tube switch | visible evidence that the IV tubing runs through the iv tube switch
11 | Reattach the water bath lid | visible evidence that the lid has been reattached to the water bath
12 | Flow IV fluid through IV tubing at the flow rate | no feedback from safety companion, the IV fluid is visible in the IV tube
13 | Remain present during the procedure to monitor the device operating parameters; take immedicate corrective action if the device operating parameters deviate from the values determined in step 1 or a fault condition is indicated by the OpenFluidWarmer device | the following operating parameters must be continually monitored by the operator and are not monitored by the OpenFluidWarmer device: the temperature of the IV fluid in the IV bag, the IV fluid flow rate, and the length of IV tubing immersed in the water bath; audible and visual indicators will alert the user if operating parameters that are monitored by the OpenFluidWarmer device deviate outside normal operating limits