## Product_Requirements
Product Requirements
Number | Type | Requirement | Justification
1 | shall | never fail to alert the operator of a hazardous condition | common safety feature of life critical medical devices; audio and visual indicators to communicate faults
2 | shall | cost less than $100 USD | reasonable target based on the initial OpenFluidWarmer cost workup
3 | should | cost less than $80 USD | reasonable target based on the initial OpenFluidWarmer cost workup
4 | shall | warm fluid to 38 +/- 2°C at a flow rate of at least 20mL/min | 38°C is assumed human body temperature, 20mL/min is lowest of "fast" flow rate defined in Field Ready's 2018 fluid warmer design study
5 | should | warm fluid to 38 +/- 2°C at a flow rate of at least 80mL/min | 38°C is assumed human body temperature, 80mL/min is typical of a fast flow rate infusion or transfusion
6 | shall | turn off the heaters and alert the user when fluid is at risk of exceeding 46°C | 42°C is when risk of hemolysis begins
7 | shall | alert the user when fluid temperature is at risk of dropping below 36°C | low temperature alarm to ensure that fluid is being warmed to near human body temperature of 38°C
8 | shall | warm up upon startup in less than 2 minutes from 4 to 36°C and in less than 45 seconds from 20 to 36°C | warm up period typical of commerical IV fluid warmers
9 | shall | be easy to clean and disinfect surfaces | common requirement of commerical IV fluid warmers, generally not critical that device surfaces be sterile
10 | shall | operate continuously when powered | common safety feature of life critical medical devices
11 | shall | meet or exceed IP54 ingress protection per IEC 60529 (partially protect against dust that may harm equipment, water splashing against the enclosure from any direction shall have no harmful effect) | IEC 60601-1
12 | shall | meet IEC electrical shock protection Class I (chassis connected to electrical earth) | common requirement of commerical IV fluid warmers
13 | shall | include protections against defibrillation voltages | IEC 60601-1, 8.5.5 defibrillation-proof applied parts
14 | shall | be constructed of common off the shelf components | so components can be easily sourced and when things break custom hardware is not required
15 | shall | be constructed using common tools and manufacturing methods | to reduce the cost of fabrication and increase likelihood that others already have the capabilites to construct the device
16 | shall | assume that the assembler has little knowledge of standard electronics assembly or troubleshooting methods | develop the assembly and user manual so users with very little knowledge of electronics can build the device
17 | should | include a method for tracking the total duration of time that the device has been in operation | to help user follow a maintainance schedule based on the number of hours the device has been in operation
18 | should | include a fault history log | to make troubleshooting fault conditions easier
19 | shall | display set point temperature on the device | ensure that users are aware of what temperature the device will warm fluid to
20 | shall | display applicable safety hazard warnings on the device | IEC 60601-1, Section 7 ME equipment identification, marking, and documents
21 | shall | properly isolate or ground all external conductive components | IEC 60601-1, 8.6 protective earthing, functional earthing, an potential equalizations of ME euipment and 8.8 insulation
22 | shall | maintain proper creepage and clearance distances | IEC 60601-1, 8.9 creepage distances and air clearances
23 | shall | meet or exceed ASTM F2172 Standard Specification for Blood/Intravenous Fluid/Irrigation Fluid Warmers, requirement for fluid warmers to minimize the hazards to the patient and operator | medical device industry standard
24 | shall | meet or exceed IEC 60601, general requirements for basic safety and essential perofrmance of medical electrical equipment | medical device industry standard
25 | shall | meet or exceed 21 CFR 820, quality requirements for the design, manufacture, packaging, labeling, storage, installation, and servicing of all finished devices intended for human use | medical device industry standard | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfcfr/CFRSearch.cfm?CFRPart=820
26 | shall | meet or exceed ISO 13485, quality management system requirements for medical devices | medical device industry standard

## Design_Targets
Design Targets
reduce operator and patient risks as-far-as-possible
minimize the total cost of the device
minimize the number of possible ways the device could be used incorrectly and cause harm
maximize heat transfer efficiency from the heating elements to the IV fluid
minimize heat loss to the environment in which the device operates
minimize the length of IV tube used by the device

## Product_Classification
Product Classification
MDD classification, Class IIb according to Rule 9