## data
sweep | point | sample | source_vln | source_i | source_pf | source_freq | Reference vln | Reference i | Reference pf | Reference kW | Reference kVAR | Reference kVA | P110M vln | P110M i | P110M kW | err_vln | err_i | err_kW | Voltage Error | Current Error | Real Power Error
current_sweep | 1 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0043 | 0.010005 | 99.8988 | 0.0011994 | -5.3539e-05 | 0.0012006 | 120.229 | 0.117 | 0.001434 | 0.22465 | 0.107 | 0.000235 | 0.1872 | 1069.4494 | 19.5604
current_sweep | 1 | 2 | 120 | 0.01 | 0.999 | 60 | 120.0023 | 0.010005 | 99.8984 | 0.0011993 | -5.3674e-05 | 0.0012006 | 120.236 | 0.109 | 0.001144 | 0.23367 | 0.098995 | -5.5e-05 | 0.1947 | 989.5081 | -4.6146 | Averaged Data | [21 empty] | Graphed Data
current_sweep | 1 | 3 | 120 | 0.01 | 0.999 | 60 | 120.0025 | 0.010006 | 99.8967 | 0.0011995 | -5.4112e-05 | 0.0012007 | 120.227 | 0.104 | 0.001356 | 0.22447 | 0.093994 | 0.000157 | 0.1871 | 939.3899 | 13.0481 | sweep | point | sample | source_vln | source_i | source_pf | source_freq | Reference vln | Reference i | Reference pf | Reference kW | Reference kVAR | Reference kVA | P110M vln | P110M i | P110M kW | err_vln | err_i | err_kW | Voltage Error | Current Error | Real Power Error | sweep | point | sample | source_vln | source_i | source_pf | source_freq | Reference vln | Reference i | Reference pf | Reference kW | Reference kVAR | Reference kVA | P110M vln | P110M i | P110M kW | err_vln | err_i | err_kW | Voltage Error | Current Error | Real Power Error
current_sweep | 2 | 1 | 120 | 0.0169 | 0.999 | 60 | 120.0019 | 0.016901 | 99.8992 | 0.0020261 | -9.0735e-05 | 0.0020282 | 120.238 | 0.111 | 0.001373 | 0.23608 | 0.094099 | -0.000653 | 0.1967 | 556.7596 | -32.2353 | current_sweep | 1 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0031 | 0.010005 | 99.8979 | 0.0011994 | -5.3775e-05 | 0.0012006 | 120.2307 | 0.11 | 0.0013113 | 0.2276 | 0.099995 | 0.00011233 | 0.18967 | 999.4491 | 9.3313 | current_sweep | 1 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0031 | 0.010005 | 99.8979 | 0.0011994 | -5.3775e-05 | 0.0012006 | 120.2307 | 0.11 | 0.0013113 | 0.2276 | 0.099995 | 0.00011233 | 0.18967 | 999.4491 | 9.3313
current_sweep | 2 | 2 | 120 | 0.0169 | 0.999 | 60 | 120.0051 | 0.016905 | 99.8993 | 0.0020266 | -9.0758e-05 | 0.0020287 | 120.239 | 0.105 | 0.001905 | 0.23391 | 0.088095 | -0.000122 | 0.1949 | 521.1199 | -6.0019 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 2 | 1 | 120 | 0.0169 | 0.999 | 60 | 120.003 | 0.016903 | 99.8994 | 0.0020263 | -9.0675e-05 | 0.0020284 | 120.237 | 0.11233 | 0.0018853 | 0.23399 | 0.095431 | -0.000141 | 0.19497 | 564.5879 | -6.9594
current_sweep | 2 | 3 | 120 | 0.0169 | 0.999 | 60 | 120.002 | 0.016902 | 99.8997 | 0.0020263 | -9.0532e-05 | 0.0020283 | 120.234 | 0.121 | 0.002378 | 0.23197 | 0.1041 | 0.000352 | 0.1933 | 615.8843 | 17.359 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 3 | 1 | 120 | 0.0284 | 0.999 | 60 | 120.0021 | 0.028403 | 99.8995 | 0.003405 | -0.0001526 | 0.0034084 | 120.2327 | 0.125 | 0.00356 | 0.23056 | 0.096597 | 0.000155 | 0.19213 | 340.092 | 4.5517
current_sweep | 3 | 1 | 120 | 0.0284 | 0.999 | 60 | 120.0032 | 0.028403 | 99.8995 | 0.003405 | -0.0001526 | 0.0034084 | 120.231 | 0.126 | 0.003418 | 0.22777 | 0.097597 | 1.3e-05 | 0.1898 | 343.6168 | 0.3814 | current_sweep | 2 | 1 | 120 | 0.0169 | 0.999 | 60 | 120.003 | 0.016903 | 99.8994 | 0.0020263 | -9.0675e-05 | 0.0020284 | 120.237 | 0.11233 | 0.0018853 | 0.23399 | 0.095431 | -0.000141 | 0.19497 | 564.5879 | -6.9594 | current_sweep | 4 | 1 | 120 | 0.0479 | 0.999 | 60 | 120.0046 | 0.047899 | 99.9 | 0.0057424 | -0.00025693 | 0.0057481 | 120.2337 | 0.12933 | 0.0057377 | 0.22908 | 0.081434 | -4.6667e-06 | 0.1909 | 170.0108 | -0.082167
current_sweep | 3 | 2 | 120 | 0.0284 | 0.999 | 60 | 120.0019 | 0.028404 | 99.8994 | 0.0034051 | -0.00015272 | 0.0034085 | 120.236 | 0.126 | 0.004034 | 0.23408 | 0.097596 | 0.000629 | 0.1951 | 343.6064 | 18.4711 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 5 | 1 | 120 | 0.0808 | 0.999 | 60 | 120.0043 | 0.080796 | 99.8998 | 0.0096862 | -0.00043381 | 0.0096959 | 120.2403 | 0.14467 | 0.0098777 | 0.23601 | 0.06387 | 0.00019167 | 0.19667 | 79.051 | 1.9767
current_sweep | 3 | 3 | 120 | 0.0284 | 0.999 | 60 | 120.0012 | 0.028403 | 99.8997 | 0.003405 | -0.00015249 | 0.0034084 | 120.231 | 0.123 | 0.003228 | 0.22984 | 0.094597 | -0.000177 | 0.1915 | 333.0527 | -5.1975 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 6 | 1 | 120 | 0.1362 | 0.999 | 60 | 120.0038 | 0.1362 | 99.9006 | 0.016328 | -0.0007273 | 0.016344 | 120.2307 | 0.18067 | 0.016497 | 0.2269 | 0.044469 | 0.00016833 | 0.1891 | 32.6499 | 1.0328
current_sweep | 4 | 1 | 120 | 0.0479 | 0.999 | 60 | 120.0049 | 0.047901 | 99.8999 | 0.0057426 | -0.00025698 | 0.0057484 | 120.233 | 0.129 | 0.005589 | 0.22812 | 0.081099 | -0.000154 | 0.1901 | 169.3056 | -2.6747 | current_sweep | 3 | 1 | 120 | 0.0284 | 0.999 | 60 | 120.0021 | 0.028403 | 99.8995 | 0.003405 | -0.0001526 | 0.0034084 | 120.2327 | 0.125 | 0.00356 | 0.23056 | 0.096597 | 0.000155 | 0.19213 | 340.092 | 4.5517 | current_sweep | 7 | 1 | 120 | 0.2297 | 0.999 | 60 | 120.004 | 0.22968 | 99.9006 | 0.027535 | -0.0012281 | 0.027562 | 120.223 | 0.25333 | 0.027338 | 0.21902 | 0.023654 | -0.00019667 | 0.1825 | 10.2987 | -0.7143
current_sweep | 4 | 2 | 120 | 0.0479 | 0.999 | 60 | 120.006 | 0.047898 | 99.9001 | 0.0057423 | -0.00025673 | 0.0057481 | 120.237 | 0.128 | 0.005813 | 0.23103 | 0.080102 | 7.1e-05 | 0.1925 | 167.2327 | 1.2304 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 8 | 1 | 120 | 0.3873 | 0.999 | 60 | 120.0035 | 0.38727 | 99.9003 | 0.046428 | -0.0020747 | 0.046474 | 120.238 | 0.39967 | 0.046845 | 0.23453 | 0.012393 | 0.000417 | 0.19547 | 3.2001 | 0.89783
current_sweep | 4 | 3 | 120 | 0.0479 | 0.999 | 60 | 120.0029 | 0.047899 | 99.8998 | 0.0057422 | -0.00025709 | 0.005748 | 120.231 | 0.131 | 0.005811 | 0.22809 | 0.083101 | 6.9e-05 | 0.1901 | 173.4942 | 1.1978 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 9 | 1 | 120 | 0.653 | 0.999 | 60 | 120.0034 | 0.65297 | 99.9001 | 0.078281 | -0.0035006 | 0.078359 | 120.223 | 0.65233 | 0.078649 | 0.21958 | -0.000641 | 0.00036833 | 0.18297 | -0.098167 | 0.47067
current_sweep | 5 | 1 | 120 | 0.0808 | 0.999 | 60 | 120.0045 | 0.080799 | 99.8998 | 0.0096865 | -0.00043386 | 0.0096962 | 120.239 | 0.146 | 0.010454 | 0.23451 | 0.065201 | 0.000768 | 0.1954 | 80.6962 | 7.9236 | current_sweep | 4 | 1 | 120 | 0.0479 | 0.999 | 60 | 120.0046 | 0.047899 | 99.9 | 0.0057424 | -0.00025693 | 0.0057481 | 120.2337 | 0.12933 | 0.0057377 | 0.22908 | 0.081434 | -4.6667e-06 | 0.1909 | 170.0108 | -0.082167 | current_sweep | 10 | 1 | 120 | 1.101 | 0.999 | 60 | 120.0057 | 1.1009 | 99.9004 | 0.13198 | -0.0058928 | 0.13211 | 120.224 | 1.0907 | 0.13191 | 0.21831 | -0.010234 | -7.6e-05 | 0.1819 | -0.9296 | -0.057467
current_sweep | 5 | 2 | 120 | 0.0808 | 0.999 | 60 | 120.0036 | 0.080795 | 99.8997 | 0.009686 | -0.00043401 | 0.0096957 | 120.238 | 0.142 | 0.009331 | 0.23438 | 0.061205 | -0.000355 | 0.1953 | 75.7533 | -3.6648 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 11 | 1 | 120 | 1.8563 | 0.999 | 60 | 120.0036 | 1.856 | 99.9003 | 0.2225 | -0.0099384 | 0.22273 | 120.22 | 1.8293 | 0.22248 | 0.21637 | -0.026655 | -2.7e-05 | 0.18033 | -1.4361 | -0.012133
current_sweep | 5 | 3 | 120 | 0.0808 | 0.999 | 60 | 120.0049 | 0.080795 | 99.8999 | 0.0096861 | -0.00043357 | 0.0096958 | 120.244 | 0.146 | 0.009848 | 0.23913 | 0.065205 | 0.000162 | 0.1993 | 80.7035 | 1.6712 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 12 | 1 | 120 | 3.1297 | 0.999 | 60 | 120.0063 | 3.1293 | 99.9004 | 0.37516 | -0.016757 | 0.37554 | 120.216 | 3.0803 | 0.37531 | 0.20973 | -0.048985 | 0.000151 | 0.1748 | -1.5654 | 0.040233
current_sweep | 6 | 1 | 120 | 0.1362 | 0.999 | 60 | 120.0044 | 0.1362 | 99.9008 | 0.016329 | -0.00072653 | 0.016345 | 120.235 | 0.179 | 0.016345 | 0.23059 | 0.042798 | 1.6e-05 | 0.1922 | 31.4226 | 0.1003 | current_sweep | 5 | 1 | 120 | 0.0808 | 0.999 | 60 | 120.0043 | 0.080796 | 99.8998 | 0.0096862 | -0.00043381 | 0.0096959 | 120.2403 | 0.14467 | 0.0098777 | 0.23601 | 0.06387 | 0.00019167 | 0.19667 | 79.051 | 1.9767 | current_sweep | 13 | 1 | 120 | 5.2767 | 0.999 | 60 | 120.01 | 5.2758 | 99.9004 | 0.63252 | -0.028245 | 0.63315 | 120.2033 | 5.1963 | 0.63213 | 0.19338 | -0.079442 | -0.000385 | 0.16113 | -1.5058 | -0.0609
current_sweep | 6 | 2 | 120 | 0.1362 | 0.999 | 60 | 120.0024 | 0.13619 | 99.9005 | 0.016327 | -0.00072776 | 0.016343 | 120.229 | 0.184 | 0.016936 | 0.22657 | 0.047812 | 0.000609 | 0.1888 | 35.107 | 3.7321 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 14 | 1 | 120 | 8.8967 | 0.999 | 60 | 120.0109 | 8.8956 | 99.9005 | 1.0665 | -0.047612 | 1.0676 | 120.1817 | 8.7543 | 1.0655 | 0.1708 | -0.14125 | -0.0010303 | 0.14233 | -1.5878 | -0.0966
current_sweep | 6 | 3 | 120 | 0.1362 | 0.999 | 60 | 120.0045 | 0.1362 | 99.9005 | 0.016329 | -0.0007276 | 0.016345 | 120.228 | 0.179 | 0.016209 | 0.22354 | 0.042796 | -0.00012 | 0.1863 | 31.4201 | -0.7341 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | current_sweep | 15 | 1 | 120 | 15 | 0.999 | 60 | 120.006 | 15.0075 | 99.9041 | 1.7993 | -0.078848 | 1.801 | 120.1273 | 14.7697 | 1.7969 | 0.12136 | -0.23788 | -0.002408 | 0.10113 | -1.5851 | -0.13383
current_sweep | 7 | 1 | 120 | 0.2297 | 0.999 | 60 | 120.0035 | 0.22968 | 99.9006 | 0.027535 | -0.0012281 | 0.027563 | 120.222 | 0.26 | 0.027512 | 0.2185 | 0.030318 | -2.3e-05 | 0.1821 | 13.1999 | -0.0844 | current_sweep | 6 | 1 | 120 | 0.1362 | 0.999 | 60 | 120.0038 | 0.1362 | 99.9006 | 0.016328 | -0.0007273 | 0.016344 | 120.2307 | 0.18067 | 0.016497 | 0.2269 | 0.044469 | 0.00016833 | 0.1891 | 32.6499 | 1.0328 | voltage_sweep | 1 | 1 | 90 | 1 | 0.999 | 60 | 90.0022 | 0.99992 | 99.9002 | 0.089906 | -0.0040195 | 0.089995 | 90.0797 | 0.98667 | 0.089869 | 0.077505 | -0.013258 | -3.6667e-05 | 0.086133 | -1.3259 | -0.041
current_sweep | 7 | 2 | 120 | 0.2297 | 0.999 | 60 | 120.0059 | 0.22967 | 99.9006 | 0.027535 | -0.001228 | 0.027562 | 120.226 | 0.253 | 0.027182 | 0.22012 | 0.023326 | -0.000353 | 0.1834 | 10.156 | -1.2815 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | voltage_sweep | 2 | 1 | 95 | 1 | 0.999 | 60 | 95.0023 | 0.99992 | 99.9002 | 0.0949 | -0.0042425 | 0.094995 | 95.0997 | 0.987 | 0.094905 | 0.09738 | -0.012922 | 4.3333e-06 | 0.1025 | -1.2923 | 0.0048
current_sweep | 7 | 3 | 120 | 0.2297 | 0.999 | 60 | 120.0026 | 0.22968 | 99.9006 | 0.027535 | -0.0012283 | 0.027562 | 120.221 | 0.247 | 0.027321 | 0.21843 | 0.017319 | -0.000214 | 0.182 | 7.5402 | -0.777 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | voltage_sweep | 3 | 1 | 100 | 1 | 0.999 | 60 | 100.0064 | 0.99994 | 99.9001 | 0.099901 | -0.004468 | 0.1 | 100.128 | 0.98733 | 0.099914 | 0.12165 | -0.012609 | 1.3667e-05 | 0.12163 | -1.261 | 0.013633
current_sweep | 8 | 1 | 120 | 0.3873 | 0.999 | 60 | 120.0036 | 0.38728 | 99.9002 | 0.046428 | -0.0020748 | 0.046475 | 120.238 | 0.399 | 0.04676 | 0.23444 | 0.011723 | 0.000332 | 0.1954 | 3.0271 | 0.7146 | current_sweep | 7 | 1 | 120 | 0.2297 | 0.999 | 60 | 120.004 | 0.22968 | 99.9006 | 0.027535 | -0.0012281 | 0.027562 | 120.223 | 0.25333 | 0.027338 | 0.21902 | 0.023654 | -0.00019667 | 0.1825 | 10.2987 | -0.7143 | voltage_sweep | 4 | 1 | 105 | 1 | 0.999 | 60 | 105.0055 | 0.99994 | 99.9001 | 0.10489 | -0.0046904 | 0.105 | 105.153 | 0.98933 | 0.10473 | 0.14749 | -0.010609 | -0.000165 | 0.14047 | -1.0609 | -0.1572
current_sweep | 8 | 2 | 120 | 0.3873 | 0.999 | 60 | 120.0033 | 0.38727 | 99.9003 | 0.046427 | -0.002074 | 0.046473 | 120.241 | 0.397 | 0.046301 | 0.23771 | 0.009734 | -0.000126 | 0.1981 | 2.5134 | -0.2712 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | voltage_sweep | 5 | 1 | 110 | 1 | 0.999 | 60 | 110.0054 | 0.99994 | 99.9002 | 0.10989 | -0.0049136 | 0.11 | 110.175 | 0.991 | 0.11 | 0.16962 | -0.0089353 | 0.000107 | 0.1542 | -0.8936 | 0.0973
current_sweep | 8 | 3 | 120 | 0.3873 | 0.999 | 60 | 120.0036 | 0.38728 | 99.9002 | 0.046428 | -0.0020752 | 0.046475 | 120.235 | 0.403 | 0.047473 | 0.23144 | 0.015723 | 0.001045 | 0.1929 | 4.0598 | 2.2501 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | voltage_sweep | 6 | 1 | 115 | 1 | 0.999 | 60 | 115.002 | 0.99996 | 99.9001 | 0.11488 | -0.0051382 | 0.115 | 115.2013 | 0.99033 | 0.11452 | 0.19933 | -0.0096233 | -0.000363 | 0.17333 | -0.96237 | -0.31583
current_sweep | 9 | 1 | 120 | 0.653 | 0.999 | 60 | 120.0032 | 0.65296 | 99.9003 | 0.078279 | -0.0034966 | 0.078357 | 120.225 | 0.653 | 0.078465 | 0.22181 | 4.2e-05 | 0.000186 | 0.1848 | 0.0065 | 0.2378 | current_sweep | 8 | 1 | 120 | 0.3873 | 0.999 | 60 | 120.0035 | 0.38727 | 99.9003 | 0.046428 | -0.0020747 | 0.046474 | 120.238 | 0.39967 | 0.046845 | 0.23453 | 0.012393 | 0.000417 | 0.19547 | 3.2001 | 0.89783 | voltage_sweep | 7 | 1 | 120 | 1 | 0.999 | 60 | 120.0049 | 0.99994 | 99.9001 | 0.11988 | -0.0053627 | 0.12 | 120.2317 | 0.992 | 0.11989 | 0.22674 | -0.0079397 | 1.4333e-05 | 0.18893 | -0.794 | 0.012167
current_sweep | 9 | 2 | 120 | 0.653 | 0.999 | 60 | 120.0043 | 0.65298 | 99.9001 | 0.078282 | -0.003502 | 0.078361 | 120.222 | 0.653 | 0.078707 | 0.21769 | 1.9e-05 | 0.000425 | 0.1814 | 0.0029 | 0.5426 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | voltage_sweep | 8 | 1 | 125 | 1 | 0.999 | 60 | 125.0061 | 0.99993 | 99.9001 | 0.12487 | -0.0055843 | 0.125 | 125.266 | 0.991 | 0.12473 | 0.25988 | -0.008931 | -0.00014133 | 0.2079 | -0.89317 | -0.1132
current_sweep | 9 | 3 | 120 | 0.653 | 0.999 | 60 | 120.0027 | 0.65298 | 99.9 | 0.078282 | -0.0035031 | 0.07836 | 120.222 | 0.651 | 0.078776 | 0.21925 | -0.001984 | 0.000494 | 0.1827 | -0.3039 | 0.6316 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | voltage_sweep | 9 | 1 | 130 | 1 | 0.999 | 60 | 130.0068 | 0.99994 | 99.9001 | 0.12987 | -0.005809 | 0.13 | 130.2853 | 0.99467 | 0.13008 | 0.27857 | -0.0052753 | 0.00021067 | 0.2143 | -0.52753 | 0.1625
current_sweep | 10 | 1 | 120 | 1.101 | 0.999 | 60 | 120.0043 | 1.1009 | 99.9004 | 0.13198 | -0.0058944 | 0.13211 | 120.225 | 1.096 | 0.13188 | 0.22067 | -0.004896 | -0.000104 | 0.1839 | -0.4447 | -0.0785 | current_sweep | 9 | 1 | 120 | 0.653 | 0.999 | 60 | 120.0034 | 0.65297 | 99.9001 | 0.078281 | -0.0035006 | 0.078359 | 120.223 | 0.65233 | 0.078649 | 0.21958 | -0.000641 | 0.00036833 | 0.18297 | -0.098167 | 0.47067 | pf_sweep | 1 | 1 | 120 | 1 | -0.999 | 60 | 120.0045 | 0.99992 | -99.8998 | 0.11987 | 0.0053705 | 0.11999 | 120.2333 | 0.98567 | 0.1198 | 0.22881 | -0.014253 | -7.0667e-05 | 0.19067 | -1.4254 | -0.059167
current_sweep | 10 | 2 | 120 | 1.101 | 0.999 | 60 | 120.0057 | 1.1009 | 99.9004 | 0.13198 | -0.0058927 | 0.13211 | 120.224 | 1.089 | 0.13207 | 0.2183 | -0.011904 | 8.7e-05 | 0.1819 | -1.0813 | 0.0657 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 2 | 1 | 120 | 1 | -0.899 | 60 | 120.005 | 0.99991 | -89.8984 | 0.10787 | 0.052555 | 0.11999 | 120.236 | 0.98767 | 0.1079 | 0.23095 | -0.012247 | 2.6e-05 | 0.19247 | -1.2248 | 0.0241
current_sweep | 10 | 3 | 120 | 1.101 | 0.999 | 60 | 120.0071 | 1.1009 | 99.9005 | 0.13198 | -0.0058913 | 0.13212 | 120.223 | 1.087 | 0.13177 | 0.21594 | -0.013902 | -0.000211 | 0.1799 | -1.2628 | -0.1596 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 3 | 1 | 120 | 1 | -0.799 | 60 | 120.006 | 0.99993 | -79.8964 | 0.095874 | 0.072164 | 0.12 | 120.233 | 0.987 | 0.095561 | 0.22699 | -0.012929 | -0.000313 | 0.18917 | -1.293 | -0.3262
current_sweep | 11 | 1 | 120 | 1.8563 | 0.999 | 60 | 120.0035 | 1.856 | 99.9003 | 0.2225 | -0.0099371 | 0.22273 | 120.222 | 1.83 | 0.22277 | 0.21848 | -0.026003 | 0.000265 | 0.1821 | -1.401 | 0.1191 | current_sweep | 10 | 1 | 120 | 1.101 | 0.999 | 60 | 120.0057 | 1.1009 | 99.9004 | 0.13198 | -0.0058928 | 0.13211 | 120.224 | 1.0907 | 0.13191 | 0.21831 | -0.010234 | -7.6e-05 | 0.1819 | -0.9296 | -0.057467 | pf_sweep | 4 | 1 | 120 | 1 | -0.699 | 60 | 120.0044 | 0.99994 | -69.8991 | 0.083877 | 0.085814 | 0.12 | 120.2373 | 0.98867 | 0.08365 | 0.23289 | -0.011275 | -0.000227 | 0.19407 | -1.1276 | -0.2704
current_sweep | 11 | 2 | 120 | 1.8563 | 0.999 | 60 | 120.004 | 1.8559 | 99.9004 | 0.2225 | -0.0099379 | 0.22272 | 120.218 | 1.831 | 0.22208 | 0.21396 | -0.024945 | -0.000416 | 0.1783 | -1.344 | -0.1869 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 5 | 1 | 120 | 1 | -0.599 | 60 | 120.0031 | 0.99991 | -59.8994 | 0.071874 | 0.096084 | 0.11999 | 120.2393 | 0.989 | 0.07184 | 0.2362 | -0.010907 | -3.3667e-05 | 0.19683 | -1.0908 | -0.047467
current_sweep | 11 | 3 | 120 | 1.8563 | 0.999 | 60 | 120.0033 | 1.856 | 99.9003 | 0.22251 | -0.0099402 | 0.22273 | 120.22 | 1.827 | 0.22258 | 0.21668 | -0.029017 | 7e-05 | 0.1806 | -1.5634 | 0.0314 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 6 | 1 | 120 | 1 | -0.5 | 60 | 120.0045 | 0.99992 | -49.9978 | 0.059995 | 0.10392 | 0.12 | 120.2397 | 0.98833 | 0.059744 | 0.2352 | -0.01159 | -0.00025133 | 0.196 | -1.1591 | -0.4184
current_sweep | 12 | 1 | 120 | 3.1297 | 0.999 | 60 | 120.0088 | 3.1293 | 99.9004 | 0.37517 | -0.016756 | 0.37555 | 120.218 | 3.081 | 0.37513 | 0.20925 | -0.048348 | -4.7e-05 | 0.1744 | -1.545 | -0.0125 | current_sweep | 11 | 1 | 120 | 1.8563 | 0.999 | 60 | 120.0036 | 1.856 | 99.9003 | 0.2225 | -0.0099384 | 0.22273 | 120.22 | 1.8293 | 0.22248 | 0.21637 | -0.026655 | -2.7e-05 | 0.18033 | -1.4361 | -0.012133 | pf_sweep | 7 | 1 | 120 | 1 | -0.4 | 60 | 120.004 | 0.99992 | -39.998 | 0.047995 | 0.10998 | 0.11999 | 120.2347 | 0.988 | 0.048092 | 0.23062 | -0.011923 | 9.6667e-05 | 0.1922 | -1.1924 | 0.20183
current_sweep | 12 | 2 | 120 | 3.1297 | 0.999 | 60 | 120.0057 | 3.1292 | 99.9003 | 0.37515 | -0.016761 | 0.37553 | 120.215 | 3.081 | 0.37569 | 0.20926 | -0.048247 | 0.000534 | 0.1744 | -1.5418 | 0.1422 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 8 | 1 | 120 | 1 | -0.3 | 60 | 120.0038 | 0.99994 | -29.9964 | 0.035995 | 0.11447 | 0.12 | 120.231 | 0.989 | 0.03605 | 0.22718 | -0.01094 | 5.5333e-05 | 0.1893 | -1.0941 | 0.15367
current_sweep | 12 | 3 | 120 | 3.1297 | 0.999 | 60 | 120.0043 | 3.1294 | 99.9004 | 0.37516 | -0.016755 | 0.37554 | 120.215 | 3.079 | 0.37513 | 0.21068 | -0.050359 | -3.4e-05 | 0.1756 | -1.6093 | -0.009 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 9 | 1 | 120 | 1 | -0.2 | 60 | 120.0039 | 0.99991 | -19.9986 | 0.023997 | 0.11757 | 0.11999 | 120.232 | 0.99267 | 0.023946 | 0.22814 | -0.007246 | -5.1333e-05 | 0.1901 | -0.72467 | -0.2137
current_sweep | 13 | 1 | 120 | 5.2767 | 0.999 | 60 | 120.0102 | 5.2758 | 99.9003 | 0.63252 | -0.028251 | 0.63315 | 120.201 | 5.195 | 0.63226 | 0.19084 | -0.080786 | -0.000261 | 0.159 | -1.5313 | -0.0413 | current_sweep | 12 | 1 | 120 | 3.1297 | 0.999 | 60 | 120.0063 | 3.1293 | 99.9004 | 0.37516 | -0.016757 | 0.37554 | 120.216 | 3.0803 | 0.37531 | 0.20973 | -0.048985 | 0.000151 | 0.1748 | -1.5654 | 0.040233 | pf_sweep | 10 | 1 | 120 | 1 | -0.1 | 60 | 120.0032 | 0.99991 | -9.9971 | 0.011996 | 0.11939 | 0.11999 | 120.2357 | 0.98933 | 0.011937 | 0.2325 | -0.010581 | -5.8667e-05 | 0.19373 | -1.0582 | -0.48993
current_sweep | 13 | 2 | 120 | 5.2767 | 0.999 | 60 | 120.01 | 5.2757 | 99.9005 | 0.63251 | -0.028234 | 0.63314 | 120.204 | 5.198 | 0.63174 | 0.19402 | -0.07772 | -0.000771 | 0.1617 | -1.4732 | -0.1219 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 11 | 1 | 120 | 1 | 0.1 | 60 | 120.0031 | 0.9999 | 10.0037 | 0.012004 | -0.11939 | 0.11999 | 120.234 | 0.99133 | 0.011877 | 0.23095 | -0.0085663 | -0.00012633 | 0.19247 | -0.85673 | -1.0513
current_sweep | 13 | 3 | 120 | 5.2767 | 0.999 | 60 | 120.0097 | 5.2758 | 99.9004 | 0.63252 | -0.028249 | 0.63315 | 120.205 | 5.196 | 0.6324 | 0.19529 | -0.07982 | -0.000123 | 0.1627 | -1.5129 | -0.0195 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 12 | 1 | 120 | 1 | 0.2 | 60 | 120.0041 | 0.99991 | 20.0059 | 0.024006 | -0.11757 | 0.11999 | 120.234 | 0.99233 | 0.023965 | 0.22992 | -0.007576 | -4.0667e-05 | 0.19163 | -0.75767 | -0.16967
current_sweep | 14 | 1 | 120 | 8.8967 | 0.999 | 60 | 120.0114 | 8.8957 | 99.9005 | 1.0665 | -0.047613 | 1.0676 | 120.181 | 8.757 | 1.0659 | 0.16965 | -0.1387 | -0.000646 | 0.1414 | -1.5592 | -0.0605 | current_sweep | 13 | 1 | 120 | 5.2767 | 0.999 | 60 | 120.01 | 5.2758 | 99.9004 | 0.63252 | -0.028245 | 0.63315 | 120.2033 | 5.1963 | 0.63213 | 0.19338 | -0.079442 | -0.000385 | 0.16113 | -1.5058 | -0.0609 | pf_sweep | 13 | 1 | 120 | 1 | 0.3 | 60 | 120.0049 | 0.99991 | 30.0037 | 0.036002 | -0.11447 | 0.11999 | 120.229 | 0.99033 | 0.036136 | 0.22408 | -0.0095723 | 0.000134 | 0.18673 | -0.95733 | 0.37173
current_sweep | 14 | 2 | 120 | 8.8967 | 0.999 | 60 | 120.0099 | 8.8955 | 99.9005 | 1.0665 | -0.047613 | 1.0675 | 120.194 | 8.752 | 1.0653 | 0.18406 | -0.1435 | -0.001182 | 0.1534 | -1.6132 | -0.1108 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 14 | 1 | 120 | 1 | 0.4 | 60 | 120.003 | 0.99991 | 40.0047 | 0.048003 | -0.10997 | 0.11999 | 120.2317 | 0.98933 | 0.048316 | 0.22867 | -0.010578 | 0.000313 | 0.19053 | -1.0578 | 0.6522
current_sweep | 14 | 3 | 120 | 8.8967 | 0.999 | 60 | 120.0113 | 8.8955 | 99.9005 | 1.0665 | -0.047611 | 1.0676 | 120.17 | 8.754 | 1.0652 | 0.15869 | -0.14154 | -0.001263 | 0.1322 | -1.5911 | -0.1185 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 15 | 1 | 120 | 1 | 0.499 | 60 | 120.0028 | 0.99991 | 49.9039 | 0.059881 | -0.10398 | 0.11999 | 120.2363 | 0.98733 | 0.059705 | 0.23353 | -0.012577 | -0.00017567 | 0.1946 | -1.2578 | -0.29333
current_sweep | 15 | 1 | 120 | 15 | 0.999 | 60 | 120.0079 | 15.0075 | 99.904 | 1.7993 | -0.078857 | 1.801 | 120.137 | 14.764 | 1.7971 | 0.12907 | -0.24346 | -0.002223 | 0.1076 | -1.6223 | -0.1235 | current_sweep | 14 | 1 | 120 | 8.8967 | 0.999 | 60 | 120.0109 | 8.8956 | 99.9005 | 1.0665 | -0.047612 | 1.0676 | 120.1817 | 8.7543 | 1.0655 | 0.1708 | -0.14125 | -0.0010303 | 0.14233 | -1.5878 | -0.0966 | pf_sweep | 16 | 1 | 120 | 1 | 0.599 | 60 | 120.0028 | 0.9999 | 59.9039 | 0.071879 | -0.096079 | 0.11999 | 120.232 | 0.99033 | 0.072035 | 0.22924 | -0.009565 | 0.000156 | 0.19103 | -0.9566 | 0.21693
current_sweep | 15 | 2 | 120 | 15 | 0.999 | 60 | 120.0064 | 15.0072 | 99.904 | 1.7992 | -0.078853 | 1.801 | 120.124 | 14.772 | 1.7972 | 0.11762 | -0.23519 | -0.002053 | 0.098 | -1.5672 | -0.1141 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 17 | 1 | 120 | 1 | 0.699 | 60 | 120.005 | 0.99991 | 69.904 | 0.083881 | -0.085806 | 0.11999 | 120.2343 | 0.98967 | 0.083666 | 0.22938 | -0.010242 | -0.00021433 | 0.19113 | -1.0243 | -0.25547
current_sweep | 15 | 3 | 120 | 15 | 0.999 | 60 | 120.0036 | 15.008 | 99.9041 | 1.7993 | -0.078835 | 1.8009 | 120.121 | 14.773 | 1.7963 | 0.11738 | -0.23499 | -0.002948 | 0.0978 | -1.5657 | -0.1639 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 18 | 1 | 120 | 1 | 0.799 | 60 | 120.003 | 0.99992 | 79.902 | 0.095877 | -0.072153 | 0.11999 | 120.231 | 0.99 | 0.09584 | 0.22797 | -0.0099213 | -3.7333e-05 | 0.18993 | -0.9922 | -0.038833
voltage_sweep | 1 | 1 | 90 | 1 | 0.999 | 60 | 90.0022 | 0.99992 | 99.9001 | 0.089905 | -0.0040211 | 0.089995 | 90.076 | 0.986 | 0.089654 | 0.073787 | -0.013921 | -0.000251 | 0.082 | -1.3922 | -0.2794 | current_sweep | 15 | 1 | 120 | 15 | 0.999 | 60 | 120.006 | 15.0075 | 99.9041 | 1.7993 | -0.078848 | 1.801 | 120.1273 | 14.7697 | 1.7969 | 0.12136 | -0.23788 | -0.002408 | 0.10113 | -1.5851 | -0.13383 | pf_sweep | 19 | 1 | 120 | 1 | 0.899 | 60 | 120.0052 | 0.99991 | 89.9023 | 0.10788 | -0.052546 | 0.11999 | 120.232 | 0.989 | 0.10754 | 0.22679 | -0.010912 | -0.00033833 | 0.189 | -1.0913 | -0.3136
voltage_sweep | 1 | 2 | 90 | 1 | 0.999 | 60 | 90.0015 | 0.99994 | 99.9002 | 0.089906 | -0.0040197 | 0.089996 | 90.083 | 0.988 | 0.090075 | 0.081505 | -0.011937 | 0.000169 | 0.0906 | -1.1938 | 0.188 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | pf_sweep | 20 | 1 | 120 | 1 | 0.999 | 60 | 120.0041 | 0.99992 | 99.9001 | 0.11987 | -0.0053605 | 0.11999 | 120.237 | 0.99067 | 0.11988 | 0.2329 | -0.0092487 | 6e-06 | 0.1941 | -0.92497 | 0.0049333
voltage_sweep | 1 | 3 | 90 | 1 | 0.999 | 60 | 90.0028 | 0.99992 | 99.9003 | 0.089905 | -0.0040177 | 0.089995 | 90.08 | 0.986 | 0.089877 | 0.077223 | -0.013916 | -2.8e-05 | 0.0858 | -1.3917 | -0.0316 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | freq_sweep | 1 | 1 | 120 | 1 | 0.999 | 50 | 120.0053 | 0.99991 | 99.9002 | 0.11987 | -0.0053599 | 0.11999 | 120.24 | 0.992 | 0.11968 | 0.23472 | -0.0079097 | -0.00019067 | 0.1956 | -0.791 | -0.15903
voltage_sweep | 2 | 1 | 95 | 1 | 0.999 | 60 | 95.0024 | 0.99998 | 99.9002 | 0.094906 | -0.0042411 | 0.095 | 95.1 | 0.986 | 0.09495 | 0.097574 | -0.013978 | 4.4e-05 | 0.1027 | -1.3978 | 0.0468 | voltage_sweep | 1 | 1 | 90 | 1 | 0.999 | 60 | 90.0022 | 0.99992 | 99.9002 | 0.089906 | -0.0040195 | 0.089995 | 90.0797 | 0.98667 | 0.089869 | 0.077505 | -0.013258 | -3.6667e-05 | 0.086133 | -1.3259 | -0.041 | freq_sweep | 2 | 1 | 120 | 1 | 0.999 | 51 | 120.0091 | 0.99993 | 99.9002 | 0.11988 | -0.0053603 | 0.12 | 120.2443 | 0.99167 | 0.11954 | 0.23523 | -0.008259 | -0.000337 | 0.196 | -0.82597 | -0.28117
voltage_sweep | 2 | 2 | 95 | 1 | 0.999 | 60 | 95.002 | 0.99989 | 99.9001 | 0.094897 | -0.0042441 | 0.094992 | 95.098 | 0.988 | 0.094934 | 0.096039 | -0.011891 | 3.7e-05 | 0.1011 | -1.1892 | 0.0393 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | freq_sweep | 3 | 1 | 120 | 1 | 0.999 | 52 | 120.0062 | 0.99992 | 99.9001 | 0.11988 | -0.0053615 | 0.12 | 120.2417 | 0.99567 | 0.11968 | 0.23546 | -0.004257 | -0.00019433 | 0.19623 | -0.42573 | -0.16197
voltage_sweep | 2 | 3 | 95 | 1 | 0.999 | 60 | 95.0025 | 0.9999 | 99.9002 | 0.094898 | -0.0042422 | 0.094993 | 95.101 | 0.987 | 0.09483 | 0.098528 | -0.012898 | -6.8e-05 | 0.1037 | -1.2899 | -0.0717 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | freq_sweep | 4 | 1 | 120 | 1 | 0.999 | 53 | 120.0064 | 0.99994 | 99.9001 | 0.11988 | -0.005362 | 0.12 | 120.2497 | 0.99 | 0.11983 | 0.24323 | -0.0099417 | -4.8e-05 | 0.2027 | -0.9942 | -0.0402
voltage_sweep | 3 | 1 | 100 | 1 | 0.999 | 60 | 100.0084 | 0.99993 | 99.9003 | 0.099902 | -0.0044646 | 0.1 | 100.127 | 0.992 | 0.099572 | 0.11857 | -0.00793 | -0.00033 | 0.1186 | -0.793 | -0.33 | voltage_sweep | 2 | 1 | 95 | 1 | 0.999 | 60 | 95.0023 | 0.99992 | 99.9002 | 0.0949 | -0.0042425 | 0.094995 | 95.0997 | 0.987 | 0.094905 | 0.09738 | -0.012922 | 4.3333e-06 | 0.1025 | -1.2923 | 0.0048 | freq_sweep | 5 | 1 | 120 | 1 | 0.999 | 54 | 120.0041 | 0.9999 | 99.9 | 0.11987 | -0.0053626 | 0.11999 | 120.2473 | 0.99367 | 0.11982 | 0.2432 | -0.006233 | -5.4e-05 | 0.20267 | -0.62337 | -0.0449
voltage_sweep | 3 | 2 | 100 | 1 | 0.999 | 60 | 100.0058 | 0.99994 | 99.8999 | 0.099899 | -0.0044728 | 0.1 | 100.126 | 0.984 | 0.099924 | 0.12022 | -0.015938 | 2.5e-05 | 0.1202 | -1.5939 | 0.0246 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | freq_sweep | 6 | 1 | 120 | 1 | 0.999 | 55 | 120.0045 | 0.99992 | 99.9002 | 0.11988 | -0.0053593 | 0.12 | 120.2447 | 0.992 | 0.12004 | 0.24015 | -0.0079207 | 0.000163 | 0.2001 | -0.79213 | 0.13607
voltage_sweep | 3 | 3 | 100 | 1 | 0.999 | 60 | 100.0049 | 0.99996 | 99.9002 | 0.099901 | -0.0044666 | 0.1 | 100.131 | 0.986 | 0.10025 | 0.12614 | -0.01396 | 0.000346 | 0.1261 | -1.3961 | 0.3463 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | freq_sweep | 7 | 1 | 120 | 1 | 0.999 | 56 | 120.0047 | 0.99991 | 99.9002 | 0.11987 | -0.0053589 | 0.11999 | 120.2423 | 0.99 | 0.11987 | 0.23763 | -0.0099093 | -6.3333e-06 | 0.198 | -0.991 | -0.0053
voltage_sweep | 4 | 1 | 105 | 1 | 0.999 | 60 | 105.0049 | 0.99996 | 99.9001 | 0.1049 | -0.0046912 | 0.105 | 105.156 | 0.992 | 0.10472 | 0.15113 | -0.007959 | -0.000178 | 0.1439 | -0.7959 | -0.1693 | voltage_sweep | 3 | 1 | 100 | 1 | 0.999 | 60 | 100.0064 | 0.99994 | 99.9001 | 0.099901 | -0.004468 | 0.1 | 100.128 | 0.98733 | 0.099914 | 0.12165 | -0.012609 | 1.3667e-05 | 0.12163 | -1.261 | 0.013633 | freq_sweep | 8 | 1 | 120 | 1 | 0.999 | 57 | 120.0036 | 0.99992 | 99.9001 | 0.11987 | -0.0053607 | 0.11999 | 120.2413 | 0.99033 | 0.1195 | 0.23769 | -0.009585 | -0.00037267 | 0.19807 | -0.9586 | -0.31093
voltage_sweep | 4 | 2 | 105 | 1 | 0.999 | 60 | 105.0053 | 0.99996 | 99.9002 | 0.1049 | -0.0046886 | 0.105 | 105.147 | 0.987 | 0.10482 | 0.14171 | -0.01296 | -7.5e-05 | 0.135 | -1.2961 | -0.0718 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | freq_sweep | 9 | 1 | 120 | 1 | 0.999 | 58 | 120.0036 | 0.99992 | 99.9001 | 0.11987 | -0.0053616 | 0.11999 | 120.236 | 0.991 | 0.11968 | 0.23236 | -0.0089173 | -0.000192 | 0.19363 | -0.89183 | -0.16007
voltage_sweep | 4 | 3 | 105 | 1 | 0.999 | 60 | 105.0064 | 0.99991 | 99.9001 | 0.10489 | -0.0046913 | 0.105 | 105.156 | 0.989 | 0.10465 | 0.14962 | -0.010907 | -0.000242 | 0.1425 | -1.0908 | -0.2305 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | freq_sweep | 10 | 1 | 120 | 1 | 0.999 | 59 | 120.0046 | 0.99991 | 99.9001 | 0.11987 | -0.0053604 | 0.11999 | 120.238 | 0.992 | 0.1199 | 0.23344 | -0.0079107 | 3.1e-05 | 0.19453 | -0.79113 | 0.0259
voltage_sweep | 5 | 1 | 110 | 1 | 0.999 | 60 | 110.0059 | 0.99996 | 99.9003 | 0.10989 | -0.0049112 | 0.11 | 110.178 | 0.989 | 0.10997 | 0.17215 | -0.010956 | 8.1e-05 | 0.1565 | -1.0956 | 0.0735 | voltage_sweep | 4 | 1 | 105 | 1 | 0.999 | 60 | 105.0055 | 0.99994 | 99.9001 | 0.10489 | -0.0046904 | 0.105 | 105.153 | 0.98933 | 0.10473 | 0.14749 | -0.010609 | -0.000165 | 0.14047 | -1.0609 | -0.1572
voltage_sweep | 5 | 2 | 110 | 1 | 0.999 | 60 | 110.0076 | 0.99994 | 99.9002 | 0.10989 | -0.0049127 | 0.11 | 110.175 | 0.991 | 0.11007 | 0.16745 | -0.008936 | 0.000175 | 0.1522 | -0.8937 | 0.1595 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
voltage_sweep | 5 | 3 | 110 | 1 | 0.999 | 60 | 110.0027 | 0.99991 | 99.9 | 0.10988 | -0.0049169 | 0.10999 | 110.172 | 0.993 | 0.10995 | 0.16928 | -0.006914 | 6.5e-05 | 0.1539 | -0.6915 | 0.0589 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
voltage_sweep | 6 | 1 | 115 | 1 | 0.999 | 60 | 115.0034 | 0.99996 | 99.9001 | 0.11488 | -0.0051369 | 0.115 | 115.201 | 0.988 | 0.11462 | 0.19764 | -0.011958 | -0.000267 | 0.1719 | -1.1959 | -0.2322 | voltage_sweep | 5 | 1 | 110 | 1 | 0.999 | 60 | 110.0054 | 0.99994 | 99.9002 | 0.10989 | -0.0049136 | 0.11 | 110.175 | 0.991 | 0.11 | 0.16962 | -0.0089353 | 0.000107 | 0.1542 | -0.8936 | 0.0973
voltage_sweep | 6 | 2 | 115 | 1 | 0.999 | 60 | 115.0014 | 0.99995 | 99.9001 | 0.11488 | -0.0051391 | 0.115 | 115.202 | 0.998 | 0.11432 | 0.2006 | -0.001952 | -0.000559 | 0.1744 | -0.1952 | -0.4866 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
voltage_sweep | 6 | 3 | 115 | 1 | 0.999 | 60 | 115.0012 | 0.99996 | 99.9001 | 0.11488 | -0.0051385 | 0.115 | 115.201 | 0.985 | 0.11462 | 0.19976 | -0.01496 | -0.000263 | 0.1737 | -1.496 | -0.2287 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
voltage_sweep | 7 | 1 | 120 | 1 | 0.999 | 60 | 120.0039 | 0.99993 | 99.9 | 0.11988 | -0.0053632 | 0.12 | 120.225 | 0.987 | 0.11994 | 0.22109 | -0.012934 | 6e-05 | 0.1842 | -1.2935 | 0.05 | voltage_sweep | 6 | 1 | 115 | 1 | 0.999 | 60 | 115.002 | 0.99996 | 99.9001 | 0.11488 | -0.0051382 | 0.115 | 115.2013 | 0.99033 | 0.11452 | 0.19933 | -0.0096233 | -0.000363 | 0.17333 | -0.96237 | -0.31583
voltage_sweep | 7 | 2 | 120 | 1 | 0.999 | 60 | 120.0053 | 0.99995 | 99.9 | 0.11988 | -0.0053644 | 0.12 | 120.239 | 0.992 | 0.1198 | 0.23366 | -0.007953 | -7.5e-05 | 0.1947 | -0.7953 | -0.0623 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
voltage_sweep | 7 | 3 | 120 | 1 | 0.999 | 60 | 120.0055 | 0.99993 | 99.9001 | 0.11988 | -0.0053605 | 0.12 | 120.231 | 0.997 | 0.11994 | 0.22546 | -0.002932 | 5.8e-05 | 0.1879 | -0.2932 | 0.0488 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
voltage_sweep | 8 | 1 | 125 | 1 | 0.999 | 60 | 125.007 | 0.99993 | 99.9003 | 0.12487 | -0.0055815 | 0.125 | 125.267 | 0.994 | 0.1245 | 0.26001 | -0.005929 | -0.000377 | 0.208 | -0.5929 | -0.3022 | voltage_sweep | 7 | 1 | 120 | 1 | 0.999 | 60 | 120.0049 | 0.99994 | 99.9001 | 0.11988 | -0.0053627 | 0.12 | 120.2317 | 0.992 | 0.11989 | 0.22674 | -0.0079397 | 1.4333e-05 | 0.18893 | -0.794 | 0.012167
voltage_sweep | 8 | 2 | 125 | 1 | 0.999 | 60 | 125.0074 | 0.99991 | 99.9001 | 0.12487 | -0.0055843 | 0.125 | 125.27 | 0.99 | 0.12474 | 0.26263 | -0.009912 | -0.000133 | 0.2101 | -0.9913 | -0.1062 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
voltage_sweep | 8 | 3 | 125 | 1 | 0.999 | 60 | 125.004 | 0.99995 | 99.9 | 0.12487 | -0.0055871 | 0.125 | 125.261 | 0.989 | 0.12496 | 0.25699 | -0.010952 | 8.6e-05 | 0.2056 | -1.0953 | 0.0688 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
voltage_sweep | 9 | 1 | 130 | 1 | 0.999 | 60 | 130.0057 | 0.99996 | 99.9001 | 0.12987 | -0.0058084 | 0.13 | 130.288 | 0.998 | 0.13062 | 0.28226 | -0.00196 | 0.00075 | 0.2171 | -0.196 | 0.5777 | voltage_sweep | 8 | 1 | 125 | 1 | 0.999 | 60 | 125.0061 | 0.99993 | 99.9001 | 0.12487 | -0.0055843 | 0.125 | 125.266 | 0.991 | 0.12473 | 0.25988 | -0.008931 | -0.00014133 | 0.2079 | -0.89317 | -0.1132
voltage_sweep | 9 | 2 | 130 | 1 | 0.999 | 60 | 130.0073 | 0.99993 | 99.9001 | 0.12987 | -0.0058093 | 0.13 | 130.284 | 0.993 | 0.12981 | 0.27672 | -0.006933 | -5.9e-05 | 0.2129 | -0.6933 | -0.0451 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
voltage_sweep | 9 | 3 | 130 | 1 | 0.999 | 60 | 130.0073 | 0.99993 | 99.9001 | 0.12987 | -0.0058093 | 0.13 | 130.284 | 0.993 | 0.12981 | 0.27672 | -0.006933 | -5.9e-05 | 0.2129 | -0.6933 | -0.0451 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 1 | 1 | 120 | 1 | -0.999 | 60 | 120.0039 | 0.99991 | -99.8998 | 0.11987 | 0.0053701 | 0.11999 | 120.237 | 0.987 | 0.11978 | 0.23309 | -0.012914 | -9.8e-05 | 0.1942 | -1.2915 | -0.0821 | voltage_sweep | 9 | 1 | 130 | 1 | 0.999 | 60 | 130.0068 | 0.99994 | 99.9001 | 0.12987 | -0.005809 | 0.13 | 130.2853 | 0.99467 | 0.13008 | 0.27857 | -0.0052753 | 0.00021067 | 0.2143 | -0.52753 | 0.1625
pf_sweep | 1 | 2 | 120 | 1 | -0.999 | 60 | 120.0043 | 0.99993 | -99.8997 | 0.11988 | 0.0053714 | 0.12 | 120.233 | 0.985 | 0.11993 | 0.22867 | -0.014935 | 5e-05 | 0.1906 | -1.4936 | 0.0416 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 1 | 3 | 120 | 1 | -0.999 | 60 | 120.0053 | 0.99991 | -99.8998 | 0.11987 | 0.0053699 | 0.11999 | 120.23 | 0.985 | 0.11971 | 0.22468 | -0.01491 | -0.000164 | 0.1872 | -1.4911 | -0.137 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 2 | 1 | 120 | 1 | -0.899 | 60 | 120.0049 | 0.99993 | -89.8976 | 0.10787 | 0.052558 | 0.12 | 120.235 | 0.99 | 0.10776 | 0.23015 | -0.009934 | -0.000116 | 0.1918 | -0.9934 | -0.1078 | pf_sweep | 1 | 1 | 120 | 1 | -0.999 | 60 | 120.0045 | 0.99992 | -99.8998 | 0.11987 | 0.0053705 | 0.11999 | 120.2333 | 0.98567 | 0.1198 | 0.22881 | -0.014253 | -7.0667e-05 | 0.19067 | -1.4254 | -0.059167
pf_sweep | 2 | 2 | 120 | 1 | -0.899 | 60 | 120.0042 | 0.9999 | -89.8991 | 0.10787 | 0.052553 | 0.11999 | 120.24 | 0.985 | 0.10802 | 0.23583 | -0.014898 | 0.000153 | 0.1965 | -1.49 | 0.1421 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 2 | 3 | 120 | 1 | -0.899 | 60 | 120.0061 | 0.99991 | -89.8986 | 0.10787 | 0.052555 | 0.12 | 120.233 | 0.988 | 0.10791 | 0.22689 | -0.011909 | 4.1e-05 | 0.1891 | -1.191 | 0.038 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 3 | 1 | 120 | 1 | -0.799 | 60 | 120.0059 | 0.99994 | -79.895 | 0.095873 | 0.072167 | 0.12 | 120.229 | 0.984 | 0.095459 | 0.22309 | -0.015936 | -0.000414 | 0.1859 | -1.5937 | -0.4314 | pf_sweep | 2 | 1 | 120 | 1 | -0.899 | 60 | 120.005 | 0.99991 | -89.8984 | 0.10787 | 0.052555 | 0.11999 | 120.236 | 0.98767 | 0.1079 | 0.23095 | -0.012247 | 2.6e-05 | 0.19247 | -1.2248 | 0.0241
pf_sweep | 3 | 2 | 120 | 1 | -0.799 | 60 | 120.0058 | 0.9999 | -79.8979 | 0.095873 | 0.07216 | 0.11999 | 120.234 | 0.99 | 0.095526 | 0.2282 | -0.009904 | -0.000347 | 0.1902 | -0.9905 | -0.3618 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 3 | 3 | 120 | 1 | -0.799 | 60 | 120.0063 | 0.99995 | -79.8965 | 0.095876 | 0.072165 | 0.12 | 120.236 | 0.987 | 0.095698 | 0.22968 | -0.012947 | -0.000178 | 0.1914 | -1.2948 | -0.1854 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 4 | 1 | 120 | 1 | -0.699 | 60 | 120.0038 | 0.99993 | -69.9007 | 0.083878 | 0.085811 | 0.12 | 120.238 | 0.989 | 0.083539 | 0.23417 | -0.010934 | -0.000339 | 0.1951 | -1.0935 | -0.4041 | pf_sweep | 3 | 1 | 120 | 1 | -0.799 | 60 | 120.006 | 0.99993 | -79.8964 | 0.095874 | 0.072164 | 0.12 | 120.233 | 0.987 | 0.095561 | 0.22699 | -0.012929 | -0.000313 | 0.18917 | -1.293 | -0.3262
pf_sweep | 4 | 2 | 120 | 1 | -0.699 | 60 | 120.0042 | 0.99993 | -69.8996 | 0.083877 | 0.085812 | 0.12 | 120.237 | 0.99 | 0.08377 | 0.23283 | -0.009933 | -0.000107 | 0.194 | -0.9933 | -0.1274 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 4 | 3 | 120 | 1 | -0.699 | 60 | 120.0053 | 0.99996 | -69.897 | 0.083877 | 0.085818 | 0.12 | 120.237 | 0.987 | 0.083642 | 0.23167 | -0.012958 | -0.000235 | 0.1931 | -1.2959 | -0.2797 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 5 | 1 | 120 | 1 | -0.599 | 60 | 120.0026 | 0.99995 | -59.8989 | 0.071876 | 0.096088 | 0.12 | 120.24 | 0.988 | 0.07212 | 0.23744 | -0.011948 | 0.000244 | 0.1979 | -1.1948 | 0.3388 | pf_sweep | 4 | 1 | 120 | 1 | -0.699 | 60 | 120.0044 | 0.99994 | -69.8991 | 0.083877 | 0.085814 | 0.12 | 120.2373 | 0.98867 | 0.08365 | 0.23289 | -0.011275 | -0.000227 | 0.19407 | -1.1276 | -0.2704
pf_sweep | 5 | 2 | 120 | 1 | -0.599 | 60 | 120.0035 | 0.99988 | -59.8975 | 0.07187 | 0.096083 | 0.11999 | 120.24 | 0.99 | 0.071684 | 0.23653 | -0.009879 | -0.000186 | 0.1971 | -0.988 | -0.2594 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 5 | 3 | 120 | 1 | -0.599 | 60 | 120.0034 | 0.99989 | -59.9016 | 0.071876 | 0.096081 | 0.11999 | 120.238 | 0.989 | 0.071717 | 0.23463 | -0.010895 | -0.000159 | 0.1955 | -1.0896 | -0.2218 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 6 | 1 | 120 | 1 | -0.5 | 60 | 120.0063 | 0.99994 | -49.997 | 0.059996 | 0.10392 | 0.12 | 120.243 | 0.99 | 0.059956 | 0.23673 | -0.009936 | -4e-05 | 0.1973 | -0.9937 | -0.0661 | pf_sweep | 5 | 1 | 120 | 1 | -0.599 | 60 | 120.0031 | 0.99991 | -59.8994 | 0.071874 | 0.096084 | 0.11999 | 120.2393 | 0.989 | 0.07184 | 0.2362 | -0.010907 | -3.3667e-05 | 0.19683 | -1.0908 | -0.047467
pf_sweep | 6 | 2 | 120 | 1 | -0.5 | 60 | 120.0031 | 0.99993 | -49.9978 | 0.059995 | 0.10392 | 0.11999 | 120.238 | 0.989 | 0.059476 | 0.23494 | -0.010928 | -0.000519 | 0.1958 | -1.0929 | -0.8644 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 6 | 3 | 120 | 1 | -0.5 | 60 | 120.0041 | 0.99991 | -49.9987 | 0.059995 | 0.10392 | 0.11999 | 120.238 | 0.986 | 0.0598 | 0.23394 | -0.013905 | -0.000195 | 0.1949 | -1.3907 | -0.3247 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 7 | 1 | 120 | 1 | -0.4 | 60 | 120.0032 | 0.99992 | -40.0001 | 0.047998 | 0.10998 | 0.11999 | 120.231 | 0.992 | 0.047999 | 0.22777 | -0.007925 | 1e-06 | 0.1898 | -0.7925 | 0.0026 | pf_sweep | 6 | 1 | 120 | 1 | -0.5 | 60 | 120.0045 | 0.99992 | -49.9978 | 0.059995 | 0.10392 | 0.12 | 120.2397 | 0.98833 | 0.059744 | 0.2352 | -0.01159 | -0.00025133 | 0.196 | -1.1591 | -0.4184
pf_sweep | 7 | 2 | 120 | 1 | -0.4 | 60 | 120.0044 | 0.9999 | -39.9982 | 0.047995 | 0.10998 | 0.11999 | 120.233 | 0.986 | 0.048122 | 0.22857 | -0.013899 | 0.000127 | 0.1905 | -1.3901 | 0.265 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 7 | 3 | 120 | 1 | -0.4 | 60 | 120.0045 | 0.99994 | -39.9956 | 0.047994 | 0.10998 | 0.12 | 120.24 | 0.986 | 0.048156 | 0.23551 | -0.013944 | 0.000162 | 0.1963 | -1.3945 | 0.3379 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 8 | 1 | 120 | 1 | -0.3 | 60 | 120.0021 | 0.99996 | -29.9972 | 0.035996 | 0.11447 | 0.12 | 120.227 | 0.99 | 0.036402 | 0.22489 | -0.009964 | 0.000406 | 0.1874 | -0.9965 | 1.128 | pf_sweep | 7 | 1 | 120 | 1 | -0.4 | 60 | 120.004 | 0.99992 | -39.998 | 0.047995 | 0.10998 | 0.11999 | 120.2347 | 0.988 | 0.048092 | 0.23062 | -0.011923 | 9.6667e-05 | 0.1922 | -1.1924 | 0.20183
pf_sweep | 8 | 2 | 120 | 1 | -0.3 | 60 | 120.0037 | 0.99994 | -29.9952 | 0.035993 | 0.11447 | 0.12 | 120.234 | 0.991 | 0.036051 | 0.23031 | -0.00894 | 5.8e-05 | 0.1919 | -0.8941 | 0.1607 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 8 | 3 | 120 | 1 | -0.3 | 60 | 120.0057 | 0.99992 | -29.9969 | 0.035995 | 0.11447 | 0.12 | 120.232 | 0.986 | 0.035697 | 0.22632 | -0.013916 | -0.000298 | 0.1886 | -1.3917 | -0.8277 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 9 | 1 | 120 | 1 | -0.2 | 60 | 120.0047 | 0.99991 | -19.997 | 0.023995 | 0.11757 | 0.11999 | 120.233 | 0.996 | 0.024339 | 0.22829 | -0.003915 | 0.000344 | 0.1902 | -0.3915 | 1.4323 | pf_sweep | 8 | 1 | 120 | 1 | -0.3 | 60 | 120.0038 | 0.99994 | -29.9964 | 0.035995 | 0.11447 | 0.12 | 120.231 | 0.989 | 0.03605 | 0.22718 | -0.01094 | 5.5333e-05 | 0.1893 | -1.0941 | 0.15367
pf_sweep | 9 | 2 | 120 | 1 | -0.2 | 60 | 120.0033 | 0.99992 | -20.0017 | 0.024001 | 0.11757 | 0.11999 | 120.23 | 0.992 | 0.023557 | 0.22665 | -0.007917 | -0.000444 | 0.1889 | -0.7918 | -1.8487 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 9 | 3 | 120 | 1 | -0.2 | 60 | 120.0035 | 0.99991 | -19.9971 | 0.023995 | 0.11757 | 0.11999 | 120.233 | 0.99 | 0.023941 | 0.22949 | -0.009906 | -5.4e-05 | 0.1912 | -0.9907 | -0.2247 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 10 | 1 | 120 | 1 | -0.1 | 60 | 120.0031 | 0.99994 | -10.0006 | 0.012 | 0.11939 | 0.12 | 120.238 | 0.989 | 0.011529 | 0.23489 | -0.010942 | -0.000471 | 0.1957 | -1.0942 | -3.9276 | pf_sweep | 9 | 1 | 120 | 1 | -0.2 | 60 | 120.0039 | 0.99991 | -19.9986 | 0.023997 | 0.11757 | 0.11999 | 120.232 | 0.99267 | 0.023946 | 0.22814 | -0.007246 | -5.1333e-05 | 0.1901 | -0.72467 | -0.2137
pf_sweep | 10 | 2 | 120 | 1 | -0.1 | 60 | 120.0041 | 0.9999 | -9.994 | 0.011992 | 0.11939 | 0.11999 | 120.233 | 0.991 | 0.012193 | 0.22888 | -0.008903 | 0.000201 | 0.1907 | -0.8904 | 1.6758 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 10 | 3 | 120 | 1 | -0.1 | 60 | 120.0023 | 0.9999 | -9.9968 | 0.011995 | 0.11939 | 0.11999 | 120.236 | 0.988 | 0.012089 | 0.23372 | -0.011899 | 9.4e-05 | 0.1948 | -1.19 | 0.782 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 11 | 1 | 120 | 1 | 0.1 | 60 | 120.0025 | 0.99989 | 10.0023 | 0.012002 | -0.11939 | 0.11999 | 120.238 | 0.991 | 0.012108 | 0.23547 | -0.008885 | 0.000106 | 0.1962 | -0.8886 | 0.8859 | pf_sweep | 10 | 1 | 120 | 1 | -0.1 | 60 | 120.0032 | 0.99991 | -9.9971 | 0.011996 | 0.11939 | 0.11999 | 120.2357 | 0.98933 | 0.011937 | 0.2325 | -0.010581 | -5.8667e-05 | 0.19373 | -1.0582 | -0.48993
pf_sweep | 11 | 2 | 120 | 1 | 0.1 | 60 | 120.0025 | 0.99991 | 10.0052 | 0.012005 | -0.11939 | 0.11999 | 120.23 | 0.993 | 0.011793 | 0.2275 | -0.00691 | -0.000212 | 0.1896 | -0.6911 | -1.769 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 11 | 3 | 120 | 1 | 0.1 | 60 | 120.0041 | 0.9999 | 10.0036 | 0.012004 | -0.11939 | 0.11999 | 120.234 | 0.99 | 0.011731 | 0.22988 | -0.009904 | -0.000273 | 0.1916 | -0.9905 | -2.2708 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 12 | 1 | 120 | 1 | 0.2 | 60 | 120.0048 | 0.99989 | 20.0091 | 0.024009 | -0.11756 | 0.11999 | 120.236 | 0.993 | 0.024028 | 0.23119 | -0.006886 | 1.9e-05 | 0.1927 | -0.6887 | 0.0785 | pf_sweep | 11 | 1 | 120 | 1 | 0.1 | 60 | 120.0031 | 0.9999 | 10.0037 | 0.012004 | -0.11939 | 0.11999 | 120.234 | 0.99133 | 0.011877 | 0.23095 | -0.0085663 | -0.00012633 | 0.19247 | -0.85673 | -1.0513
pf_sweep | 12 | 2 | 120 | 1 | 0.2 | 60 | 120.004 | 0.99995 | 20.0043 | 0.024005 | -0.11757 | 0.12 | 120.234 | 0.992 | 0.023676 | 0.23 | -0.007947 | -0.000329 | 0.1917 | -0.7947 | -1.3693 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 12 | 3 | 120 | 1 | 0.2 | 60 | 120.0034 | 0.99989 | 20.0043 | 0.024003 | -0.11757 | 0.11999 | 120.232 | 0.992 | 0.024191 | 0.22857 | -0.007895 | 0.000188 | 0.1905 | -0.7896 | 0.7818 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 13 | 1 | 120 | 1 | 0.3 | 60 | 120.0063 | 0.99989 | 30.0051 | 0.036004 | -0.11446 | 0.11999 | 120.234 | 0.988 | 0.03617 | 0.22766 | -0.011895 | 0.000166 | 0.1897 | -1.1896 | 0.4606 | pf_sweep | 12 | 1 | 120 | 1 | 0.2 | 60 | 120.0041 | 0.99991 | 20.0059 | 0.024006 | -0.11757 | 0.11999 | 120.234 | 0.99233 | 0.023965 | 0.22992 | -0.007576 | -4.0667e-05 | 0.19163 | -0.75767 | -0.16967
pf_sweep | 13 | 2 | 120 | 1 | 0.3 | 60 | 120.0053 | 0.99992 | 30.0063 | 0.036006 | -0.11447 | 0.12 | 120.227 | 0.991 | 0.036274 | 0.22174 | -0.008922 | 0.000268 | 0.1848 | -0.8923 | 0.7434 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 13 | 3 | 120 | 1 | 0.3 | 60 | 120.0032 | 0.9999 | 29.9997 | 0.035997 | -0.11446 | 0.11999 | 120.226 | 0.992 | 0.035965 | 0.22285 | -0.0079 | -3.2e-05 | 0.1857 | -0.7901 | -0.0888 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 14 | 1 | 120 | 1 | 0.4 | 60 | 120.0032 | 0.99989 | 40.0039 | 0.048001 | -0.10997 | 0.11999 | 120.231 | 0.986 | 0.048004 | 0.22781 | -0.013892 | 3e-06 | 0.1898 | -1.3894 | 0.0067 | pf_sweep | 13 | 1 | 120 | 1 | 0.3 | 60 | 120.0049 | 0.99991 | 30.0037 | 0.036002 | -0.11447 | 0.11999 | 120.229 | 0.99033 | 0.036136 | 0.22408 | -0.0095723 | 0.000134 | 0.18673 | -0.95733 | 0.37173
pf_sweep | 14 | 2 | 120 | 1 | 0.4 | 60 | 120.0027 | 0.99992 | 40.0067 | 0.048005 | -0.10997 | 0.11999 | 120.228 | 0.99 | 0.048585 | 0.22529 | -0.009924 | 0.00058 | 0.1877 | -0.9924 | 1.2073 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 14 | 3 | 120 | 1 | 0.4 | 60 | 120.0031 | 0.99992 | 40.0036 | 0.048002 | -0.10997 | 0.11999 | 120.236 | 0.992 | 0.048358 | 0.23289 | -0.007917 | 0.000356 | 0.1941 | -0.7917 | 0.7426 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 15 | 1 | 120 | 1 | 0.499 | 60 | 120.0022 | 0.99991 | 49.9028 | 0.059879 | -0.10398 | 0.11999 | 120.238 | 0.989 | 0.059824 | 0.2358 | -0.010913 | -5.5e-05 | 0.1965 | -1.0914 | -0.0922 | pf_sweep | 14 | 1 | 120 | 1 | 0.4 | 60 | 120.003 | 0.99991 | 40.0047 | 0.048003 | -0.10997 | 0.11999 | 120.2317 | 0.98933 | 0.048316 | 0.22867 | -0.010578 | 0.000313 | 0.19053 | -1.0578 | 0.6522
pf_sweep | 15 | 2 | 120 | 1 | 0.499 | 60 | 120.0027 | 0.9999 | 49.9028 | 0.059879 | -0.10398 | 0.11999 | 120.238 | 0.986 | 0.059557 | 0.23534 | -0.013904 | -0.000322 | 0.1961 | -1.3905 | -0.5375 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 15 | 3 | 120 | 1 | 0.499 | 60 | 120.0035 | 0.99991 | 49.9061 | 0.059884 | -0.10398 | 0.11999 | 120.233 | 0.987 | 0.059734 | 0.22945 | -0.012914 | -0.00015 | 0.1912 | -1.2915 | -0.2503 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 16 | 1 | 120 | 1 | 0.599 | 60 | 120.0034 | 0.9999 | 59.9045 | 0.07188 | -0.096079 | 0.11999 | 120.233 | 0.991 | 0.071434 | 0.22964 | -0.008899 | -0.000446 | 0.1914 | -0.89 | -0.6207 | pf_sweep | 15 | 1 | 120 | 1 | 0.499 | 60 | 120.0028 | 0.99991 | 49.9039 | 0.059881 | -0.10398 | 0.11999 | 120.2363 | 0.98733 | 0.059705 | 0.23353 | -0.012577 | -0.00017567 | 0.1946 | -1.2578 | -0.29333
pf_sweep | 16 | 2 | 120 | 1 | 0.599 | 60 | 120.0026 | 0.9999 | 59.9037 | 0.071879 | -0.096079 | 0.11999 | 120.235 | 0.988 | 0.072302 | 0.23237 | -0.011901 | 0.000423 | 0.1936 | -1.1902 | 0.5886 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 16 | 3 | 120 | 1 | 0.599 | 60 | 120.0023 | 0.9999 | 59.9036 | 0.071878 | -0.096078 | 0.11999 | 120.228 | 0.992 | 0.072369 | 0.22571 | -0.007895 | 0.000491 | 0.1881 | -0.7896 | 0.6829 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 17 | 1 | 120 | 1 | 0.699 | 60 | 120.0053 | 0.9999 | 69.905 | 0.083881 | -0.085804 | 0.11999 | 120.234 | 0.99 | 0.084118 | 0.22874 | -0.0099 | 0.000237 | 0.1906 | -0.9901 | 0.2821 | pf_sweep | 16 | 1 | 120 | 1 | 0.599 | 60 | 120.0028 | 0.9999 | 59.9039 | 0.071879 | -0.096079 | 0.11999 | 120.232 | 0.99033 | 0.072035 | 0.22924 | -0.009565 | 0.000156 | 0.19103 | -0.9566 | 0.21693
pf_sweep | 17 | 2 | 120 | 1 | 0.699 | 60 | 120.0042 | 0.99991 | 69.9039 | 0.08388 | -0.085805 | 0.11999 | 120.235 | 0.99 | 0.083629 | 0.23083 | -0.009908 | -0.000251 | 0.1923 | -0.9909 | -0.2991 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 17 | 3 | 120 | 1 | 0.699 | 60 | 120.0054 | 0.99992 | 69.9031 | 0.083881 | -0.085808 | 0.12 | 120.234 | 0.989 | 0.083252 | 0.22856 | -0.010917 | -0.000629 | 0.1905 | -1.0918 | -0.7494 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 18 | 1 | 120 | 1 | 0.799 | 60 | 120.0036 | 0.99991 | 79.9029 | 0.095878 | -0.072151 | 0.11999 | 120.235 | 0.99 | 0.095936 | 0.23141 | -0.009911 | 5.8e-05 | 0.1928 | -0.9912 | 0.0608 | pf_sweep | 17 | 1 | 120 | 1 | 0.699 | 60 | 120.005 | 0.99991 | 69.904 | 0.083881 | -0.085806 | 0.11999 | 120.2343 | 0.98967 | 0.083666 | 0.22938 | -0.010242 | -0.00021433 | 0.19113 | -1.0243 | -0.25547
pf_sweep | 18 | 2 | 120 | 1 | 0.799 | 60 | 120.003 | 0.99993 | 79.9007 | 0.095877 | -0.072156 | 0.12 | 120.226 | 0.989 | 0.09569 | 0.22299 | -0.010935 | -0.000187 | 0.1858 | -1.0935 | -0.195 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 18 | 3 | 120 | 1 | 0.799 | 60 | 120.0025 | 0.99992 | 79.9024 | 0.095877 | -0.072151 | 0.11999 | 120.232 | 0.991 | 0.095894 | 0.2295 | -0.008918 | 1.7e-05 | 0.1912 | -0.8919 | 0.0177 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 19 | 1 | 120 | 1 | 0.899 | 60 | 120.0068 | 0.9999 | 89.9021 | 0.10788 | -0.052546 | 0.11999 | 120.227 | 0.99 | 0.10794 | 0.22016 | -0.009895 | 6.2e-05 | 0.1835 | -0.9896 | 0.0571 | pf_sweep | 18 | 1 | 120 | 1 | 0.799 | 60 | 120.003 | 0.99992 | 79.902 | 0.095877 | -0.072153 | 0.11999 | 120.231 | 0.99 | 0.09584 | 0.22797 | -0.0099213 | -3.7333e-05 | 0.18993 | -0.9922 | -0.038833
pf_sweep | 19 | 2 | 120 | 1 | 0.899 | 60 | 120.0036 | 0.99991 | 89.9016 | 0.10788 | -0.052547 | 0.11999 | 120.236 | 0.989 | 0.10725 | 0.23241 | -0.010907 | -0.000629 | 0.1937 | -1.0908 | -0.5831 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 19 | 3 | 120 | 1 | 0.899 | 60 | 120.0052 | 0.99993 | 89.9034 | 0.10788 | -0.052544 | 0.12 | 120.233 | 0.988 | 0.10743 | 0.2278 | -0.011934 | -0.000448 | 0.1898 | -1.1935 | -0.4148 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 20 | 1 | 120 | 1 | 0.999 | 60 | 120.0038 | 0.99991 | 99.9002 | 0.11987 | -0.0053594 | 0.11999 | 120.237 | 0.99 | 0.11986 | 0.23324 | -0.009906 | -1.6e-05 | 0.1944 | -0.9907 | -0.0131 | pf_sweep | 19 | 1 | 120 | 1 | 0.899 | 60 | 120.0052 | 0.99991 | 89.9023 | 0.10788 | -0.052546 | 0.11999 | 120.232 | 0.989 | 0.10754 | 0.22679 | -0.010912 | -0.00033833 | 0.189 | -1.0913 | -0.3136
pf_sweep | 20 | 2 | 120 | 1 | 0.999 | 60 | 120.0051 | 0.99993 | 99.9001 | 0.11988 | -0.0053606 | 0.12 | 120.241 | 0.992 | 0.11994 | 0.23587 | -0.007925 | 6.5e-05 | 0.1966 | -0.7926 | 0.054 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
pf_sweep | 20 | 3 | 120 | 1 | 0.999 | 60 | 120.0034 | 0.99991 | 99.9001 | 0.11987 | -0.0053615 | 0.11999 | 120.233 | 0.99 | 0.11984 | 0.22957 | -0.009915 | -3.1e-05 | 0.1913 | -0.9916 | -0.0261 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 1 | 1 | 120 | 1 | 0.999 | 50 | 120.0055 | 0.99988 | 99.9002 | 0.11987 | -0.0053596 | 0.11999 | 120.236 | 0.989 | 0.1195 | 0.23052 | -0.010885 | -0.000377 | 0.1921 | -1.0886 | -0.3144 | pf_sweep | 20 | 1 | 120 | 1 | 0.999 | 60 | 120.0041 | 0.99992 | 99.9001 | 0.11987 | -0.0053605 | 0.11999 | 120.237 | 0.99067 | 0.11988 | 0.2329 | -0.0092487 | 6e-06 | 0.1941 | -0.92497 | 0.0049333
freq_sweep | 1 | 2 | 120 | 1 | 0.999 | 50 | 120.007 | 0.99993 | 99.9002 | 0.11988 | -0.0053601 | 0.12 | 120.241 | 0.994 | 0.1196 | 0.23405 | -0.005932 | -0.000277 | 0.195 | -0.5932 | -0.2311 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 1 | 3 | 120 | 1 | 0.999 | 50 | 120.0034 | 0.99991 | 99.9002 | 0.11987 | -0.0053598 | 0.11999 | 120.243 | 0.993 | 0.11995 | 0.2396 | -0.006912 | 8.2e-05 | 0.1997 | -0.6912 | 0.0684 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 2 | 1 | 120 | 1 | 0.999 | 51 | 120.0077 | 0.99993 | 99.9003 | 0.11988 | -0.0053582 | 0.12 | 120.246 | 0.992 | 0.11965 | 0.23834 | -0.007934 | -0.000227 | 0.1986 | -0.7934 | -0.1894 | freq_sweep | 1 | 1 | 120 | 1 | 0.999 | 50 | 120.0053 | 0.99991 | 99.9002 | 0.11987 | -0.0053599 | 0.11999 | 120.24 | 0.992 | 0.11968 | 0.23472 | -0.0079097 | -0.00019067 | 0.1956 | -0.791 | -0.15903
freq_sweep | 2 | 2 | 120 | 1 | 0.999 | 51 | 120.0142 | 0.99993 | 99.9 | 0.11989 | -0.0053648 | 0.12001 | 120.241 | 0.992 | 0.11951 | 0.22683 | -0.007931 | -0.000377 | 0.189 | -0.7932 | -0.3144 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 2 | 3 | 120 | 1 | 0.999 | 51 | 120.0055 | 0.99991 | 99.9003 | 0.11988 | -0.0053578 | 0.11999 | 120.246 | 0.991 | 0.11947 | 0.24051 | -0.008912 | -0.000407 | 0.2004 | -0.8913 | -0.3397 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 3 | 1 | 120 | 1 | 0.999 | 52 | 120.0081 | 0.99994 | 99.9002 | 0.11988 | -0.005361 | 0.12 | 120.245 | 0.997 | 0.12006 | 0.23688 | -0.002937 | 0.000177 | 0.1974 | -0.2937 | 0.1479 | freq_sweep | 2 | 1 | 120 | 1 | 0.999 | 51 | 120.0091 | 0.99993 | 99.9002 | 0.11988 | -0.0053603 | 0.12 | 120.2443 | 0.99167 | 0.11954 | 0.23523 | -0.008259 | -0.000337 | 0.196 | -0.82597 | -0.28117
freq_sweep | 3 | 2 | 120 | 1 | 0.999 | 52 | 120.0032 | 0.99991 | 99.9001 | 0.11987 | -0.0053629 | 0.11999 | 120.237 | 0.991 | 0.11968 | 0.23383 | -0.008914 | -0.000194 | 0.1949 | -0.8915 | -0.1618 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 3 | 3 | 120 | 1 | 0.999 | 52 | 120.0073 | 0.99992 | 99.9001 | 0.11988 | -0.0053606 | 0.12 | 120.243 | 0.999 | 0.11931 | 0.23567 | -0.00092 | -0.000566 | 0.1964 | -0.092 | -0.472 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 4 | 1 | 120 | 1 | 0.999 | 53 | 120.0073 | 0.99994 | 99.9002 | 0.11988 | -0.0053607 | 0.12 | 120.248 | 0.989 | 0.11999 | 0.24072 | -0.010937 | 0.000109 | 0.2006 | -1.0937 | 0.091 | freq_sweep | 3 | 1 | 120 | 1 | 0.999 | 52 | 120.0062 | 0.99992 | 99.9001 | 0.11988 | -0.0053615 | 0.12 | 120.2417 | 0.99567 | 0.11968 | 0.23546 | -0.004257 | -0.00019433 | 0.19623 | -0.42573 | -0.16197
freq_sweep | 4 | 2 | 120 | 1 | 0.999 | 53 | 120.0066 | 0.99994 | 99.9001 | 0.11988 | -0.0053604 | 0.12 | 120.248 | 0.988 | 0.1198 | 0.2414 | -0.011938 | -7.5e-05 | 0.2012 | -1.1939 | -0.0628 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 4 | 3 | 120 | 1 | 0.999 | 53 | 120.0054 | 0.99995 | 99.9 | 0.11988 | -0.0053648 | 0.12 | 120.253 | 0.993 | 0.1197 | 0.24756 | -0.00695 | -0.000178 | 0.2063 | -0.695 | -0.1488 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 5 | 1 | 120 | 1 | 0.999 | 54 | 120.0039 | 0.99989 | 99.9 | 0.11987 | -0.0053639 | 0.11999 | 120.253 | 0.992 | 0.11943 | 0.24915 | -0.007892 | -0.000443 | 0.2076 | -0.7892 | -0.3694 | freq_sweep | 4 | 1 | 120 | 1 | 0.999 | 53 | 120.0064 | 0.99994 | 99.9001 | 0.11988 | -0.005362 | 0.12 | 120.2497 | 0.99 | 0.11983 | 0.24323 | -0.0099417 | -4.8e-05 | 0.2027 | -0.9942 | -0.0402 | [22 empty] | .
freq_sweep | 5 | 2 | 120 | 1 | 0.999 | 54 | 120.0041 | 0.99991 | 99.9 | 0.11987 | -0.0053621 | 0.11999 | 120.244 | 0.996 | 0.11998 | 0.23989 | -0.003914 | 0.000106 | 0.1999 | -0.3915 | 0.0885 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 5 | 3 | 120 | 1 | 0.999 | 54 | 120.0044 | 0.99989 | 99.9001 | 0.11987 | -0.0053619 | 0.11999 | 120.245 | 0.993 | 0.12005 | 0.24055 | -0.006893 | 0.000175 | 0.2005 | -0.6894 | 0.1462 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 6 | 1 | 120 | 1 | 0.999 | 55 | 120.0047 | 0.99994 | 99.9003 | 0.11988 | -0.0053577 | 0.12 | 120.241 | 0.992 | 0.11986 | 0.23631 | -0.007942 | -1.3e-05 | 0.1969 | -0.7942 | -0.0109 | freq_sweep | 5 | 1 | 120 | 1 | 0.999 | 54 | 120.0041 | 0.9999 | 99.9 | 0.11987 | -0.0053626 | 0.11999 | 120.2473 | 0.99367 | 0.11982 | 0.2432 | -0.006233 | -5.4e-05 | 0.20267 | -0.62337 | -0.0449
freq_sweep | 6 | 2 | 120 | 1 | 0.999 | 55 | 120.005 | 0.99988 | 99.9002 | 0.11987 | -0.0053595 | 0.11999 | 120.249 | 0.992 | 0.12056 | 0.24396 | -0.00788 | 0.000689 | 0.2033 | -0.7881 | 0.5749 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 6 | 3 | 120 | 1 | 0.999 | 55 | 120.0038 | 0.99994 | 99.9001 | 0.11988 | -0.0053608 | 0.12 | 120.244 | 0.992 | 0.11969 | 0.24019 | -0.00794 | -0.000187 | 0.2001 | -0.7941 | -0.1558 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 7 | 1 | 120 | 1 | 0.999 | 56 | 120.0035 | 0.99991 | 99.9002 | 0.11987 | -0.0053603 | 0.11999 | 120.238 | 0.99 | 0.11999 | 0.23451 | -0.009909 | 0.000114 | 0.1954 | -0.991 | 0.0953 | freq_sweep | 6 | 1 | 120 | 1 | 0.999 | 55 | 120.0045 | 0.99992 | 99.9002 | 0.11988 | -0.0053593 | 0.12 | 120.2447 | 0.992 | 0.12004 | 0.24015 | -0.0079207 | 0.000163 | 0.2001 | -0.79213 | 0.13607
freq_sweep | 7 | 2 | 120 | 1 | 0.999 | 56 | 120.0039 | 0.99992 | 99.9 | 0.11987 | -0.0053632 | 0.11999 | 120.24 | 0.99 | 0.11984 | 0.23609 | -0.009917 | -2.9e-05 | 0.1967 | -0.9917 | -0.0241 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 7 | 3 | 120 | 1 | 0.999 | 56 | 120.0067 | 0.9999 | 99.9004 | 0.11988 | -0.0053533 | 0.11999 | 120.249 | 0.99 | 0.11977 | 0.24228 | -0.009902 | -0.000104 | 0.2019 | -0.9903 | -0.0871 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 8 | 1 | 120 | 1 | 0.999 | 57 | 120.0035 | 0.99991 | 99.9001 | 0.11987 | -0.0053619 | 0.11999 | 120.243 | 0.99 | 0.11954 | 0.23945 | -0.00991 | -0.000335 | 0.1995 | -0.9911 | -0.2794 | freq_sweep | 7 | 1 | 120 | 1 | 0.999 | 56 | 120.0047 | 0.99991 | 99.9002 | 0.11987 | -0.0053589 | 0.11999 | 120.2423 | 0.99 | 0.11987 | 0.23763 | -0.0099093 | -6.3333e-06 | 0.198 | -0.991 | -0.0053
freq_sweep | 8 | 2 | 120 | 1 | 0.999 | 57 | 120.0038 | 0.99992 | 99.9002 | 0.11987 | -0.0053601 | 0.11999 | 120.245 | 0.992 | 0.11945 | 0.24119 | -0.007917 | -0.000428 | 0.201 | -0.7918 | -0.3571 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 8 | 3 | 120 | 1 | 0.999 | 57 | 120.0036 | 0.99993 | 99.9002 | 0.11988 | -0.00536 | 0.11999 | 120.236 | 0.989 | 0.11952 | 0.23242 | -0.010928 | -0.000355 | 0.1937 | -1.0929 | -0.2963 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 9 | 1 | 120 | 1 | 0.999 | 58 | 120.0043 | 0.99992 | 99.9001 | 0.11987 | -0.0053618 | 0.11999 | 120.235 | 0.99 | 0.11989 | 0.2307 | -0.009919 | 1.9e-05 | 0.1922 | -0.992 | 0.0161 | freq_sweep | 8 | 1 | 120 | 1 | 0.999 | 57 | 120.0036 | 0.99992 | 99.9001 | 0.11987 | -0.0053607 | 0.11999 | 120.2413 | 0.99033 | 0.1195 | 0.23769 | -0.009585 | -0.00037267 | 0.19807 | -0.9586 | -0.31093
freq_sweep | 9 | 2 | 120 | 1 | 0.999 | 58 | 120.0028 | 0.99994 | 99.9001 | 0.11988 | -0.0053635 | 0.12 | 120.238 | 0.99 | 0.11937 | 0.23516 | -0.009938 | -0.000507 | 0.196 | -0.9939 | -0.4225 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 9 | 3 | 120 | 1 | 0.999 | 58 | 120.0038 | 0.9999 | 99.9002 | 0.11987 | -0.0053594 | 0.11999 | 120.235 | 0.993 | 0.11978 | 0.23122 | -0.006895 | -8.8e-05 | 0.1927 | -0.6896 | -0.0738 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 10 | 1 | 120 | 1 | 0.999 | 59 | 120.0034 | 0.99989 | 99.9001 | 0.11987 | -0.0053616 | 0.11999 | 120.244 | 0.992 | 0.12052 | 0.2406 | -0.007893 | 0.000648 | 0.2005 | -0.7894 | 0.5409 | freq_sweep | 9 | 1 | 120 | 1 | 0.999 | 58 | 120.0036 | 0.99992 | 99.9001 | 0.11987 | -0.0053616 | 0.11999 | 120.236 | 0.991 | 0.11968 | 0.23236 | -0.0089173 | -0.000192 | 0.19363 | -0.89183 | -0.16007
freq_sweep | 10 | 2 | 120 | 1 | 0.999 | 59 | 120.0064 | 0.99994 | 99.9002 | 0.11988 | -0.0053578 | 0.12 | 120.235 | 0.994 | 0.11969 | 0.22861 | -0.005938 | -0.000192 | 0.1905 | -0.5938 | -0.1603 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 10 | 3 | 120 | 1 | 0.999 | 59 | 120.0039 | 0.9999 | 99.9001 | 0.11987 | -0.0053618 | 0.11999 | 120.235 | 0.99 | 0.11951 | 0.23111 | -0.009901 | -0.000363 | 0.1926 | -0.9902 | -0.3029 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 11 | 1 | 120 | 1 | 0.999 | 60 | 120.0033 | 0.99993 | 99.9002 | 0.11988 | -0.0053591 | 0.12 | 120.238 | 0.989 | 0.11959 | 0.23466 | -0.010932 | -0.000288 | 0.1955 | -1.0933 | -0.2406 | freq_sweep | 10 | 1 | 120 | 1 | 0.999 | 59 | 120.0046 | 0.99991 | 99.9001 | 0.11987 | -0.0053604 | 0.11999 | 120.238 | 0.992 | 0.1199 | 0.23344 | -0.0079107 | 3.1e-05 | 0.19453 | -0.79113 | 0.0259
freq_sweep | 11 | 2 | 120 | 1 | 0.999 | 60 | 120.0035 | 0.99989 | 99.9 | 0.11987 | -0.0053616 | 0.11999 | 120.237 | 0.985 | 0.11968 | 0.23349 | -0.014894 | -0.000188 | 0.1946 | -1.4895 | -0.1567 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
freq_sweep | 11 | 3 | 120 | 1 | 0.999 | 60 | 120.0035 | 0.99996 | 99.9001 | 0.11988 | -0.0053621 | 0.12 | 120.231 | 0.992 | 0.12039 | 0.22749 | -0.007955 | 0.000516 | 0.1896 | -0.7956 | 0.4302 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
.
5/4/2026
> Testing on Firmware: 1.3.0, Hardware: 1.0
> Current error is terrible but kW and Vac look reasonable
> I did not look at energy accuracy. It's presumably derived from the power readings so should be similarly accurate.
Summary - Don't trust the current reading from this device.
> Since the device does not report power factor you can't try to derive current from the other measurements.
RAW DATA
AVERAGED DATA