## Overview
Overview: | 2026-09-30
> Characterizing the accuracy of the Meross, EM16P energy meter
> Focusing on basic accuracy "sweeps" where a single parameter is varied [Voltage, Current, PF, Freq] while others remain constant
Setup:
Equipment
> Reference | ION9000 - A Class 0.1S power quality and energy meter [calibrated 2025]
> Source | Rotek 8000
> DUT | Meross EM16P [FW: 13.1.4]
Wiring
> Channels A/B/C on the DUT and REF were connected
> 200A CT's [CH: 1,7] used on phase A/B were looped 3 times to provide a reasonable comparison
NOTE
Because the voltages were loaded by the DUT, some minor instability was noted on the voltage readings of the reference
BUG FOUND???
At particularly terrible power factors the device reports the current as zero
In my original test I noted that the current disappeared in my -0.1PF test point.

## Current_Sweep
sweep | point | sample | channel | phase | source_vln | source_i | source_pf | source_freq | ref_vln | ref_i | ref_pf | ref_kW | ref_kVAR | ref_kVA | dut_vln | dut_i | dut_kW | err_vln | err_i | err_kW | pct_err_vln | pct_err_i | pct_err_kW
current_sweep | 1 | 1 | 1 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0066 | 0.010003 | 99.8999 | 0.0011992 | -5.3309e-05 | 0.0012004 | 119.747 | 0 | 0 | -0.25963 | -0.010003 | -0.001199 | -0.2163 | -100 | -100
current_sweep | 1 | 1 | 2 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0066 | 0.010003 | 99.8999 | 0.0011992 | -5.3309e-05 | 0.0012004 | 119.747 | 0 | 0 | -0.25963 | -0.010003 | -0.001199 | -0.2163 | -100 | -100
current_sweep | 1 | 1 | 3 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0066 | 0.010003 | 99.8999 | 0.0011992 | -5.3309e-05 | 0.0012004 | 119.747 | 0 | 0 | -0.25963 | -0.010003 | -0.001199 | -0.2163 | -100 | -100
current_sweep | 1 | 1 | 4 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0066 | 0.010003 | 99.8999 | 0.0011992 | -5.3309e-05 | 0.0012004 | 119.747 | 0 | 0 | -0.25963 | -0.010003 | -0.001199 | -0.2163 | -100 | -100
current_sweep | 1 | 1 | 5 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0066 | 0.010003 | 99.8999 | 0.0011992 | -5.3309e-05 | 0.0012004 | 119.747 | 0 | 0 | -0.25963 | -0.010003 | -0.001199 | -0.2163 | -100 | -100
current_sweep | 1 | 1 | 6 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0066 | 0.010003 | 99.8999 | 0.0011992 | -5.3309e-05 | 0.0012004 | 119.747 | 0 | 0 | -0.25963 | -0.010003 | -0.001199 | -0.2163 | -100 | -100
current_sweep | 1 | 1 | 7 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9781 | 0.010004 | 99.9027 | 0.0011991 | -5.2591e-05 | 0.0012003 | 119.76 | 0 | 0 | -0.21814 | -0.010004 | -0.001199 | -0.1818 | -100 | -100
current_sweep | 1 | 1 | 8 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9781 | 0.010004 | 99.9027 | 0.0011991 | -5.2591e-05 | 0.0012003 | 119.76 | 0 | 0 | -0.21814 | -0.010004 | -0.001199 | -0.1818 | -100 | -100
current_sweep | 1 | 1 | 9 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9781 | 0.010004 | 99.9027 | 0.0011991 | -5.2591e-05 | 0.0012003 | 119.76 | 0 | 0 | -0.21814 | -0.010004 | -0.001199 | -0.1818 | -100 | -100
current_sweep | 1 | 1 | 10 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9781 | 0.010004 | 99.9027 | 0.0011991 | -5.2591e-05 | 0.0012003 | 119.76 | 0 | 0 | -0.21814 | -0.010004 | -0.001199 | -0.1818 | -100 | -100
current_sweep | 1 | 1 | 11 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9781 | 0.010004 | 99.9027 | 0.0011991 | -5.2591e-05 | 0.0012003 | 119.76 | 0 | 0 | -0.21814 | -0.010004 | -0.001199 | -0.1818 | -100 | -100
current_sweep | 1 | 1 | 12 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9781 | 0.010004 | 99.9027 | 0.0011991 | -5.2591e-05 | 0.0012003 | 119.76 | 0 | 0 | -0.21814 | -0.010004 | -0.001199 | -0.1818 | -100 | -100
current_sweep | 1 | 1 | 13 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9899 | 0.010004 | 99.8954 | 0.0011991 | -5.3175e-05 | 0.0012004 | 119.783 | 0 | 0 | -0.20691 | -0.010004 | -0.001199 | -0.1724 | -100 | -100
current_sweep | 1 | 1 | 14 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9899 | 0.010004 | 99.8954 | 0.0011991 | -5.3175e-05 | 0.0012004 | 119.783 | 0 | 0 | -0.20691 | -0.010004 | -0.001199 | -0.1724 | -100 | -100
current_sweep | 1 | 1 | 15 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9899 | 0.010004 | 99.8954 | 0.0011991 | -5.3175e-05 | 0.0012004 | 119.783 | 0 | 0 | -0.20691 | -0.010004 | -0.001199 | -0.1724 | -100 | -100
current_sweep | 1 | 1 | 16 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9899 | 0.010004 | 99.8954 | 0.0011991 | -5.3175e-05 | 0.0012004 | 119.783 | 0 | 0 | -0.20691 | -0.010004 | -0.001199 | -0.1724 | -100 | -100
current_sweep | 1 | 1 | 17 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9899 | 0.010004 | 99.8954 | 0.0011991 | -5.3175e-05 | 0.0012004 | 119.783 | 0 | 0 | -0.20691 | -0.010004 | -0.001199 | -0.1724 | -100 | -100
current_sweep | 1 | 1 | 18 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9899 | 0.010004 | 99.8954 | 0.0011991 | -5.3175e-05 | 0.0012004 | 119.783 | 0 | 0 | -0.20691 | -0.010004 | -0.001199 | -0.1724 | -100 | -100
current_sweep | 1 | 2 | 1 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0039 | 0.010002 | 99.8987 | 0.001199 | -5.353e-05 | 0.0012003 | 119.776 | 0 | 0 | -0.2279 | -0.010002 | -0.001199 | -0.1899 | -100 | -100
current_sweep | 1 | 2 | 2 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0039 | 0.010002 | 99.8987 | 0.001199 | -5.353e-05 | 0.0012003 | 119.776 | 0 | 0 | -0.2279 | -0.010002 | -0.001199 | -0.1899 | -100 | -100
current_sweep | 1 | 2 | 3 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0039 | 0.010002 | 99.8987 | 0.001199 | -5.353e-05 | 0.0012003 | 119.776 | 0 | 0 | -0.2279 | -0.010002 | -0.001199 | -0.1899 | -100 | -100
current_sweep | 1 | 2 | 4 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0039 | 0.010002 | 99.8987 | 0.001199 | -5.353e-05 | 0.0012003 | 119.776 | 0 | 0 | -0.2279 | -0.010002 | -0.001199 | -0.1899 | -100 | -100
current_sweep | 1 | 2 | 5 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0039 | 0.010002 | 99.8987 | 0.001199 | -5.353e-05 | 0.0012003 | 119.776 | 0 | 0 | -0.2279 | -0.010002 | -0.001199 | -0.1899 | -100 | -100
current_sweep | 1 | 2 | 6 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0039 | 0.010002 | 99.8987 | 0.001199 | -5.353e-05 | 0.0012003 | 119.776 | 0 | 0 | -0.2279 | -0.010002 | -0.001199 | -0.1899 | -100 | -100
current_sweep | 1 | 2 | 7 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9868 | 0.010004 | 99.902 | 0.0011991 | -5.2785e-05 | 0.0012003 | 119.769 | 0 | 0 | -0.21782 | -0.010004 | -0.001199 | -0.1815 | -100 | -100
current_sweep | 1 | 2 | 8 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9868 | 0.010004 | 99.902 | 0.0011991 | -5.2785e-05 | 0.0012003 | 119.769 | 0 | 0 | -0.21782 | -0.010004 | -0.001199 | -0.1815 | -100 | -100
current_sweep | 1 | 2 | 9 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9868 | 0.010004 | 99.902 | 0.0011991 | -5.2785e-05 | 0.0012003 | 119.769 | 0 | 0 | -0.21782 | -0.010004 | -0.001199 | -0.1815 | -100 | -100
current_sweep | 1 | 2 | 10 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9868 | 0.010004 | 99.902 | 0.0011991 | -5.2785e-05 | 0.0012003 | 119.769 | 0 | 0 | -0.21782 | -0.010004 | -0.001199 | -0.1815 | -100 | -100
current_sweep | 1 | 2 | 11 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9868 | 0.010004 | 99.902 | 0.0011991 | -5.2785e-05 | 0.0012003 | 119.769 | 0 | 0 | -0.21782 | -0.010004 | -0.001199 | -0.1815 | -100 | -100
current_sweep | 1 | 2 | 12 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9868 | 0.010004 | 99.902 | 0.0011991 | -5.2785e-05 | 0.0012003 | 119.769 | 0 | 0 | -0.21782 | -0.010004 | -0.001199 | -0.1815 | -100 | -100
current_sweep | 1 | 2 | 13 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9906 | 0.010003 | 99.8997 | 0.001199 | -5.3185e-05 | 0.0012002 | 119.771 | 0 | 0 | -0.21956 | -0.010003 | -0.001199 | -0.183 | -100 | -100
current_sweep | 1 | 2 | 14 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9906 | 0.010003 | 99.8997 | 0.001199 | -5.3185e-05 | 0.0012002 | 119.771 | 0 | 0 | -0.21956 | -0.010003 | -0.001199 | -0.183 | -100 | -100
current_sweep | 1 | 2 | 15 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9906 | 0.010003 | 99.8997 | 0.001199 | -5.3185e-05 | 0.0012002 | 119.771 | 0 | 0 | -0.21956 | -0.010003 | -0.001199 | -0.183 | -100 | -100
current_sweep | 1 | 2 | 16 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9906 | 0.010003 | 99.8997 | 0.001199 | -5.3185e-05 | 0.0012002 | 119.771 | 0 | 0 | -0.21956 | -0.010003 | -0.001199 | -0.183 | -100 | -100
current_sweep | 1 | 2 | 17 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9906 | 0.010003 | 99.8997 | 0.001199 | -5.3185e-05 | 0.0012002 | 119.771 | 0 | 0 | -0.21956 | -0.010003 | -0.001199 | -0.183 | -100 | -100
current_sweep | 1 | 2 | 18 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9906 | 0.010003 | 99.8997 | 0.001199 | -5.3185e-05 | 0.0012002 | 119.771 | 0 | 0 | -0.21956 | -0.010003 | -0.001199 | -0.183 | -100 | -100
current_sweep | 1 | 3 | 1 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0003 | 0.010003 | 99.8998 | 0.0011992 | -5.3337e-05 | 0.0012004 | 119.756 | 0 | 0 | -0.24433 | -0.010003 | -0.001199 | -0.2036 | -100 | -100
current_sweep | 1 | 3 | 2 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0003 | 0.010003 | 99.8998 | 0.0011992 | -5.3337e-05 | 0.0012004 | 119.756 | 0 | 0 | -0.24433 | -0.010003 | -0.001199 | -0.2036 | -100 | -100
current_sweep | 1 | 3 | 3 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0003 | 0.010003 | 99.8998 | 0.0011992 | -5.3337e-05 | 0.0012004 | 119.756 | 0 | 0 | -0.24433 | -0.010003 | -0.001199 | -0.2036 | -100 | -100
current_sweep | 1 | 3 | 4 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0003 | 0.010003 | 99.8998 | 0.0011992 | -5.3337e-05 | 0.0012004 | 119.756 | 0 | 0 | -0.24433 | -0.010003 | -0.001199 | -0.2036 | -100 | -100
current_sweep | 1 | 3 | 5 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0003 | 0.010003 | 99.8998 | 0.0011992 | -5.3337e-05 | 0.0012004 | 119.756 | 0 | 0 | -0.24433 | -0.010003 | -0.001199 | -0.2036 | -100 | -100
current_sweep | 1 | 3 | 6 | 1 | 120 | 0.01 | 0.999 | 60 | 120.0003 | 0.010003 | 99.8998 | 0.0011992 | -5.3337e-05 | 0.0012004 | 119.756 | 0 | 0 | -0.24433 | -0.010003 | -0.001199 | -0.2036 | -100 | -100
current_sweep | 1 | 3 | 7 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9771 | 0.010005 | 99.9026 | 0.0011992 | -5.2617e-05 | 0.0012003 | 119.761 | 0 | 0 | -0.21612 | -0.010005 | -0.001199 | -0.1801 | -100 | -100
current_sweep | 1 | 3 | 8 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9771 | 0.010005 | 99.9026 | 0.0011992 | -5.2617e-05 | 0.0012003 | 119.761 | 0 | 0 | -0.21612 | -0.010005 | -0.001199 | -0.1801 | -100 | -100
current_sweep | 1 | 3 | 9 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9771 | 0.010005 | 99.9026 | 0.0011992 | -5.2617e-05 | 0.0012003 | 119.761 | 0 | 0 | -0.21612 | -0.010005 | -0.001199 | -0.1801 | -100 | -100
current_sweep | 1 | 3 | 10 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9771 | 0.010005 | 99.9026 | 0.0011992 | -5.2617e-05 | 0.0012003 | 119.761 | 0 | 0 | -0.21612 | -0.010005 | -0.001199 | -0.1801 | -100 | -100
current_sweep | 1 | 3 | 11 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9771 | 0.010005 | 99.9026 | 0.0011992 | -5.2617e-05 | 0.0012003 | 119.761 | 0 | 0 | -0.21612 | -0.010005 | -0.001199 | -0.1801 | -100 | -100
current_sweep | 1 | 3 | 12 | 2 | 120 | 0.01 | 0.999 | 60 | 119.9771 | 0.010005 | 99.9026 | 0.0011992 | -5.2617e-05 | 0.0012003 | 119.761 | 0 | 0 | -0.21612 | -0.010005 | -0.001199 | -0.1801 | -100 | -100
current_sweep | 1 | 3 | 13 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9931 | 0.010004 | 99.8912 | 0.0011992 | -5.3455e-05 | 0.0012005 | 119.769 | 0 | 0 | -0.22412 | -0.010004 | -0.001199 | -0.1868 | -100 | -100
current_sweep | 1 | 3 | 14 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9931 | 0.010004 | 99.8912 | 0.0011992 | -5.3455e-05 | 0.0012005 | 119.769 | 0 | 0 | -0.22412 | -0.010004 | -0.001199 | -0.1868 | -100 | -100
current_sweep | 1 | 3 | 15 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9931 | 0.010004 | 99.8912 | 0.0011992 | -5.3455e-05 | 0.0012005 | 119.769 | 0 | 0 | -0.22412 | -0.010004 | -0.001199 | -0.1868 | -100 | -100
current_sweep | 1 | 3 | 16 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9931 | 0.010004 | 99.8912 | 0.0011992 | -5.3455e-05 | 0.0012005 | 119.769 | 0 | 0 | -0.22412 | -0.010004 | -0.001199 | -0.1868 | -100 | -100
current_sweep | 1 | 3 | 17 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9931 | 0.010004 | 99.8912 | 0.0011992 | -5.3455e-05 | 0.0012005 | 119.769 | 0 | 0 | -0.22412 | -0.010004 | -0.001199 | -0.1868 | -100 | -100
current_sweep | 1 | 3 | 18 | 3 | 120 | 0.01 | 0.999 | 60 | 119.9931 | 0.010004 | 99.8912 | 0.0011992 | -5.3455e-05 | 0.0012005 | 119.769 | 0 | 0 | -0.22412 | -0.010004 | -0.001199 | -0.1868 | -100 | -100
current_sweep | 2 | 1 | 1 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.016 | 0.017204 | 99.9004 | 0.0020627 | -9.1787e-05 | 0.0020648 | 119.765 | 0.048 | 0.005348 | -0.25098 | 0.030796 | 0.003285 | -0.2091 | -6.9985 | -13.5761
current_sweep | 2 | 1 | 2 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.016 | 0.017204 | 99.9004 | 0.0020627 | -9.1787e-05 | 0.0020648 | 119.765 | 0.018 | 0.002076 | -0.25098 | 0.000796 | 1.3e-05 | -0.2091 | 4.6267 | 0.6448
current_sweep | 2 | 1 | 3 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.016 | 0.017204 | 99.9004 | 0.0020627 | -9.1787e-05 | 0.0020648 | 119.765 | 0.018 | 0.002069 | -0.25098 | 0.000796 | 6e-06 | -0.2091 | 4.6267 | 0.3054
current_sweep | 2 | 1 | 4 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.016 | 0.017204 | 99.9004 | 0.0020627 | -9.1787e-05 | 0.0020648 | 119.765 | 0.018 | 0.002065 | -0.25098 | 0.000796 | 2e-06 | -0.2091 | 4.6267 | 0.1115
current_sweep | 2 | 1 | 5 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.016 | 0.017204 | 99.9004 | 0.0020627 | -9.1787e-05 | 0.0020648 | 119.765 | 0.018 | 0.002065 | -0.25098 | 0.000796 | 2e-06 | -0.2091 | 4.6267 | 0.1115
current_sweep | 2 | 1 | 6 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.016 | 0.017204 | 99.9004 | 0.0020627 | -9.1787e-05 | 0.0020648 | 119.765 | 0.018 | 0.00205 | -0.25098 | 0.000796 | -1.3e-05 | -0.2091 | 4.6267 | -0.6157
current_sweep | 2 | 1 | 7 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9804 | 0.017207 | 99.903 | 0.0020625 | -9.069e-05 | 0.0020645 | 119.779 | 0.051 | 0.005686 | -0.20139 | 0.033793 | 0.003623 | -0.1678 | -1.203 | -8.1051
current_sweep | 2 | 1 | 8 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9804 | 0.017207 | 99.903 | 0.0020625 | -9.069e-05 | 0.0020645 | 119.779 | 0.017 | 0.001951 | -0.20139 | -0.000207 | -0.000112 | -0.1678 | -1.203 | -5.4061
current_sweep | 2 | 1 | 9 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9804 | 0.017207 | 99.903 | 0.0020625 | -9.069e-05 | 0.0020645 | 119.779 | 0.018 | 0.002074 | -0.20139 | 0.000793 | 1.1e-05 | -0.1678 | 4.6086 | 0.5576
current_sweep | 2 | 1 | 10 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9804 | 0.017207 | 99.903 | 0.0020625 | -9.069e-05 | 0.0020645 | 119.779 | 0.018 | 0.002097 | -0.20139 | 0.000793 | 3.4e-05 | -0.1678 | 4.6086 | 1.6727
current_sweep | 2 | 1 | 11 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9804 | 0.017207 | 99.903 | 0.0020625 | -9.069e-05 | 0.0020645 | 119.779 | 0.018 | 0.002093 | -0.20139 | 0.000793 | 3e-05 | -0.1678 | 4.6086 | 1.4788
current_sweep | 2 | 1 | 12 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9804 | 0.017207 | 99.903 | 0.0020625 | -9.069e-05 | 0.0020645 | 119.779 | 0.018 | 0.002091 | -0.20139 | 0.000793 | 2.8e-05 | -0.1678 | 4.6086 | 1.3818
current_sweep | 2 | 1 | 13 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9935 | 0.017203 | 99.9005 | 0.0020622 | -9.172e-05 | 0.0020643 | 119.765 | 0.018 | 0.002079 | -0.22848 | 0.000797 | 1.7e-05 | -0.1904 | 4.6315 | 0.8136
current_sweep | 2 | 1 | 14 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9935 | 0.017203 | 99.9005 | 0.0020622 | -9.172e-05 | 0.0020643 | 119.765 | 0.018 | 0.002081 | -0.22848 | 0.000797 | 1.9e-05 | -0.1904 | 4.6315 | 0.9106
current_sweep | 2 | 1 | 15 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9935 | 0.017203 | 99.9005 | 0.0020622 | -9.172e-05 | 0.0020643 | 119.765 | 0.018 | 0.002083 | -0.22848 | 0.000797 | 2.1e-05 | -0.1904 | 4.6315 | 1.0075
current_sweep | 2 | 1 | 16 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9935 | 0.017203 | 99.9005 | 0.0020622 | -9.172e-05 | 0.0020643 | 119.765 | 0.018 | 0.002053 | -0.22848 | 0.000797 | -9e-06 | -0.1904 | 4.6315 | -0.4472
current_sweep | 2 | 1 | 17 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9935 | 0.017203 | 99.9005 | 0.0020622 | -9.172e-05 | 0.0020643 | 119.765 | 0.018 | 0.002068 | -0.22848 | 0.000797 | 6e-06 | -0.1904 | 4.6315 | 0.2802
current_sweep | 2 | 1 | 18 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9935 | 0.017203 | 99.9005 | 0.0020622 | -9.172e-05 | 0.0020643 | 119.765 | 0.018 | 0.002056 | -0.22848 | 0.000797 | -6e-06 | -0.1904 | 4.6315 | -0.3017
current_sweep | 2 | 2 | 1 | 1 | 120 | 0.0172 | 0.999 | 60 | 119.9918 | 0.017204 | 99.9006 | 0.0020623 | -9.1722e-05 | 0.0020643 | 119.759 | 0.048 | 0.005325 | -0.23278 | 0.030796 | 0.003263 | -0.194 | -6.9969 | -13.9291
current_sweep | 2 | 2 | 2 | 1 | 120 | 0.0172 | 0.999 | 60 | 119.9918 | 0.017204 | 99.9006 | 0.0020623 | -9.1722e-05 | 0.0020643 | 119.759 | 0.018 | 0.002066 | -0.23278 | 0.000796 | 4e-06 | -0.194 | 4.6285 | 0.1816
current_sweep | 2 | 2 | 3 | 1 | 120 | 0.0172 | 0.999 | 60 | 119.9918 | 0.017204 | 99.9006 | 0.0020623 | -9.1722e-05 | 0.0020643 | 119.759 | 0.018 | 0.002079 | -0.23278 | 0.000796 | 1.7e-05 | -0.194 | 4.6285 | 0.812
current_sweep | 2 | 2 | 4 | 1 | 120 | 0.0172 | 0.999 | 60 | 119.9918 | 0.017204 | 99.9006 | 0.0020623 | -9.1722e-05 | 0.0020643 | 119.759 | 0.018 | 0.002066 | -0.23278 | 0.000796 | 4e-06 | -0.194 | 4.6285 | 0.1816
current_sweep | 2 | 2 | 5 | 1 | 120 | 0.0172 | 0.999 | 60 | 119.9918 | 0.017204 | 99.9006 | 0.0020623 | -9.1722e-05 | 0.0020643 | 119.759 | 0.018 | 0.002065 | -0.23278 | 0.000796 | 3e-06 | -0.194 | 4.6285 | 0.1331
current_sweep | 2 | 2 | 6 | 1 | 120 | 0.0172 | 0.999 | 60 | 119.9918 | 0.017204 | 99.9006 | 0.0020623 | -9.1722e-05 | 0.0020643 | 119.759 | 0.018 | 0.00207 | -0.23278 | 0.000796 | 8e-06 | -0.194 | 4.6285 | 0.3756
current_sweep | 2 | 2 | 7 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9788 | 0.017207 | 99.9032 | 0.0020625 | -9.0599e-05 | 0.0020645 | 119.773 | 0.051 | 0.005712 | -0.20576 | 0.033793 | 0.00365 | -0.1715 | -1.2027 | -7.6835
current_sweep | 2 | 2 | 8 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9788 | 0.017207 | 99.9032 | 0.0020625 | -9.0599e-05 | 0.0020645 | 119.773 | 0.018 | 0.00196 | -0.20576 | 0.000793 | -0.000102 | -0.1715 | 4.6089 | -4.9683
current_sweep | 2 | 2 | 9 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9788 | 0.017207 | 99.9032 | 0.0020625 | -9.0599e-05 | 0.0020645 | 119.773 | 0.018 | 0.002064 | -0.20576 | 0.000793 | 2e-06 | -0.1715 | 4.6089 | 0.0742
current_sweep | 2 | 2 | 10 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9788 | 0.017207 | 99.9032 | 0.0020625 | -9.0599e-05 | 0.0020645 | 119.773 | 0.018 | 0.002097 | -0.20576 | 0.000793 | 3.5e-05 | -0.1715 | 4.6089 | 1.6742
current_sweep | 2 | 2 | 11 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9788 | 0.017207 | 99.9032 | 0.0020625 | -9.0599e-05 | 0.0020645 | 119.773 | 0.018 | 0.002083 | -0.20576 | 0.000793 | 2.1e-05 | -0.1715 | 4.6089 | 0.9954
current_sweep | 2 | 2 | 12 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9788 | 0.017207 | 99.9032 | 0.0020625 | -9.0599e-05 | 0.0020645 | 119.773 | 0.018 | 0.002091 | -0.20576 | 0.000793 | 2.9e-05 | -0.1715 | 4.6089 | 1.3833
current_sweep | 2 | 2 | 13 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9921 | 0.017203 | 99.8999 | 0.0020622 | -9.196e-05 | 0.0020643 | 119.77 | 0.018 | 0.002079 | -0.22213 | 0.000797 | 1.7e-05 | -0.1851 | 4.6299 | 0.8138
current_sweep | 2 | 2 | 14 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9921 | 0.017203 | 99.8999 | 0.0020622 | -9.196e-05 | 0.0020643 | 119.77 | 0.018 | 0.002072 | -0.22213 | 0.000797 | 1e-05 | -0.1851 | 4.6299 | 0.4744
current_sweep | 2 | 2 | 15 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9921 | 0.017203 | 99.8999 | 0.0020622 | -9.196e-05 | 0.0020643 | 119.77 | 0.018 | 0.002064 | -0.22213 | 0.000797 | 2e-06 | -0.1851 | 4.6299 | 0.0864
current_sweep | 2 | 2 | 16 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9921 | 0.017203 | 99.8999 | 0.0020622 | -9.196e-05 | 0.0020643 | 119.77 | 0.018 | 0.002053 | -0.22213 | 0.000797 | -9e-06 | -0.1851 | 4.6299 | -0.447
current_sweep | 2 | 2 | 17 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9921 | 0.017203 | 99.8999 | 0.0020622 | -9.196e-05 | 0.0020643 | 119.77 | 0.018 | 0.002058 | -0.22213 | 0.000797 | -4e-06 | -0.1851 | 4.6299 | -0.2045
current_sweep | 2 | 2 | 18 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9921 | 0.017203 | 99.8999 | 0.0020622 | -9.196e-05 | 0.0020643 | 119.77 | 0.018 | 0.002056 | -0.22213 | 0.000797 | -6e-06 | -0.1851 | 4.6299 | -0.3015
current_sweep | 2 | 3 | 1 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.0442 | 0.017202 | 99.9002 | 0.0020629 | -9.1757e-05 | 0.002065 | 119.637 | 0.048 | 0.005348 | -0.40719 | 0.030798 | 0.003285 | -0.3392 | -6.9856 | -13.5843
current_sweep | 2 | 3 | 2 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.0442 | 0.017202 | 99.9002 | 0.0020629 | -9.1757e-05 | 0.002065 | 119.637 | 0.018 | 0.002065 | -0.40719 | 0.000798 | 2e-06 | -0.3392 | 4.6412 | 0.1019
current_sweep | 2 | 3 | 3 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.0442 | 0.017202 | 99.9002 | 0.0020629 | -9.1757e-05 | 0.002065 | 119.637 | 0.018 | 0.002077 | -0.40719 | 0.000798 | 1.4e-05 | -0.3392 | 4.6412 | 0.6837
current_sweep | 2 | 3 | 4 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.0442 | 0.017202 | 99.9002 | 0.0020629 | -9.1757e-05 | 0.002065 | 119.637 | 0.018 | 0.002065 | -0.40719 | 0.000798 | 2e-06 | -0.3392 | 4.6412 | 0.1019
current_sweep | 2 | 3 | 5 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.0442 | 0.017202 | 99.9002 | 0.0020629 | -9.1757e-05 | 0.002065 | 119.637 | 0.018 | 0.002054 | -0.40719 | 0.000798 | -9e-06 | -0.3392 | 4.6412 | -0.4313
current_sweep | 2 | 3 | 6 | 1 | 120 | 0.0172 | 0.999 | 60 | 120.0442 | 0.017202 | 99.9002 | 0.0020629 | -9.1757e-05 | 0.002065 | 119.637 | 0.018 | 0.002058 | -0.40719 | 0.000798 | -5e-06 | -0.3392 | 4.6412 | -0.2374
current_sweep | 2 | 3 | 7 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9832 | 0.017207 | 99.9035 | 0.0020626 | -9.0492e-05 | 0.0020646 | 119.754 | 0.051 | 0.005687 | -0.22925 | 0.033793 | 0.003624 | -0.1911 | -1.205 | -8.0934
current_sweep | 2 | 3 | 8 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9832 | 0.017207 | 99.9035 | 0.0020626 | -9.0492e-05 | 0.0020646 | 119.754 | 0.017 | 0.001941 | -0.22925 | -0.000207 | -0.000122 | -0.1911 | -1.205 | -5.8955
current_sweep | 2 | 3 | 9 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9832 | 0.017207 | 99.9035 | 0.0020626 | -9.0492e-05 | 0.0020646 | 119.754 | 0.018 | 0.002063 | -0.22925 | 0.000793 | 0 | -0.1911 | 4.6064 | 0.0193
current_sweep | 2 | 3 | 10 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9832 | 0.017207 | 99.9035 | 0.0020626 | -9.0492e-05 | 0.0020646 | 119.754 | 0.018 | 0.002077 | -0.22925 | 0.000793 | 1.4e-05 | -0.1911 | 4.6064 | 0.6981
current_sweep | 2 | 3 | 11 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9832 | 0.017207 | 99.9035 | 0.0020626 | -9.0492e-05 | 0.0020646 | 119.754 | 0.018 | 0.002083 | -0.22925 | 0.000793 | 2e-05 | -0.1911 | 4.6064 | 0.989
current_sweep | 2 | 3 | 12 | 2 | 120 | 0.0172 | 0.999 | 60 | 119.9832 | 0.017207 | 99.9035 | 0.0020626 | -9.0492e-05 | 0.0020646 | 119.754 | 0.018 | 0.002091 | -0.22925 | 0.000793 | 2.8e-05 | -0.1911 | 4.6064 | 1.3768
current_sweep | 2 | 3 | 13 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9892 | 0.017204 | 99.8993 | 0.0020622 | -9.2121e-05 | 0.0020643 | 119.773 | 0.018 | 0.00207 | -0.2162 | 0.000796 | 8e-06 | -0.1802 | 4.6266 | 0.3773
current_sweep | 2 | 3 | 14 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9892 | 0.017204 | 99.8993 | 0.0020622 | -9.2121e-05 | 0.0020643 | 119.773 | 0.018 | 0.002072 | -0.2162 | 0.000796 | 1e-05 | -0.1802 | 4.6266 | 0.4743
current_sweep | 2 | 3 | 15 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9892 | 0.017204 | 99.8993 | 0.0020622 | -9.2121e-05 | 0.0020643 | 119.773 | 0.018 | 0.002074 | -0.2162 | 0.000796 | 1.2e-05 | -0.1802 | 4.6266 | 0.5713
current_sweep | 2 | 3 | 16 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9892 | 0.017204 | 99.8993 | 0.0020622 | -9.2121e-05 | 0.0020643 | 119.773 | 0.018 | 0.002053 | -0.2162 | 0.000796 | -9e-06 | -0.1802 | 4.6266 | -0.4471
current_sweep | 2 | 3 | 17 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9892 | 0.017204 | 99.8993 | 0.0020622 | -9.2121e-05 | 0.0020643 | 119.773 | 0.018 | 0.002078 | -0.2162 | 0.000796 | 1.6e-05 | -0.1802 | 4.6266 | 0.7652
current_sweep | 2 | 3 | 18 | 3 | 120 | 0.0172 | 0.999 | 60 | 119.9892 | 0.017204 | 99.8993 | 0.0020622 | -9.2121e-05 | 0.0020643 | 119.773 | 0.018 | 0.002066 | -0.2162 | 0.000796 | 4e-06 | -0.1802 | 4.6266 | 0.1833
current_sweep | 3 | 1 | 1 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0095 | 0.029602 | 99.9006 | 0.003549 | -0.00015824 | 0.0035526 | 119.76 | 0.079 | 0.00921 | -0.24952 | 0.049398 | 0.005661 | -0.2079 | -11.0432 | -13.4976
current_sweep | 3 | 1 | 2 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0095 | 0.029602 | 99.9006 | 0.003549 | -0.00015824 | 0.0035526 | 119.76 | 0.03 | 0.003561 | -0.24952 | 0.000398 | 1.2e-05 | -0.2079 | 1.3432 | 0.3371
current_sweep | 3 | 1 | 3 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0095 | 0.029602 | 99.9006 | 0.003549 | -0.00015824 | 0.0035526 | 119.76 | 0.03 | 0.003559 | -0.24952 | 0.000398 | 1e-05 | -0.2079 | 1.3432 | 0.2808
current_sweep | 3 | 1 | 4 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0095 | 0.029602 | 99.9006 | 0.003549 | -0.00015824 | 0.0035526 | 119.76 | 0.03 | 0.003563 | -0.24952 | 0.000398 | 1.4e-05 | -0.2079 | 1.3432 | 0.3935
current_sweep | 3 | 1 | 5 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0095 | 0.029602 | 99.9006 | 0.003549 | -0.00015824 | 0.0035526 | 119.76 | 0.03 | 0.003561 | -0.24952 | 0.000398 | 1.2e-05 | -0.2079 | 1.3432 | 0.3371
current_sweep | 3 | 1 | 6 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0095 | 0.029602 | 99.9006 | 0.003549 | -0.00015824 | 0.0035526 | 119.76 | 0.03 | 0.00355 | -0.24952 | 0.000398 | 1e-06 | -0.2079 | 1.3432 | 0.0272
current_sweep | 3 | 1 | 7 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9798 | 0.02961 | 99.9034 | 0.0035492 | -0.00015598 | 0.0035526 | 119.764 | 0.083 | 0.009811 | -0.2158 | 0.05339 | 0.006262 | -0.1799 | -6.5638 | -7.8571
current_sweep | 3 | 1 | 8 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9798 | 0.02961 | 99.9034 | 0.0035492 | -0.00015598 | 0.0035526 | 119.764 | 0.029 | 0.003446 | -0.2158 | -0.00061 | -0.000103 | -0.1799 | -2.0608 | -2.9076
current_sweep | 3 | 1 | 9 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9798 | 0.02961 | 99.9034 | 0.0035492 | -0.00015598 | 0.0035526 | 119.764 | 0.03 | 0.003578 | -0.2158 | 0.00039 | 2.9e-05 | -0.1799 | 1.3164 | 0.8116
current_sweep | 3 | 1 | 10 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9798 | 0.02961 | 99.9034 | 0.0035492 | -0.00015598 | 0.0035526 | 119.764 | 0.03 | 0.003594 | -0.2158 | 0.00039 | 4.5e-05 | -0.1799 | 1.3164 | 1.2624
current_sweep | 3 | 1 | 11 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9798 | 0.02961 | 99.9034 | 0.0035492 | -0.00015598 | 0.0035526 | 119.764 | 0.03 | 0.003576 | -0.2158 | 0.00039 | 2.7e-05 | -0.1799 | 1.3164 | 0.7552
current_sweep | 3 | 1 | 12 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9798 | 0.02961 | 99.9034 | 0.0035492 | -0.00015598 | 0.0035526 | 119.764 | 0.03 | 0.003573 | -0.2158 | 0.00039 | 2.4e-05 | -0.1799 | 1.3164 | 0.6707
current_sweep | 3 | 1 | 13 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9921 | 0.029603 | 99.9004 | 0.0035486 | -0.00015828 | 0.0035521 | 119.777 | 0.03 | 0.003572 | -0.21512 | 0.000397 | 2.3e-05 | -0.1793 | 1.3414 | 0.6601
current_sweep | 3 | 1 | 14 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9921 | 0.029603 | 99.9004 | 0.0035486 | -0.00015828 | 0.0035521 | 119.777 | 0.03 | 0.003566 | -0.21512 | 0.000397 | 1.7e-05 | -0.1793 | 1.3414 | 0.4911
current_sweep | 3 | 1 | 15 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9921 | 0.029603 | 99.9004 | 0.0035486 | -0.00015828 | 0.0035521 | 119.777 | 0.03 | 0.003556 | -0.21512 | 0.000397 | 7e-06 | -0.1793 | 1.3414 | 0.2093
current_sweep | 3 | 1 | 16 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9921 | 0.029603 | 99.9004 | 0.0035486 | -0.00015828 | 0.0035521 | 119.777 | 0.03 | 0.003564 | -0.21512 | 0.000397 | 1.5e-05 | -0.1793 | 1.3414 | 0.4347
current_sweep | 3 | 1 | 17 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9921 | 0.029603 | 99.9004 | 0.0035486 | -0.00015828 | 0.0035521 | 119.777 | 0.03 | 0.003564 | -0.21512 | 0.000397 | 1.5e-05 | -0.1793 | 1.3414 | 0.4347
current_sweep | 3 | 1 | 18 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9921 | 0.029603 | 99.9004 | 0.0035486 | -0.00015828 | 0.0035521 | 119.777 | 0.03 | 0.003579 | -0.21512 | 0.000397 | 3e-05 | -0.1793 | 1.3414 | 0.8574
current_sweep | 3 | 2 | 1 | 1 | 120 | 0.0296 | 0.999 | 60 | 119.9823 | 0.029602 | 99.9 | 0.0035481 | -0.00015817 | 0.0035517 | 119.737 | 0.079 | 0.00916 | -0.24526 | 0.049398 | 0.005612 | -0.2044 | -11.0406 | -13.9446
current_sweep | 3 | 2 | 2 | 1 | 120 | 0.0296 | 0.999 | 60 | 119.9823 | 0.029602 | 99.9 | 0.0035481 | -0.00015817 | 0.0035517 | 119.737 | 0.03 | 0.00356 | -0.24526 | 0.000398 | 1.2e-05 | -0.2044 | 1.3462 | 0.3353
current_sweep | 3 | 2 | 3 | 1 | 120 | 0.0296 | 0.999 | 60 | 119.9823 | 0.029602 | 99.9 | 0.0035481 | -0.00015817 | 0.0035517 | 119.737 | 0.03 | 0.003549 | -0.24526 | 0.000398 | 1e-06 | -0.2044 | 1.3462 | 0.0253
current_sweep | 3 | 2 | 4 | 1 | 120 | 0.0296 | 0.999 | 60 | 119.9823 | 0.029602 | 99.9 | 0.0035481 | -0.00015817 | 0.0035517 | 119.737 | 0.03 | 0.003572 | -0.24526 | 0.000398 | 2.4e-05 | -0.2044 | 1.3462 | 0.6735
current_sweep | 3 | 2 | 5 | 1 | 120 | 0.0296 | 0.999 | 60 | 119.9823 | 0.029602 | 99.9 | 0.0035481 | -0.00015817 | 0.0035517 | 119.737 | 0.03 | 0.003569 | -0.24526 | 0.000398 | 2.1e-05 | -0.2044 | 1.3462 | 0.5889
current_sweep | 3 | 2 | 6 | 1 | 120 | 0.0296 | 0.999 | 60 | 119.9823 | 0.029602 | 99.9 | 0.0035481 | -0.00015817 | 0.0035517 | 119.737 | 0.03 | 0.00354 | -0.24526 | 0.000398 | -8e-06 | -0.2044 | 1.3462 | -0.2284
current_sweep | 3 | 2 | 7 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9795 | 0.029611 | 99.9037 | 0.0035492 | -0.00015576 | 0.0035527 | 119.78 | 0.083 | 0.009785 | -0.19949 | 0.053389 | 0.006236 | -0.1663 | -6.5652 | -8.1026
current_sweep | 3 | 2 | 8 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9795 | 0.029611 | 99.9037 | 0.0035492 | -0.00015576 | 0.0035527 | 119.78 | 0.029 | 0.003455 | -0.19949 | -0.000611 | -9.4e-05 | -0.1663 | -2.0623 | -2.6554
current_sweep | 3 | 2 | 9 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9795 | 0.029611 | 99.9037 | 0.0035492 | -0.00015576 | 0.0035527 | 119.78 | 0.03 | 0.003568 | -0.19949 | 0.000389 | 1.9e-05 | -0.1663 | 1.3149 | 0.5283
current_sweep | 3 | 2 | 10 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9795 | 0.029611 | 99.9037 | 0.0035492 | -0.00015576 | 0.0035527 | 119.78 | 0.03 | 0.003594 | -0.19949 | 0.000389 | 4.5e-05 | -0.1663 | 1.3149 | 1.2609
current_sweep | 3 | 2 | 11 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9795 | 0.029611 | 99.9037 | 0.0035492 | -0.00015576 | 0.0035527 | 119.78 | 0.03 | 0.003585 | -0.19949 | 0.000389 | 3.6e-05 | -0.1663 | 1.3149 | 1.0073
current_sweep | 3 | 2 | 12 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9795 | 0.029611 | 99.9037 | 0.0035492 | -0.00015576 | 0.0035527 | 119.78 | 0.03 | 0.003573 | -0.19949 | 0.000389 | 2.4e-05 | -0.1663 | 1.3149 | 0.6692
current_sweep | 3 | 2 | 13 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9914 | 0.029604 | 99.9 | 0.0035487 | -0.00015848 | 0.0035522 | 119.782 | 0.03 | 0.003582 | -0.20941 | 0.000396 | 3.3e-05 | -0.1745 | 1.3373 | 0.9388
current_sweep | 3 | 2 | 14 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9914 | 0.029604 | 99.9 | 0.0035487 | -0.00015848 | 0.0035522 | 119.782 | 0.03 | 0.003585 | -0.20941 | 0.000396 | 3.6e-05 | -0.1745 | 1.3373 | 1.0233
current_sweep | 3 | 2 | 15 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9914 | 0.029604 | 99.9 | 0.0035487 | -0.00015848 | 0.0035522 | 119.782 | 0.03 | 0.003556 | -0.20941 | 0.000396 | 7e-06 | -0.1745 | 1.3373 | 0.2061
current_sweep | 3 | 2 | 16 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9914 | 0.029604 | 99.9 | 0.0035487 | -0.00015848 | 0.0035522 | 119.782 | 0.03 | 0.003564 | -0.20941 | 0.000396 | 1.5e-05 | -0.1745 | 1.3373 | 0.4315
current_sweep | 3 | 2 | 17 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9914 | 0.029604 | 99.9 | 0.0035487 | -0.00015848 | 0.0035522 | 119.782 | 0.03 | 0.003564 | -0.20941 | 0.000396 | 1.5e-05 | -0.1745 | 1.3373 | 0.4315
current_sweep | 3 | 2 | 18 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9914 | 0.029604 | 99.9 | 0.0035487 | -0.00015848 | 0.0035522 | 119.782 | 0.03 | 0.00358 | -0.20941 | 0.000396 | 3.1e-05 | -0.1745 | 1.3373 | 0.8824
current_sweep | 3 | 3 | 1 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0131 | 0.029602 | 99.9009 | 0.0035491 | -0.00015789 | 0.0035526 | 119.737 | 0.079 | 0.009235 | -0.27606 | 0.049398 | 0.005686 | -0.23 | -11.0418 | -13.2643
current_sweep | 3 | 3 | 2 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0131 | 0.029602 | 99.9009 | 0.0035491 | -0.00015789 | 0.0035526 | 119.737 | 0.03 | 0.00357 | -0.27606 | 0.000398 | 2.1e-05 | -0.23 | 1.3448 | 0.589
current_sweep | 3 | 3 | 3 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0131 | 0.029602 | 99.9009 | 0.0035491 | -0.00015789 | 0.0035526 | 119.737 | 0.03 | 0.003558 | -0.27606 | 0.000398 | 9e-06 | -0.23 | 1.3448 | 0.2508
current_sweep | 3 | 3 | 4 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0131 | 0.029602 | 99.9009 | 0.0035491 | -0.00015789 | 0.0035526 | 119.737 | 0.03 | 0.003544 | -0.27606 | 0.000398 | -5e-06 | -0.23 | 1.3448 | -0.1436
current_sweep | 3 | 3 | 5 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0131 | 0.029602 | 99.9009 | 0.0035491 | -0.00015789 | 0.0035526 | 119.737 | 0.03 | 0.003579 | -0.27606 | 0.000398 | 3e-05 | -0.23 | 1.3448 | 0.8425
current_sweep | 3 | 3 | 6 | 1 | 120 | 0.0296 | 0.999 | 60 | 120.0131 | 0.029602 | 99.9009 | 0.0035491 | -0.00015789 | 0.0035526 | 119.737 | 0.03 | 0.00355 | -0.27606 | 0.000398 | 1e-06 | -0.23 | 1.3448 | 0.0254
current_sweep | 3 | 3 | 7 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9793 | 0.029611 | 99.9038 | 0.0035492 | -0.00015568 | 0.0035527 | 119.768 | 0.084 | 0.009836 | -0.21127 | 0.054389 | 0.006287 | -0.1761 | -5.4394 | -7.6236
current_sweep | 3 | 3 | 8 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9793 | 0.029611 | 99.9038 | 0.0035492 | -0.00015568 | 0.0035527 | 119.768 | 0.029 | 0.003465 | -0.21127 | -0.000611 | -8.4e-05 | -0.1761 | -2.0622 | -2.3736
current_sweep | 3 | 3 | 9 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9793 | 0.029611 | 99.9038 | 0.0035492 | -0.00015568 | 0.0035527 | 119.768 | 0.03 | 0.003559 | -0.21127 | 0.000389 | 1e-05 | -0.1761 | 1.3149 | 0.2748
current_sweep | 3 | 3 | 10 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9793 | 0.029611 | 99.9038 | 0.0035492 | -0.00015568 | 0.0035527 | 119.768 | 0.03 | 0.003585 | -0.21127 | 0.000389 | 3.6e-05 | -0.1761 | 1.3149 | 1.0074
current_sweep | 3 | 3 | 11 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9793 | 0.029611 | 99.9038 | 0.0035492 | -0.00015568 | 0.0035527 | 119.768 | 0.03 | 0.003565 | -0.21127 | 0.000389 | 1.6e-05 | -0.1761 | 1.3149 | 0.4439
current_sweep | 3 | 3 | 12 | 2 | 120 | 0.0296 | 0.999 | 60 | 119.9793 | 0.029611 | 99.9038 | 0.0035492 | -0.00015568 | 0.0035527 | 119.768 | 0.03 | 0.003592 | -0.21127 | 0.000389 | 4.3e-05 | -0.1761 | 1.3149 | 1.2046
current_sweep | 3 | 3 | 13 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9878 | 0.029604 | 99.9003 | 0.0035485 | -0.00015828 | 0.0035521 | 119.78 | 0.03 | 0.003562 | -0.20778 | 0.000396 | 1.3e-05 | -0.1732 | 1.3387 | 0.3794
current_sweep | 3 | 3 | 14 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9878 | 0.029604 | 99.9003 | 0.0035485 | -0.00015828 | 0.0035521 | 119.78 | 0.03 | 0.003556 | -0.20778 | 0.000396 | 7e-06 | -0.1732 | 1.3387 | 0.2103
current_sweep | 3 | 3 | 15 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9878 | 0.029604 | 99.9003 | 0.0035485 | -0.00015828 | 0.0035521 | 119.78 | 0.03 | 0.003576 | -0.20778 | 0.000396 | 2.7e-05 | -0.1732 | 1.3387 | 0.7739
current_sweep | 3 | 3 | 16 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9878 | 0.029604 | 99.9003 | 0.0035485 | -0.00015828 | 0.0035521 | 119.78 | 0.03 | 0.003545 | -0.20778 | 0.000396 | -4e-06 | -0.1732 | 1.3387 | -0.0997
current_sweep | 3 | 3 | 17 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9878 | 0.029604 | 99.9003 | 0.0035485 | -0.00015828 | 0.0035521 | 119.78 | 0.03 | 0.003564 | -0.20778 | 0.000396 | 1.5e-05 | -0.1732 | 1.3387 | 0.4357
current_sweep | 3 | 3 | 18 | 3 | 120 | 0.0296 | 0.999 | 60 | 119.9878 | 0.029604 | 99.9003 | 0.0035485 | -0.00015828 | 0.0035521 | 119.78 | 0.03 | 0.00358 | -0.20778 | 0.000396 | 3.1e-05 | -0.1732 | 1.3387 | 0.8866
current_sweep | 4 | 1 | 1 | 1 | 120 | 0.051 | 0.999 | 60 | 119.9965 | 0.050999 | 99.9006 | 0.0061136 | -0.00027235 | 0.0061197 | 119.748 | 0.135 | 0.016039 | -0.24854 | 0.084001 | 0.009925 | -0.2071 | -11.7632 | -12.5507
current_sweep | 4 | 1 | 2 | 1 | 120 | 0.051 | 0.999 | 60 | 119.9965 | 0.050999 | 99.9006 | 0.0061136 | -0.00027235 | 0.0061197 | 119.748 | 0.051 | 0.006159 | -0.24854 | 1e-06 | 4.5e-05 | -0.2071 | 0.0017 | 0.742
current_sweep | 4 | 1 | 3 | 1 | 120 | 0.051 | 0.999 | 60 | 119.9965 | 0.050999 | 99.9006 | 0.0061136 | -0.00027235 | 0.0061197 | 119.748 | 0.051 | 0.006159 | -0.24854 | 1e-06 | 4.5e-05 | -0.2071 | 0.0017 | 0.742
current_sweep | 4 | 1 | 4 | 1 | 120 | 0.051 | 0.999 | 60 | 119.9965 | 0.050999 | 99.9006 | 0.0061136 | -0.00027235 | 0.0061197 | 119.748 | 0.051 | 0.006137 | -0.24854 | 1e-06 | 2.3e-05 | -0.2071 | 0.0017 | 0.3821
current_sweep | 4 | 1 | 5 | 1 | 120 | 0.051 | 0.999 | 60 | 119.9965 | 0.050999 | 99.9006 | 0.0061136 | -0.00027235 | 0.0061197 | 119.748 | 0.051 | 0.006135 | -0.24854 | 1e-06 | 2.1e-05 | -0.2071 | 0.0017 | 0.3494
current_sweep | 4 | 1 | 6 | 1 | 120 | 0.051 | 0.999 | 60 | 119.9965 | 0.050999 | 99.9006 | 0.0061136 | -0.00027235 | 0.0061197 | 119.748 | 0.051 | 0.006133 | -0.24854 | 1e-06 | 1.9e-05 | -0.2071 | 0.0017 | 0.3167
current_sweep | 4 | 1 | 7 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9737 | 0.051016 | 99.9037 | 0.0061147 | -0.00026838 | 0.0061206 | 119.781 | 0.143 | 0.017037 | -0.19267 | 0.091984 | 0.010922 | -0.1606 | -6.5653 | -7.1252
current_sweep | 4 | 1 | 8 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9737 | 0.051016 | 99.9037 | 0.0061147 | -0.00026838 | 0.0061206 | 119.781 | 0.05 | 0.006038 | -0.19267 | -0.001016 | -7.7e-05 | -0.1606 | -1.9915 | -1.2541
current_sweep | 4 | 1 | 9 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9737 | 0.051016 | 99.9037 | 0.0061147 | -0.00026838 | 0.0061206 | 119.781 | 0.051 | 0.006135 | -0.19267 | -1.6e-05 | 2e-05 | -0.1606 | -0.0314 | 0.3322
current_sweep | 4 | 1 | 10 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9737 | 0.051016 | 99.9037 | 0.0061147 | -0.00026838 | 0.0061206 | 119.781 | 0.051 | 0.006146 | -0.19267 | -1.6e-05 | 3.1e-05 | -0.1606 | -0.0314 | 0.5121
current_sweep | 4 | 1 | 11 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9737 | 0.051016 | 99.9037 | 0.0061147 | -0.00026838 | 0.0061206 | 119.781 | 0.051 | 0.006154 | -0.19267 | -1.6e-05 | 3.9e-05 | -0.1606 | -0.0314 | 0.643
current_sweep | 4 | 1 | 12 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9737 | 0.051016 | 99.9037 | 0.0061147 | -0.00026838 | 0.0061206 | 119.781 | 0.051 | 0.006149 | -0.19267 | -1.6e-05 | 3.4e-05 | -0.1606 | -0.0314 | 0.5612
current_sweep | 4 | 1 | 13 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9905 | 0.051001 | 99.8999 | 0.0061135 | -0.00027324 | 0.0061196 | 119.765 | 0.051 | 0.006133 | -0.22549 | -1e-06 | 2e-05 | -0.1879 | -0.0017 | 0.3191
current_sweep | 4 | 1 | 14 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9905 | 0.051001 | 99.8999 | 0.0061135 | -0.00027324 | 0.0061196 | 119.765 | 0.051 | 0.006149 | -0.22549 | -1e-06 | 3.6e-05 | -0.1879 | -0.0017 | 0.5808
current_sweep | 4 | 1 | 15 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9905 | 0.051001 | 99.8999 | 0.0061135 | -0.00027324 | 0.0061196 | 119.765 | 0.051 | 0.006134 | -0.22549 | -1e-06 | 2.1e-05 | -0.1879 | -0.0017 | 0.3354
current_sweep | 4 | 1 | 16 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9905 | 0.051001 | 99.8999 | 0.0061135 | -0.00027324 | 0.0061196 | 119.765 | 0.051 | 0.00615 | -0.22549 | -1e-06 | 3.7e-05 | -0.1879 | -0.0017 | 0.5971
current_sweep | 4 | 1 | 17 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9905 | 0.051001 | 99.8999 | 0.0061135 | -0.00027324 | 0.0061196 | 119.765 | 0.051 | 0.006167 | -0.22549 | -1e-06 | 5.4e-05 | -0.1879 | -0.0017 | 0.8752
current_sweep | 4 | 1 | 18 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9905 | 0.051001 | 99.8999 | 0.0061135 | -0.00027324 | 0.0061196 | 119.765 | 0.051 | 0.00617 | -0.22549 | -1e-06 | 5.7e-05 | -0.1879 | -0.0017 | 0.9243
current_sweep | 4 | 2 | 1 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0123 | 0.050998 | 99.9007 | 0.0061144 | -0.00027241 | 0.0061204 | 119.746 | 0.135 | 0.016049 | -0.26634 | 0.084002 | 0.009935 | -0.2219 | -11.7618 | -12.5064
current_sweep | 4 | 2 | 2 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0123 | 0.050998 | 99.9007 | 0.0061144 | -0.00027241 | 0.0061204 | 119.746 | 0.051 | 0.00616 | -0.26634 | 2e-06 | 4.6e-05 | -0.2219 | 0.0033 | 0.7466
current_sweep | 4 | 2 | 3 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0123 | 0.050998 | 99.9007 | 0.0061144 | -0.00027241 | 0.0061204 | 119.746 | 0.051 | 0.006158 | -0.26634 | 2e-06 | 4.4e-05 | -0.2219 | 0.0033 | 0.7138
current_sweep | 4 | 2 | 4 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0123 | 0.050998 | 99.9007 | 0.0061144 | -0.00027241 | 0.0061204 | 119.746 | 0.051 | 0.006131 | -0.26634 | 2e-06 | 1.7e-05 | -0.2219 | 0.0033 | 0.2723
current_sweep | 4 | 2 | 5 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0123 | 0.050998 | 99.9007 | 0.0061144 | -0.00027241 | 0.0061204 | 119.746 | 0.051 | 0.006146 | -0.26634 | 2e-06 | 3.2e-05 | -0.2219 | 0.0033 | 0.5176
current_sweep | 4 | 2 | 6 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0123 | 0.050998 | 99.9007 | 0.0061144 | -0.00027241 | 0.0061204 | 119.746 | 0.051 | 0.006122 | -0.26634 | 2e-06 | 8e-06 | -0.2219 | 0.0033 | 0.1251
current_sweep | 4 | 2 | 7 | 2 | 120 | 0.051 | 0.999 | 60 | 119.976 | 0.051016 | 99.9043 | 0.0061148 | -0.00026765 | 0.0061207 | 119.748 | 0.143 | 0.017032 | -0.22804 | 0.091984 | 0.010917 | -0.1901 | -6.5647 | -7.1543
current_sweep | 4 | 2 | 8 | 2 | 120 | 0.051 | 0.999 | 60 | 119.976 | 0.051016 | 99.9043 | 0.0061148 | -0.00026765 | 0.0061207 | 119.748 | 0.05 | 0.006037 | -0.22804 | -0.001016 | -7.8e-05 | -0.1901 | -1.991 | -1.2724
current_sweep | 4 | 2 | 9 | 2 | 120 | 0.051 | 0.999 | 60 | 119.976 | 0.051016 | 99.9043 | 0.0061148 | -0.00026765 | 0.0061207 | 119.748 | 0.051 | 0.006153 | -0.22804 | -1.6e-05 | 3.8e-05 | -0.1901 | -0.0308 | 0.6247
current_sweep | 4 | 2 | 10 | 2 | 120 | 0.051 | 0.999 | 60 | 119.976 | 0.051016 | 99.9043 | 0.0061148 | -0.00026765 | 0.0061207 | 119.748 | 0.051 | 0.006144 | -0.22804 | -1.6e-05 | 2.9e-05 | -0.1901 | -0.0308 | 0.4775
current_sweep | 4 | 2 | 11 | 2 | 120 | 0.051 | 0.999 | 60 | 119.976 | 0.051016 | 99.9043 | 0.0061148 | -0.00026765 | 0.0061207 | 119.748 | 0.051 | 0.006153 | -0.22804 | -1.6e-05 | 3.8e-05 | -0.1901 | -0.0308 | 0.6247
current_sweep | 4 | 2 | 12 | 2 | 120 | 0.051 | 0.999 | 60 | 119.976 | 0.051016 | 99.9043 | 0.0061148 | -0.00026765 | 0.0061207 | 119.748 | 0.051 | 0.006148 | -0.22804 | -1.6e-05 | 3.3e-05 | -0.1901 | -0.0308 | 0.5429
current_sweep | 4 | 2 | 13 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9889 | 0.051 | 99.9 | 0.0061134 | -0.00027333 | 0.0061195 | 119.77 | 0.051 | 0.006152 | -0.21885 | 0 | 3.9e-05 | -0.1824 | -0.0009 | 0.6319
current_sweep | 4 | 2 | 14 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9889 | 0.051 | 99.9 | 0.0061134 | -0.00027333 | 0.0061195 | 119.77 | 0.051 | 0.00612 | -0.21885 | 0 | 7e-06 | -0.1824 | -0.0009 | 0.1084
current_sweep | 4 | 2 | 15 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9889 | 0.051 | 99.9 | 0.0061134 | -0.00027333 | 0.0061195 | 119.77 | 0.051 | 0.006124 | -0.21885 | 0 | 1.1e-05 | -0.1824 | -0.0009 | 0.1739
current_sweep | 4 | 2 | 16 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9889 | 0.051 | 99.9 | 0.0061134 | -0.00027333 | 0.0061195 | 119.77 | 0.051 | 0.006149 | -0.21885 | 0 | 3.6e-05 | -0.1824 | -0.0009 | 0.5828
current_sweep | 4 | 2 | 17 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9889 | 0.051 | 99.9 | 0.0061134 | -0.00027333 | 0.0061195 | 119.77 | 0.051 | 0.006167 | -0.21885 | 0 | 5.4e-05 | -0.1824 | -0.0009 | 0.8772
current_sweep | 4 | 2 | 18 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9889 | 0.051 | 99.9 | 0.0061134 | -0.00027333 | 0.0061195 | 119.77 | 0.051 | 0.00617 | -0.21885 | 0 | 5.7e-05 | -0.1824 | -0.0009 | 0.9263
current_sweep | 4 | 3 | 1 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0122 | 0.050998 | 99.9005 | 0.0061143 | -0.00027264 | 0.0061204 | 119.746 | 0.135 | 0.016025 | -0.26615 | 0.084002 | 0.009911 | -0.2218 | -11.7615 | -12.6366
current_sweep | 4 | 3 | 2 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0122 | 0.050998 | 99.9005 | 0.0061143 | -0.00027264 | 0.0061204 | 119.746 | 0.051 | 0.006149 | -0.26615 | 2e-06 | 3.5e-05 | -0.2218 | 0.0036 | 0.5673
current_sweep | 4 | 3 | 3 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0122 | 0.050998 | 99.9005 | 0.0061143 | -0.00027264 | 0.0061204 | 119.746 | 0.051 | 0.006149 | -0.26615 | 2e-06 | 3.5e-05 | -0.2218 | 0.0036 | 0.5673
current_sweep | 4 | 3 | 4 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0122 | 0.050998 | 99.9005 | 0.0061143 | -0.00027264 | 0.0061204 | 119.746 | 0.051 | 0.006131 | -0.26615 | 2e-06 | 1.7e-05 | -0.2218 | 0.0036 | 0.273
current_sweep | 4 | 3 | 5 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0122 | 0.050998 | 99.9005 | 0.0061143 | -0.00027264 | 0.0061204 | 119.746 | 0.051 | 0.006135 | -0.26615 | 2e-06 | 2.1e-05 | -0.2218 | 0.0036 | 0.3384
current_sweep | 4 | 3 | 6 | 1 | 120 | 0.051 | 0.999 | 60 | 120.0122 | 0.050998 | 99.9005 | 0.0061143 | -0.00027264 | 0.0061204 | 119.746 | 0.051 | 0.006123 | -0.26615 | 2e-06 | 9e-06 | -0.2218 | 0.0036 | 0.1421
current_sweep | 4 | 3 | 7 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9743 | 0.051015 | 99.9039 | 0.0061146 | -0.00026824 | 0.0061205 | 119.774 | 0.142 | 0.017036 | -0.20031 | 0.090985 | 0.010921 | -0.167 | -7.2169 | -7.1296
current_sweep | 4 | 3 | 8 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9743 | 0.051015 | 99.9039 | 0.0061146 | -0.00026824 | 0.0061205 | 119.774 | 0.05 | 0.006037 | -0.20031 | -0.001015 | -7.8e-05 | -0.167 | -1.9897 | -1.2693
current_sweep | 4 | 3 | 9 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9743 | 0.051015 | 99.9039 | 0.0061146 | -0.00026824 | 0.0061205 | 119.774 | 0.051 | 0.006154 | -0.20031 | -1.5e-05 | 3.9e-05 | -0.167 | -0.0295 | 0.6442
current_sweep | 4 | 3 | 10 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9743 | 0.051015 | 99.9039 | 0.0061146 | -0.00026824 | 0.0061205 | 119.774 | 0.051 | 0.006155 | -0.20031 | -1.5e-05 | 4e-05 | -0.167 | -0.0295 | 0.6605
current_sweep | 4 | 3 | 11 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9743 | 0.051015 | 99.9039 | 0.0061146 | -0.00026824 | 0.0061205 | 119.774 | 0.051 | 0.006143 | -0.20031 | -1.5e-05 | 2.8e-05 | -0.167 | -0.0295 | 0.4643
current_sweep | 4 | 3 | 12 | 2 | 120 | 0.051 | 0.999 | 60 | 119.9743 | 0.051015 | 99.9039 | 0.0061146 | -0.00026824 | 0.0061205 | 119.774 | 0.051 | 0.006149 | -0.20031 | -1.5e-05 | 3.4e-05 | -0.167 | -0.0295 | 0.5624
current_sweep | 4 | 3 | 13 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9879 | 0.051006 | 99.8986 | 0.0061139 | -0.00027323 | 0.0061201 | 119.764 | 0.051 | 0.006142 | -0.22394 | -6e-06 | 2.8e-05 | -0.1866 | -0.0115 | 0.4598
current_sweep | 4 | 3 | 14 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9879 | 0.051006 | 99.8986 | 0.0061139 | -0.00027323 | 0.0061201 | 119.764 | 0.051 | 0.00614 | -0.22394 | -6e-06 | 2.6e-05 | -0.1866 | -0.0115 | 0.4271
current_sweep | 4 | 3 | 15 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9879 | 0.051006 | 99.8986 | 0.0061139 | -0.00027323 | 0.0061201 | 119.764 | 0.051 | 0.006144 | -0.22394 | -6e-06 | 3e-05 | -0.1866 | -0.0115 | 0.4925
current_sweep | 4 | 3 | 16 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9879 | 0.051006 | 99.8986 | 0.0061139 | -0.00027323 | 0.0061201 | 119.764 | 0.051 | 0.006149 | -0.22394 | -6e-06 | 3.5e-05 | -0.1866 | -0.0115 | 0.5743
current_sweep | 4 | 3 | 17 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9879 | 0.051006 | 99.8986 | 0.0061139 | -0.00027323 | 0.0061201 | 119.764 | 0.051 | 0.006167 | -0.22394 | -6e-06 | 5.3e-05 | -0.1866 | -0.0115 | 0.8687
current_sweep | 4 | 3 | 18 | 3 | 120 | 0.051 | 0.999 | 60 | 119.9879 | 0.051006 | 99.8986 | 0.0061139 | -0.00027323 | 0.0061201 | 119.764 | 0.051 | 0.006151 | -0.22394 | -6e-06 | 3.7e-05 | -0.1866 | -0.0115 | 0.607
current_sweep | 5 | 1 | 1 | 1 | 120 | 0.0877 | 0.999 | 60 | 119.9972 | 0.087697 | 99.9001 | 0.010513 | -0.00046943 | 0.010523 | 119.736 | 0.239 | 0.02873 | -0.26124 | 0.1513 | 0.018217 | -0.2177 | -9.1574 | -8.9059
current_sweep | 5 | 1 | 2 | 1 | 120 | 0.0877 | 0.999 | 60 | 119.9972 | 0.087697 | 99.9001 | 0.010513 | -0.00046943 | 0.010523 | 119.736 | 0.088 | 0.010563 | -0.26124 | 0.000303 | 5e-05 | -0.2177 | 0.345 | 0.4762
current_sweep | 5 | 1 | 3 | 1 | 120 | 0.0877 | 0.999 | 60 | 119.9972 | 0.087697 | 99.9001 | 0.010513 | -0.00046943 | 0.010523 | 119.736 | 0.088 | 0.01056 | -0.26124 | 0.000303 | 4.7e-05 | -0.2177 | 0.345 | 0.4477
current_sweep | 5 | 1 | 4 | 1 | 120 | 0.0877 | 0.999 | 60 | 119.9972 | 0.087697 | 99.9001 | 0.010513 | -0.00046943 | 0.010523 | 119.736 | 0.088 | 0.01055 | -0.26124 | 0.000303 | 3.7e-05 | -0.2177 | 0.345 | 0.3526
current_sweep | 5 | 1 | 5 | 1 | 120 | 0.0877 | 0.999 | 60 | 119.9972 | 0.087697 | 99.9001 | 0.010513 | -0.00046943 | 0.010523 | 119.736 | 0.088 | 0.010558 | -0.26124 | 0.000303 | 4.5e-05 | -0.2177 | 0.345 | 0.4287
current_sweep | 5 | 1 | 6 | 1 | 120 | 0.0877 | 0.999 | 60 | 119.9972 | 0.087697 | 99.9001 | 0.010513 | -0.00046943 | 0.010523 | 119.736 | 0.087 | 0.010534 | -0.26124 | -0.000697 | 2.1e-05 | -0.2177 | -0.7953 | 0.2004
current_sweep | 5 | 1 | 7 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.98 | 0.087723 | 99.9037 | 0.010515 | -0.0004618 | 0.010525 | 119.778 | 0.248 | 0.02978 | -0.20196 | 0.16028 | 0.019265 | -0.1683 | -5.7641 | -5.594
current_sweep | 5 | 1 | 8 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.98 | 0.087723 | 99.9037 | 0.010515 | -0.0004618 | 0.010525 | 119.778 | 0.087 | 0.010456 | -0.20196 | -0.000723 | -5.9e-05 | -0.1683 | -0.8243 | -0.5599
current_sweep | 5 | 1 | 9 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.98 | 0.087723 | 99.9037 | 0.010515 | -0.0004618 | 0.010525 | 119.778 | 0.088 | 0.010553 | -0.20196 | 0.000277 | 3.8e-05 | -0.1683 | 0.3157 | 0.3626
current_sweep | 5 | 1 | 10 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.98 | 0.087723 | 99.9037 | 0.010515 | -0.0004618 | 0.010525 | 119.778 | 0.088 | 0.010571 | -0.20196 | 0.000277 | 5.6e-05 | -0.1683 | 0.3157 | 0.5338
current_sweep | 5 | 1 | 11 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.98 | 0.087723 | 99.9037 | 0.010515 | -0.0004618 | 0.010525 | 119.778 | 0.088 | 0.010561 | -0.20196 | 0.000277 | 4.6e-05 | -0.1683 | 0.3157 | 0.4387
current_sweep | 5 | 1 | 12 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.98 | 0.087723 | 99.9037 | 0.010515 | -0.0004618 | 0.010525 | 119.778 | 0.088 | 0.010567 | -0.20196 | 0.000277 | 5.2e-05 | -0.1683 | 0.3157 | 0.4957
current_sweep | 5 | 1 | 13 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9902 | 0.087705 | 99.9008 | 0.010513 | -0.00046843 | 0.010524 | 119.772 | 0.088 | 0.010554 | -0.2182 | 0.000295 | 4.1e-05 | -0.1818 | 0.3364 | 0.3872
current_sweep | 5 | 1 | 14 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9902 | 0.087705 | 99.9008 | 0.010513 | -0.00046843 | 0.010524 | 119.772 | 0.088 | 0.010549 | -0.2182 | 0.000295 | 3.6e-05 | -0.1818 | 0.3364 | 0.3396
current_sweep | 5 | 1 | 15 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9902 | 0.087705 | 99.9008 | 0.010513 | -0.00046843 | 0.010524 | 119.772 | 0.088 | 0.010557 | -0.2182 | 0.000295 | 4.4e-05 | -0.1818 | 0.3364 | 0.4157
current_sweep | 5 | 1 | 16 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9902 | 0.087705 | 99.9008 | 0.010513 | -0.00046843 | 0.010524 | 119.772 | 0.088 | 0.010618 | -0.2182 | 0.000295 | 0.000105 | -0.1818 | 0.3364 | 0.9959
current_sweep | 5 | 1 | 17 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9902 | 0.087705 | 99.9008 | 0.010513 | -0.00046843 | 0.010524 | 119.772 | 0.088 | 0.010618 | -0.2182 | 0.000295 | 0.000105 | -0.1818 | 0.3364 | 0.9959
current_sweep | 5 | 1 | 18 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9902 | 0.087705 | 99.9008 | 0.010513 | -0.00046843 | 0.010524 | 119.772 | 0.088 | 0.010617 | -0.2182 | 0.000295 | 0.000104 | -0.1818 | 0.3364 | 0.9864
current_sweep | 5 | 2 | 1 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.01 | 0.087699 | 99.9006 | 0.010514 | -0.00046905 | 0.010525 | 119.728 | 0.24 | 0.028758 | -0.28205 | 0.1523 | 0.018244 | -0.235 | -8.7786 | -8.8286
current_sweep | 5 | 2 | 2 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.01 | 0.087699 | 99.9006 | 0.010514 | -0.00046905 | 0.010525 | 119.728 | 0.088 | 0.01057 | -0.28205 | 0.000301 | 5.6e-05 | -0.235 | 0.3435 | 0.5301
current_sweep | 5 | 2 | 3 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.01 | 0.087699 | 99.9006 | 0.010514 | -0.00046905 | 0.010525 | 119.728 | 0.088 | 0.010563 | -0.28205 | 0.000301 | 4.9e-05 | -0.235 | 0.3435 | 0.4635
current_sweep | 5 | 2 | 4 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.01 | 0.087699 | 99.9006 | 0.010514 | -0.00046905 | 0.010525 | 119.728 | 0.088 | 0.010548 | -0.28205 | 0.000301 | 3.4e-05 | -0.235 | 0.3435 | 0.3208
current_sweep | 5 | 2 | 5 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.01 | 0.087699 | 99.9006 | 0.010514 | -0.00046905 | 0.010525 | 119.728 | 0.088 | 0.010562 | -0.28205 | 0.000301 | 4.8e-05 | -0.235 | 0.3435 | 0.454
current_sweep | 5 | 2 | 6 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.01 | 0.087699 | 99.9006 | 0.010514 | -0.00046905 | 0.010525 | 119.728 | 0.088 | 0.010538 | -0.28205 | 0.000301 | 2.4e-05 | -0.235 | 0.3435 | 0.2257
current_sweep | 5 | 2 | 7 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.983 | 0.087724 | 99.9036 | 0.010515 | -0.00046207 | 0.010525 | 119.754 | 0.248 | 0.029764 | -0.22902 | 0.16028 | 0.019249 | -0.1909 | -5.7653 | -5.6483
current_sweep | 5 | 2 | 8 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.983 | 0.087724 | 99.9036 | 0.010515 | -0.00046207 | 0.010525 | 119.754 | 0.087 | 0.010455 | -0.22902 | -0.000724 | -6e-05 | -0.1909 | -0.8255 | -0.5731
current_sweep | 5 | 2 | 9 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.983 | 0.087724 | 99.9036 | 0.010515 | -0.00046207 | 0.010525 | 119.754 | 0.088 | 0.010559 | -0.22902 | 0.000276 | 4.4e-05 | -0.1909 | 0.3144 | 0.4159
current_sweep | 5 | 2 | 10 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.983 | 0.087724 | 99.9036 | 0.010515 | -0.00046207 | 0.010525 | 119.754 | 0.088 | 0.010578 | -0.22902 | 0.000276 | 6.3e-05 | -0.1909 | 0.3144 | 0.5966
current_sweep | 5 | 2 | 11 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.983 | 0.087724 | 99.9036 | 0.010515 | -0.00046207 | 0.010525 | 119.754 | 0.088 | 0.010561 | -0.22902 | 0.000276 | 4.6e-05 | -0.1909 | 0.3144 | 0.435
current_sweep | 5 | 2 | 12 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.983 | 0.087724 | 99.9036 | 0.010515 | -0.00046207 | 0.010525 | 119.754 | 0.088 | 0.010555 | -0.22902 | 0.000276 | 4e-05 | -0.1909 | 0.3144 | 0.3779
current_sweep | 5 | 2 | 13 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.991 | 0.087705 | 99.9008 | 0.010513 | -0.00046841 | 0.010524 | 119.78 | 0.088 | 0.010545 | -0.21098 | 0.000295 | 3.2e-05 | -0.1758 | 0.3367 | 0.3011
current_sweep | 5 | 2 | 14 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.991 | 0.087705 | 99.9008 | 0.010513 | -0.00046841 | 0.010524 | 119.78 | 0.088 | 0.010544 | -0.21098 | 0.000295 | 3.1e-05 | -0.1758 | 0.3367 | 0.2916
current_sweep | 5 | 2 | 15 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.991 | 0.087705 | 99.9008 | 0.010513 | -0.00046841 | 0.010524 | 119.78 | 0.088 | 0.010554 | -0.21098 | 0.000295 | 4.1e-05 | -0.1758 | 0.3367 | 0.3867
current_sweep | 5 | 2 | 16 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.991 | 0.087705 | 99.9008 | 0.010513 | -0.00046841 | 0.010524 | 119.78 | 0.088 | 0.010637 | -0.21098 | 0.000295 | 0.000124 | -0.1758 | 0.3367 | 1.1762
current_sweep | 5 | 2 | 17 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.991 | 0.087705 | 99.9008 | 0.010513 | -0.00046841 | 0.010524 | 119.78 | 0.088 | 0.010625 | -0.21098 | 0.000295 | 0.000112 | -0.1758 | 0.3367 | 1.0621
current_sweep | 5 | 2 | 18 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.991 | 0.087705 | 99.9008 | 0.010513 | -0.00046841 | 0.010524 | 119.78 | 0.088 | 0.010608 | -0.21098 | 0.000295 | 9.5e-05 | -0.1758 | 0.3367 | 0.9004
current_sweep | 5 | 3 | 1 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.0128 | 0.087698 | 99.9004 | 0.010514 | -0.00046886 | 0.010525 | 119.75 | 0.239 | 0.028736 | -0.26277 | 0.1513 | 0.018222 | -0.219 | -9.1582 | -8.8997
current_sweep | 5 | 3 | 2 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.0128 | 0.087698 | 99.9004 | 0.010514 | -0.00046886 | 0.010525 | 119.75 | 0.088 | 0.010561 | -0.26277 | 0.000302 | 4.7e-05 | -0.219 | 0.3441 | 0.443
current_sweep | 5 | 3 | 3 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.0128 | 0.087698 | 99.9004 | 0.010514 | -0.00046886 | 0.010525 | 119.75 | 0.088 | 0.010575 | -0.26277 | 0.000302 | 6.1e-05 | -0.219 | 0.3441 | 0.5762
current_sweep | 5 | 3 | 4 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.0128 | 0.087698 | 99.9004 | 0.010514 | -0.00046886 | 0.010525 | 119.75 | 0.088 | 0.010571 | -0.26277 | 0.000302 | 5.7e-05 | -0.219 | 0.3441 | 0.5381
current_sweep | 5 | 3 | 5 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.0128 | 0.087698 | 99.9004 | 0.010514 | -0.00046886 | 0.010525 | 119.75 | 0.088 | 0.010563 | -0.26277 | 0.000302 | 4.9e-05 | -0.219 | 0.3441 | 0.462
current_sweep | 5 | 3 | 6 | 1 | 120 | 0.0877 | 0.999 | 60 | 120.0128 | 0.087698 | 99.9004 | 0.010514 | -0.00046886 | 0.010525 | 119.75 | 0.088 | 0.01054 | -0.26277 | 0.000302 | 2.6e-05 | -0.219 | 0.3441 | 0.2433
current_sweep | 5 | 3 | 7 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.9818 | 0.087725 | 99.9037 | 0.010515 | -0.00046175 | 0.010525 | 119.773 | 0.248 | 0.029797 | -0.20875 | 0.16028 | 0.019282 | -0.174 | -5.766 | -5.5435
current_sweep | 5 | 3 | 8 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.9818 | 0.087725 | 99.9037 | 0.010515 | -0.00046175 | 0.010525 | 119.773 | 0.087 | 0.010463 | -0.20875 | -0.000725 | -5.2e-05 | -0.174 | -0.8263 | -0.4969
current_sweep | 5 | 3 | 9 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.9818 | 0.087725 | 99.9037 | 0.010515 | -0.00046175 | 0.010525 | 119.773 | 0.087 | 0.010524 | -0.20875 | -0.000725 | 9e-06 | -0.174 | -0.8263 | 0.0833
current_sweep | 5 | 3 | 10 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.9818 | 0.087725 | 99.9037 | 0.010515 | -0.00046175 | 0.010525 | 119.773 | 0.088 | 0.010571 | -0.20875 | 0.000275 | 5.6e-05 | -0.174 | 0.3136 | 0.5302
current_sweep | 5 | 3 | 11 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.9818 | 0.087725 | 99.9037 | 0.010515 | -0.00046175 | 0.010525 | 119.773 | 0.088 | 0.010551 | -0.20875 | 0.000275 | 3.6e-05 | -0.174 | 0.3136 | 0.34
current_sweep | 5 | 3 | 12 | 2 | 120 | 0.0877 | 0.999 | 60 | 119.9818 | 0.087725 | 99.9037 | 0.010515 | -0.00046175 | 0.010525 | 119.773 | 0.088 | 0.010573 | -0.20875 | 0.000275 | 5.8e-05 | -0.174 | 0.3136 | 0.5492
current_sweep | 5 | 3 | 13 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9895 | 0.087704 | 99.9007 | 0.010513 | -0.00046879 | 0.010524 | 119.785 | 0.088 | 0.010551 | -0.20453 | 0.000296 | 3.8e-05 | -0.1705 | 0.3371 | 0.36
current_sweep | 5 | 3 | 14 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9895 | 0.087704 | 99.9007 | 0.010513 | -0.00046879 | 0.010524 | 119.785 | 0.088 | 0.010553 | -0.20453 | 0.000296 | 4e-05 | -0.1705 | 0.3371 | 0.379
current_sweep | 5 | 3 | 15 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9895 | 0.087704 | 99.9007 | 0.010513 | -0.00046879 | 0.010524 | 119.785 | 0.088 | 0.010563 | -0.20453 | 0.000296 | 5e-05 | -0.1705 | 0.3371 | 0.4741
current_sweep | 5 | 3 | 16 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9895 | 0.087704 | 99.9007 | 0.010513 | -0.00046879 | 0.010524 | 119.785 | 0.088 | 0.010637 | -0.20453 | 0.000296 | 0.000124 | -0.1705 | 0.3371 | 1.178
current_sweep | 5 | 3 | 17 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9895 | 0.087704 | 99.9007 | 0.010513 | -0.00046879 | 0.010524 | 119.785 | 0.088 | 0.010617 | -0.20453 | 0.000296 | 0.000104 | -0.1705 | 0.3371 | 0.9877
current_sweep | 5 | 3 | 18 | 3 | 120 | 0.0877 | 0.999 | 60 | 119.9895 | 0.087704 | 99.9007 | 0.010513 | -0.00046879 | 0.010524 | 119.785 | 0.088 | 0.010613 | -0.20453 | 0.000296 | 0.0001 | -0.1705 | 0.3371 | 0.9497
current_sweep | 6 | 1 | 1 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0051 | 0.151 | 99.9011 | 0.018103 | -0.00080442 | 0.018121 | 119.732 | 0.426 | 0.051085 | -0.27307 | 0.275 | 0.032982 | -0.2275 | -5.9611 | -5.9364
current_sweep | 6 | 1 | 2 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0051 | 0.151 | 99.9011 | 0.018103 | -0.00080442 | 0.018121 | 119.732 | 0.151 | 0.018168 | -0.27307 | -1e-06 | 6.5e-05 | -0.2275 | -0.0009 | 0.359
current_sweep | 6 | 1 | 3 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0051 | 0.151 | 99.9011 | 0.018103 | -0.00080442 | 0.018121 | 119.732 | 0.151 | 0.018173 | -0.27307 | -1e-06 | 7e-05 | -0.2275 | -0.0009 | 0.3867
current_sweep | 6 | 1 | 4 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0051 | 0.151 | 99.9011 | 0.018103 | -0.00080442 | 0.018121 | 119.732 | 0.151 | 0.018118 | -0.27307 | -1e-06 | 1.5e-05 | -0.2275 | -0.0009 | 0.0828
current_sweep | 6 | 1 | 5 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0051 | 0.151 | 99.9011 | 0.018103 | -0.00080442 | 0.018121 | 119.732 | 0.151 | 0.018142 | -0.27307 | -1e-06 | 3.9e-05 | -0.2275 | -0.0009 | 0.2154
current_sweep | 6 | 1 | 6 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0051 | 0.151 | 99.9011 | 0.018103 | -0.00080442 | 0.018121 | 119.732 | 0.15 | 0.018078 | -0.27307 | -0.001001 | -2.5e-05 | -0.2275 | -0.6631 | -0.1381
current_sweep | 6 | 1 | 7 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079667 | 0.018119 | 119.77 | 0.436 | 0.052239 | -0.20804 | 0.28498 | 0.034137 | -0.1734 | -3.766 | -3.8043
current_sweep | 6 | 1 | 8 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079667 | 0.018119 | 119.77 | 0.15 | 0.018042 | -0.20804 | -0.001021 | -6e-05 | -0.1734 | -0.6759 | -0.3295
current_sweep | 6 | 1 | 9 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079667 | 0.018119 | 119.77 | 0.151 | 0.018154 | -0.20804 | -2.1e-05 | 5.2e-05 | -0.1734 | -0.0137 | 0.2893
current_sweep | 6 | 1 | 10 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079667 | 0.018119 | 119.77 | 0.151 | 0.01815 | -0.20804 | -2.1e-05 | 4.8e-05 | -0.1734 | -0.0137 | 0.2672
current_sweep | 6 | 1 | 11 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079667 | 0.018119 | 119.77 | 0.151 | 0.018103 | -0.20804 | -2.1e-05 | 1e-06 | -0.1734 | -0.0137 | 0.0075
current_sweep | 6 | 1 | 12 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079667 | 0.018119 | 119.77 | 0.151 | 0.018121 | -0.20804 | -2.1e-05 | 1.9e-05 | -0.1734 | -0.0137 | 0.107
current_sweep | 6 | 1 | 13 | 3 | 120 | 0.151 | 0.999 | 60 | 119.993 | 0.15102 | 99.9007 | 0.018104 | -0.00080709 | 0.018122 | 119.766 | 0.151 | 0.018142 | -0.227 | -2.3e-05 | 3.8e-05 | -0.1892 | -0.0153 | 0.2114
current_sweep | 6 | 1 | 14 | 3 | 120 | 0.151 | 0.999 | 60 | 119.993 | 0.15102 | 99.9007 | 0.018104 | -0.00080709 | 0.018122 | 119.766 | 0.151 | 0.018125 | -0.227 | -2.3e-05 | 2.1e-05 | -0.1892 | -0.0153 | 0.1175
current_sweep | 6 | 1 | 15 | 3 | 120 | 0.151 | 0.999 | 60 | 119.993 | 0.15102 | 99.9007 | 0.018104 | -0.00080709 | 0.018122 | 119.766 | 0.151 | 0.018157 | -0.227 | -2.3e-05 | 5.3e-05 | -0.1892 | -0.0153 | 0.2943
current_sweep | 6 | 1 | 16 | 3 | 120 | 0.151 | 0.999 | 60 | 119.993 | 0.15102 | 99.9007 | 0.018104 | -0.00080709 | 0.018122 | 119.766 | 0.15 | 0.018076 | -0.227 | -0.001023 | -2.8e-05 | -0.1892 | -0.6774 | -0.1531
current_sweep | 6 | 1 | 17 | 3 | 120 | 0.151 | 0.999 | 60 | 119.993 | 0.15102 | 99.9007 | 0.018104 | -0.00080709 | 0.018122 | 119.766 | 0.15 | 0.018083 | -0.227 | -0.001023 | -2.1e-05 | -0.1892 | -0.6774 | -0.1145
current_sweep | 6 | 1 | 18 | 3 | 120 | 0.151 | 0.999 | 60 | 119.993 | 0.15102 | 99.9007 | 0.018104 | -0.00080709 | 0.018122 | 119.766 | 0.151 | 0.018085 | -0.227 | -2.3e-05 | -1.9e-05 | -0.1892 | -0.0153 | -0.1034
current_sweep | 6 | 2 | 1 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0102 | 0.151 | 99.9015 | 0.018104 | -0.00080324 | 0.018121 | 119.775 | 0.427 | 0.051127 | -0.23516 | 0.276 | 0.033023 | -0.196 | -5.739 | -5.8622
current_sweep | 6 | 2 | 2 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0102 | 0.151 | 99.9015 | 0.018104 | -0.00080324 | 0.018121 | 119.775 | 0.151 | 0.018168 | -0.23516 | 1e-06 | 6.4e-05 | -0.196 | 0.0005 | 0.3558
current_sweep | 6 | 2 | 3 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0102 | 0.151 | 99.9015 | 0.018104 | -0.00080324 | 0.018121 | 119.775 | 0.151 | 0.018162 | -0.23516 | 1e-06 | 5.8e-05 | -0.196 | 0.0005 | 0.3226
current_sweep | 6 | 2 | 4 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0102 | 0.151 | 99.9015 | 0.018104 | -0.00080324 | 0.018121 | 119.775 | 0.151 | 0.018095 | -0.23516 | 1e-06 | -9e-06 | -0.196 | 0.0005 | -0.0475
current_sweep | 6 | 2 | 5 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0102 | 0.151 | 99.9015 | 0.018104 | -0.00080324 | 0.018121 | 119.775 | 0.151 | 0.018144 | -0.23516 | 1e-06 | 4e-05 | -0.196 | 0.0005 | 0.2232
current_sweep | 6 | 2 | 6 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0102 | 0.151 | 99.9015 | 0.018104 | -0.00080324 | 0.018121 | 119.775 | 0.151 | 0.018105 | -0.23516 | 1e-06 | 1e-06 | -0.196 | 0.0005 | 0.0078
current_sweep | 6 | 2 | 7 | 2 | 120 | 0.151 | 0.999 | 60 | 119.9805 | 0.15102 | 99.9031 | 0.018102 | -0.00079712 | 0.01812 | 119.775 | 0.436 | 0.052272 | -0.20549 | 0.28498 | 0.03417 | -0.1713 | -3.766 | -3.7454
current_sweep | 6 | 2 | 8 | 2 | 120 | 0.151 | 0.999 | 60 | 119.9805 | 0.15102 | 99.9031 | 0.018102 | -0.00079712 | 0.01812 | 119.775 | 0.15 | 0.018046 | -0.20549 | -0.001021 | -5.6e-05 | -0.1713 | -0.6759 | -0.3093
current_sweep | 6 | 2 | 9 | 2 | 120 | 0.151 | 0.999 | 60 | 119.9805 | 0.15102 | 99.9031 | 0.018102 | -0.00079712 | 0.01812 | 119.775 | 0.151 | 0.018153 | -0.20549 | -2.1e-05 | 5.1e-05 | -0.1713 | -0.0137 | 0.2818
current_sweep | 6 | 2 | 10 | 2 | 120 | 0.151 | 0.999 | 60 | 119.9805 | 0.15102 | 99.9031 | 0.018102 | -0.00079712 | 0.01812 | 119.775 | 0.151 | 0.018149 | -0.20549 | -2.1e-05 | 4.7e-05 | -0.1713 | -0.0137 | 0.2597
current_sweep | 6 | 2 | 11 | 2 | 120 | 0.151 | 0.999 | 60 | 119.9805 | 0.15102 | 99.9031 | 0.018102 | -0.00079712 | 0.01812 | 119.775 | 0.151 | 0.018092 | -0.20549 | -2.1e-05 | -1e-05 | -0.1713 | -0.0137 | -0.0552
current_sweep | 6 | 2 | 12 | 2 | 120 | 0.151 | 0.999 | 60 | 119.9805 | 0.15102 | 99.9031 | 0.018102 | -0.00079712 | 0.01812 | 119.775 | 0.151 | 0.018133 | -0.20549 | -2.1e-05 | 3.1e-05 | -0.1713 | -0.0137 | 0.1713
current_sweep | 6 | 2 | 13 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9925 | 0.15102 | 99.9005 | 0.018103 | -0.00080768 | 0.018121 | 119.766 | 0.151 | 0.018141 | -0.22646 | -1.6e-05 | 3.8e-05 | -0.1887 | -0.0108 | 0.211
current_sweep | 6 | 2 | 14 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9925 | 0.15102 | 99.9005 | 0.018103 | -0.00080768 | 0.018121 | 119.766 | 0.151 | 0.018126 | -0.22646 | -1.6e-05 | 2.3e-05 | -0.1887 | -0.0108 | 0.1282
current_sweep | 6 | 2 | 15 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9925 | 0.15102 | 99.9005 | 0.018103 | -0.00080768 | 0.018121 | 119.766 | 0.151 | 0.018156 | -0.22646 | -1.6e-05 | 5.3e-05 | -0.1887 | -0.0108 | 0.2939
current_sweep | 6 | 2 | 16 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9925 | 0.15102 | 99.9005 | 0.018103 | -0.00080768 | 0.018121 | 119.766 | 0.15 | 0.018077 | -0.22646 | -0.001016 | -2.6e-05 | -0.1887 | -0.673 | -0.1425
current_sweep | 6 | 2 | 17 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9925 | 0.15102 | 99.9005 | 0.018103 | -0.00080768 | 0.018121 | 119.766 | 0.15 | 0.018061 | -0.22646 | -0.001016 | -4.2e-05 | -0.1887 | -0.673 | -0.2309
current_sweep | 6 | 2 | 18 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9925 | 0.15102 | 99.9005 | 0.018103 | -0.00080768 | 0.018121 | 119.766 | 0.15 | 0.018069 | -0.22646 | -0.001016 | -3.4e-05 | -0.1887 | -0.673 | -0.1867
current_sweep | 6 | 3 | 1 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0081 | 0.15099 | 99.9016 | 0.018102 | -0.00080284 | 0.018119 | 119.737 | 0.426 | 0.051087 | -0.27108 | 0.27502 | 0.032985 | -0.2259 | -5.9512 | -5.9257
current_sweep | 6 | 3 | 2 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0081 | 0.15099 | 99.9016 | 0.018102 | -0.00080284 | 0.018119 | 119.737 | 0.151 | 0.018166 | -0.27108 | 1.5e-05 | 6.4e-05 | -0.2259 | 0.0096 | 0.3555
current_sweep | 6 | 3 | 3 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0081 | 0.15099 | 99.9016 | 0.018102 | -0.00080284 | 0.018119 | 119.737 | 0.151 | 0.018161 | -0.27108 | 1.5e-05 | 5.9e-05 | -0.2259 | 0.0096 | 0.3279
current_sweep | 6 | 3 | 4 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0081 | 0.15099 | 99.9016 | 0.018102 | -0.00080284 | 0.018119 | 119.737 | 0.151 | 0.018116 | -0.27108 | 1.5e-05 | 1.4e-05 | -0.2259 | 0.0096 | 0.0793
current_sweep | 6 | 3 | 5 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0081 | 0.15099 | 99.9016 | 0.018102 | -0.00080284 | 0.018119 | 119.737 | 0.151 | 0.018142 | -0.27108 | 1.5e-05 | 4e-05 | -0.2259 | 0.0096 | 0.223
current_sweep | 6 | 3 | 6 | 1 | 120 | 0.151 | 0.999 | 60 | 120.0081 | 0.15099 | 99.9016 | 0.018102 | -0.00080284 | 0.018119 | 119.737 | 0.151 | 0.018098 | -0.27108 | 1.5e-05 | -4e-06 | -0.2259 | 0.0096 | -0.0201
current_sweep | 6 | 3 | 7 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079645 | 0.018119 | 119.773 | 0.436 | 0.052267 | -0.20503 | 0.28498 | 0.034165 | -0.1709 | -3.7652 | -3.7521
current_sweep | 6 | 3 | 8 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079645 | 0.018119 | 119.773 | 0.15 | 0.018061 | -0.20503 | -0.00102 | -4.1e-05 | -0.1709 | -0.6751 | -0.2238
current_sweep | 6 | 3 | 9 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079645 | 0.018119 | 119.773 | 0.151 | 0.018154 | -0.20503 | -2e-05 | 5.2e-05 | -0.1709 | -0.013 | 0.29
current_sweep | 6 | 3 | 10 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079645 | 0.018119 | 119.773 | 0.151 | 0.018144 | -0.20503 | -2e-05 | 4.2e-05 | -0.1709 | -0.013 | 0.2347
current_sweep | 6 | 3 | 11 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079645 | 0.018119 | 119.773 | 0.151 | 0.018087 | -0.20503 | -2e-05 | -1.5e-05 | -0.1709 | -0.013 | -0.0802
current_sweep | 6 | 3 | 12 | 2 | 120 | 0.151 | 0.999 | 60 | 119.978 | 0.15102 | 99.9033 | 0.018102 | -0.00079645 | 0.018119 | 119.773 | 0.151 | 0.018128 | -0.20503 | -2e-05 | 2.6e-05 | -0.1709 | -0.013 | 0.1463
current_sweep | 6 | 3 | 13 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9893 | 0.15101 | 99.9008 | 0.018101 | -0.0008068 | 0.018119 | 119.768 | 0.151 | 0.018137 | -0.22133 | -8e-06 | 3.6e-05 | -0.1845 | -0.0053 | 0.1968
current_sweep | 6 | 3 | 14 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9893 | 0.15101 | 99.9008 | 0.018101 | -0.0008068 | 0.018119 | 119.768 | 0.151 | 0.018136 | -0.22133 | -8e-06 | 3.5e-05 | -0.1845 | -0.0053 | 0.1913
current_sweep | 6 | 3 | 15 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9893 | 0.15101 | 99.9008 | 0.018101 | -0.0008068 | 0.018119 | 119.768 | 0.151 | 0.018141 | -0.22133 | -8e-06 | 4e-05 | -0.1845 | -0.0053 | 0.2189
current_sweep | 6 | 3 | 16 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9893 | 0.15101 | 99.9008 | 0.018101 | -0.0008068 | 0.018119 | 119.768 | 0.151 | 0.018086 | -0.22133 | -8e-06 | -1.5e-05 | -0.1845 | -0.0053 | -0.0849
current_sweep | 6 | 3 | 17 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9893 | 0.15101 | 99.9008 | 0.018101 | -0.0008068 | 0.018119 | 119.768 | 0.15 | 0.018068 | -0.22133 | -0.001008 | -3.3e-05 | -0.1845 | -0.6675 | -0.1844
current_sweep | 6 | 3 | 18 | 3 | 120 | 0.151 | 0.999 | 60 | 119.9893 | 0.15101 | 99.9008 | 0.018101 | -0.0008068 | 0.018119 | 119.768 | 0.151 | 0.018088 | -0.22133 | -8e-06 | -1.3e-05 | -0.1845 | -0.0053 | -0.0739
current_sweep | 7 | 1 | 1 | 1 | 120 | 0.2599 | 0.999 | 60 | 119.9937 | 0.25989 | 99.9009 | 0.031154 | -0.0013849 | 0.031185 | 119.769 | 0.748 | 0.08958 | -0.22466 | 0.48811 | 0.058426 | -0.1872 | -4.0611 | -4.1535
current_sweep | 7 | 1 | 2 | 1 | 120 | 0.2599 | 0.999 | 60 | 119.9937 | 0.25989 | 99.9009 | 0.031154 | -0.0013849 | 0.031185 | 119.769 | 0.26 | 0.03117 | -0.22466 | 0.000112 | 1.6e-05 | -0.1872 | 0.0432 | 0.0515
current_sweep | 7 | 1 | 3 | 1 | 120 | 0.2599 | 0.999 | 60 | 119.9937 | 0.25989 | 99.9009 | 0.031154 | -0.0013849 | 0.031185 | 119.769 | 0.26 | 0.031205 | -0.22466 | 0.000112 | 5.1e-05 | -0.1872 | 0.0432 | 0.1638
current_sweep | 7 | 1 | 4 | 1 | 120 | 0.2599 | 0.999 | 60 | 119.9937 | 0.25989 | 99.9009 | 0.031154 | -0.0013849 | 0.031185 | 119.769 | 0.259 | 0.031117 | -0.22466 | -0.000888 | -3.7e-05 | -0.1872 | -0.3416 | -0.1187
current_sweep | 7 | 1 | 5 | 1 | 120 | 0.2599 | 0.999 | 60 | 119.9937 | 0.25989 | 99.9009 | 0.031154 | -0.0013849 | 0.031185 | 119.769 | 0.259 | 0.031129 | -0.22466 | -0.000888 | -2.5e-05 | -0.1872 | -0.3416 | -0.0801
current_sweep | 7 | 1 | 6 | 1 | 120 | 0.2599 | 0.999 | 60 | 119.9937 | 0.25989 | 99.9009 | 0.031154 | -0.0013849 | 0.031185 | 119.769 | 0.259 | 0.031067 | -0.22466 | -0.000888 | -8.7e-05 | -0.1872 | -0.3416 | -0.2791
current_sweep | 7 | 1 | 7 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9806 | 0.25994 | 99.9031 | 0.031158 | -0.0013728 | 0.031188 | 119.769 | 0.757 | 0.090767 | -0.21163 | 0.49706 | 0.059609 | -0.1764 | -2.9263 | -2.8945
current_sweep | 7 | 1 | 8 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9806 | 0.25994 | 99.9031 | 0.031158 | -0.0013728 | 0.031188 | 119.769 | 0.258 | 0.03097 | -0.21163 | -0.00194 | -0.000188 | -0.1764 | -0.7463 | -0.6019
current_sweep | 7 | 1 | 9 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9806 | 0.25994 | 99.9031 | 0.031158 | -0.0013728 | 0.031188 | 119.769 | 0.259 | 0.031123 | -0.21163 | -0.00094 | -3.5e-05 | -0.1764 | -0.3616 | -0.1108
current_sweep | 7 | 1 | 10 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9806 | 0.25994 | 99.9031 | 0.031158 | -0.0013728 | 0.031188 | 119.769 | 0.259 | 0.031124 | -0.21163 | -0.00094 | -3.4e-05 | -0.1764 | -0.3616 | -0.1076
current_sweep | 7 | 1 | 11 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9806 | 0.25994 | 99.9031 | 0.031158 | -0.0013728 | 0.031188 | 119.769 | 0.259 | 0.031091 | -0.21163 | -0.00094 | -6.7e-05 | -0.1764 | -0.3616 | -0.2135
current_sweep | 7 | 1 | 12 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9806 | 0.25994 | 99.9031 | 0.031158 | -0.0013728 | 0.031188 | 119.769 | 0.259 | 0.031111 | -0.21163 | -0.00094 | -4.7e-05 | -0.1764 | -0.3616 | -0.1493
current_sweep | 7 | 1 | 13 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9954 | 0.25991 | 99.9006 | 0.031157 | -0.0013897 | 0.031188 | 119.781 | 0.259 | 0.03113 | -0.21437 | -0.000907 | -2.7e-05 | -0.1786 | -0.349 | -0.0855
current_sweep | 7 | 1 | 14 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9954 | 0.25991 | 99.9006 | 0.031157 | -0.0013897 | 0.031188 | 119.781 | 0.259 | 0.031122 | -0.21437 | -0.000907 | -3.5e-05 | -0.1786 | -0.349 | -0.1112
current_sweep | 7 | 1 | 15 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9954 | 0.25991 | 99.9006 | 0.031157 | -0.0013897 | 0.031188 | 119.781 | 0.26 | 0.031149 | -0.21437 | 9.3e-05 | -8e-06 | -0.1786 | 0.0358 | -0.0245
current_sweep | 7 | 1 | 16 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9954 | 0.25991 | 99.9006 | 0.031157 | -0.0013897 | 0.031188 | 119.781 | 0.259 | 0.031098 | -0.21437 | -0.000907 | -5.9e-05 | -0.1786 | -0.349 | -0.1882
current_sweep | 7 | 1 | 17 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9954 | 0.25991 | 99.9006 | 0.031157 | -0.0013897 | 0.031188 | 119.781 | 0.259 | 0.031119 | -0.21437 | -0.000907 | -3.8e-05 | -0.1786 | -0.349 | -0.1208
current_sweep | 7 | 1 | 18 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9954 | 0.25991 | 99.9006 | 0.031157 | -0.0013897 | 0.031188 | 119.781 | 0.259 | 0.031127 | -0.21437 | -0.000907 | -3e-05 | -0.1786 | -0.349 | -0.0952
current_sweep | 7 | 2 | 1 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0085 | 0.25989 | 99.9013 | 0.031158 | -0.0013843 | 0.031189 | 119.735 | 0.747 | 0.089556 | -0.27347 | 0.48711 | 0.058398 | -0.2279 | -4.1901 | -4.1921
current_sweep | 7 | 2 | 2 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0085 | 0.25989 | 99.9013 | 0.031158 | -0.0013843 | 0.031189 | 119.735 | 0.26 | 0.031174 | -0.27347 | 0.00011 | 1.6e-05 | -0.2279 | 0.0425 | 0.0508
current_sweep | 7 | 2 | 3 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0085 | 0.25989 | 99.9013 | 0.031158 | -0.0013843 | 0.031189 | 119.735 | 0.26 | 0.03117 | -0.27347 | 0.00011 | 1.2e-05 | -0.2279 | 0.0425 | 0.0379
current_sweep | 7 | 2 | 4 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0085 | 0.25989 | 99.9013 | 0.031158 | -0.0013843 | 0.031189 | 119.735 | 0.259 | 0.031107 | -0.27347 | -0.00089 | -5.1e-05 | -0.2279 | -0.3423 | -0.1642
current_sweep | 7 | 2 | 5 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0085 | 0.25989 | 99.9013 | 0.031158 | -0.0013843 | 0.031189 | 119.735 | 0.259 | 0.031128 | -0.27347 | -0.00089 | -3e-05 | -0.2279 | -0.3423 | -0.0969
current_sweep | 7 | 2 | 6 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0085 | 0.25989 | 99.9013 | 0.031158 | -0.0013843 | 0.031189 | 119.735 | 0.259 | 0.031025 | -0.27347 | -0.00089 | -0.000133 | -0.2279 | -0.3423 | -0.4274
current_sweep | 7 | 2 | 7 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9781 | 0.25993 | 99.9031 | 0.031156 | -0.0013723 | 0.031186 | 119.781 | 0.758 | 0.090809 | -0.1971 | 0.49807 | 0.059653 | -0.1643 | -2.7943 | -2.8439
current_sweep | 7 | 2 | 8 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9781 | 0.25993 | 99.9031 | 0.031156 | -0.0013723 | 0.031186 | 119.781 | 0.258 | 0.030989 | -0.1971 | -0.00193 | -0.000167 | -0.1643 | -0.7425 | -0.535
current_sweep | 7 | 2 | 9 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9781 | 0.25993 | 99.9031 | 0.031156 | -0.0013723 | 0.031186 | 119.781 | 0.259 | 0.031114 | -0.1971 | -0.00093 | -4.2e-05 | -0.1643 | -0.3578 | -0.1338
current_sweep | 7 | 2 | 10 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9781 | 0.25993 | 99.9031 | 0.031156 | -0.0013723 | 0.031186 | 119.781 | 0.259 | 0.031126 | -0.1971 | -0.00093 | -3e-05 | -0.1643 | -0.3578 | -0.0953
current_sweep | 7 | 2 | 11 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9781 | 0.25993 | 99.9031 | 0.031156 | -0.0013723 | 0.031186 | 119.781 | 0.259 | 0.031081 | -0.1971 | -0.00093 | -7.5e-05 | -0.1643 | -0.3578 | -0.2397
current_sweep | 7 | 2 | 12 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9781 | 0.25993 | 99.9031 | 0.031156 | -0.0013723 | 0.031186 | 119.781 | 0.259 | 0.031121 | -0.1971 | -0.00093 | -3.5e-05 | -0.1643 | -0.3578 | -0.1114
current_sweep | 7 | 2 | 13 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9927 | 0.2599 | 99.9007 | 0.031156 | -0.0013897 | 0.031187 | 119.778 | 0.259 | 0.03113 | -0.2147 | -0.000905 | -2.6e-05 | -0.1789 | -0.3482 | -0.0825
current_sweep | 7 | 2 | 14 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9927 | 0.2599 | 99.9007 | 0.031156 | -0.0013897 | 0.031187 | 119.778 | 0.259 | 0.031112 | -0.2147 | -0.000905 | -4.4e-05 | -0.1789 | -0.3482 | -0.1403
current_sweep | 7 | 2 | 15 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9927 | 0.2599 | 99.9007 | 0.031156 | -0.0013897 | 0.031187 | 119.778 | 0.26 | 0.03115 | -0.2147 | 9.5e-05 | -6e-06 | -0.1789 | 0.0366 | -0.0183
current_sweep | 7 | 2 | 16 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9927 | 0.2599 | 99.9007 | 0.031156 | -0.0013897 | 0.031187 | 119.778 | 0.259 | 0.031103 | -0.2147 | -0.000905 | -5.3e-05 | -0.1789 | -0.3482 | -0.1692
current_sweep | 7 | 2 | 17 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9927 | 0.2599 | 99.9007 | 0.031156 | -0.0013897 | 0.031187 | 119.778 | 0.259 | 0.031123 | -0.2147 | -0.000905 | -3.3e-05 | -0.1789 | -0.3482 | -0.105
current_sweep | 7 | 2 | 18 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.9927 | 0.2599 | 99.9007 | 0.031156 | -0.0013897 | 0.031187 | 119.778 | 0.259 | 0.031132 | -0.2147 | -0.000905 | -2.4e-05 | -0.1789 | -0.3482 | -0.0761
current_sweep | 7 | 3 | 1 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0105 | 0.25989 | 99.9011 | 0.031158 | -0.0013853 | 0.031189 | 119.738 | 0.748 | 0.08958 | -0.27248 | 0.48811 | 0.058422 | -0.227 | -4.061 | -4.167
current_sweep | 7 | 3 | 2 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0105 | 0.25989 | 99.9011 | 0.031158 | -0.0013853 | 0.031189 | 119.738 | 0.26 | 0.031184 | -0.27248 | 0.000113 | 2.6e-05 | -0.227 | 0.0433 | 0.0823
current_sweep | 7 | 3 | 3 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0105 | 0.25989 | 99.9011 | 0.031158 | -0.0013853 | 0.031189 | 119.738 | 0.26 | 0.031158 | -0.27248 | 0.000113 | 0 | -0.227 | 0.0433 | -0.0012
current_sweep | 7 | 3 | 4 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0105 | 0.25989 | 99.9011 | 0.031158 | -0.0013853 | 0.031189 | 119.738 | 0.259 | 0.031116 | -0.27248 | -0.000887 | -4.2e-05 | -0.227 | -0.3415 | -0.136
current_sweep | 7 | 3 | 5 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0105 | 0.25989 | 99.9011 | 0.031158 | -0.0013853 | 0.031189 | 119.738 | 0.259 | 0.031133 | -0.27248 | -0.000887 | -2.5e-05 | -0.227 | -0.3415 | -0.0814
current_sweep | 7 | 3 | 6 | 1 | 120 | 0.2599 | 0.999 | 60 | 120.0105 | 0.25989 | 99.9011 | 0.031158 | -0.0013853 | 0.031189 | 119.738 | 0.259 | 0.031025 | -0.27248 | -0.000887 | -0.000133 | -0.227 | -0.3415 | -0.428
current_sweep | 7 | 3 | 7 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9784 | 0.25994 | 99.9033 | 0.031156 | -0.0013711 | 0.031187 | 119.778 | 0.758 | 0.090857 | -0.2004 | 0.49806 | 0.059701 | -0.167 | -2.7964 | -2.795
current_sweep | 7 | 3 | 8 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9784 | 0.25994 | 99.9033 | 0.031156 | -0.0013711 | 0.031187 | 119.778 | 0.258 | 0.030976 | -0.2004 | -0.001935 | -0.00018 | -0.167 | -0.7446 | -0.5793
current_sweep | 7 | 3 | 9 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9784 | 0.25994 | 99.9033 | 0.031156 | -0.0013711 | 0.031187 | 119.778 | 0.259 | 0.031118 | -0.2004 | -0.000935 | -3.8e-05 | -0.167 | -0.3599 | -0.1235
current_sweep | 7 | 3 | 10 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9784 | 0.25994 | 99.9033 | 0.031156 | -0.0013711 | 0.031187 | 119.778 | 0.259 | 0.031114 | -0.2004 | -0.000935 | -4.2e-05 | -0.167 | -0.3599 | -0.1364
current_sweep | 7 | 3 | 11 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9784 | 0.25994 | 99.9033 | 0.031156 | -0.0013711 | 0.031187 | 119.778 | 0.259 | 0.031089 | -0.2004 | -0.000935 | -6.7e-05 | -0.167 | -0.3599 | -0.2166
current_sweep | 7 | 3 | 12 | 2 | 120 | 0.2599 | 0.999 | 60 | 119.9784 | 0.25994 | 99.9033 | 0.031156 | -0.0013711 | 0.031187 | 119.778 | 0.259 | 0.031111 | -0.2004 | -0.000935 | -4.5e-05 | -0.167 | -0.3599 | -0.146
current_sweep | 7 | 3 | 13 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.99 | 0.25991 | 99.9007 | 0.031156 | -0.0013894 | 0.031187 | 119.769 | 0.259 | 0.031124 | -0.22102 | -0.000913 | -3.2e-05 | -0.1842 | -0.3512 | -0.1026
current_sweep | 7 | 3 | 14 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.99 | 0.25991 | 99.9007 | 0.031156 | -0.0013894 | 0.031187 | 119.769 | 0.259 | 0.031127 | -0.22102 | -0.000913 | -2.9e-05 | -0.1842 | -0.3512 | -0.093
current_sweep | 7 | 3 | 15 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.99 | 0.25991 | 99.9007 | 0.031156 | -0.0013894 | 0.031187 | 119.769 | 0.26 | 0.03115 | -0.22102 | 8.7e-05 | -6e-06 | -0.1842 | 0.0335 | -0.0192
current_sweep | 7 | 3 | 16 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.99 | 0.25991 | 99.9007 | 0.031156 | -0.0013894 | 0.031187 | 119.769 | 0.259 | 0.031103 | -0.22102 | -0.000913 | -5.3e-05 | -0.1842 | -0.3512 | -0.17
current_sweep | 7 | 3 | 17 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.99 | 0.25991 | 99.9007 | 0.031156 | -0.0013894 | 0.031187 | 119.769 | 0.259 | 0.031121 | -0.22102 | -0.000913 | -3.5e-05 | -0.1842 | -0.3512 | -0.1123
current_sweep | 7 | 3 | 18 | 3 | 120 | 0.2599 | 0.999 | 60 | 119.99 | 0.25991 | 99.9007 | 0.031156 | -0.0013894 | 0.031187 | 119.769 | 0.259 | 0.031127 | -0.22102 | -0.000913 | -2.9e-05 | -0.1842 | -0.3512 | -0.093
current_sweep | 8 | 1 | 1 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0126 | 0.44719 | 99.9008 | 0.053615 | -0.0023881 | 0.053668 | 119.768 | 1.305 | 0.15638 | -0.2446 | 0.85781 | 0.10276 | -0.2038 | -2.725 | -2.776
current_sweep | 8 | 1 | 2 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0126 | 0.44719 | 99.9008 | 0.053615 | -0.0023881 | 0.053668 | 119.768 | 0.447 | 0.053638 | -0.2446 | -0.000186 | 2.3e-05 | -0.2038 | -0.0415 | 0.0435
current_sweep | 8 | 1 | 3 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0126 | 0.44719 | 99.9008 | 0.053615 | -0.0023881 | 0.053668 | 119.768 | 0.447 | 0.053607 | -0.2446 | -0.000186 | -8e-06 | -0.2038 | -0.0415 | -0.0143
current_sweep | 8 | 1 | 4 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0126 | 0.44719 | 99.9008 | 0.053615 | -0.0023881 | 0.053668 | 119.768 | 0.447 | 0.053567 | -0.2446 | -0.000186 | -4.8e-05 | -0.2038 | -0.0415 | -0.0889
current_sweep | 8 | 1 | 5 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0126 | 0.44719 | 99.9008 | 0.053615 | -0.0023881 | 0.053668 | 119.768 | 0.447 | 0.053622 | -0.2446 | -0.000186 | 7e-06 | -0.2038 | -0.0415 | 0.0137
current_sweep | 8 | 1 | 6 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0126 | 0.44719 | 99.9008 | 0.053615 | -0.0023881 | 0.053668 | 119.768 | 0.446 | 0.053437 | -0.2446 | -0.001186 | -0.000178 | -0.2038 | -0.2651 | -0.3314
current_sweep | 8 | 1 | 7 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.982 | 0.4473 | 99.9035 | 0.053616 | -0.0023562 | 0.053668 | 119.767 | 1.317 | 0.15784 | -0.21499 | 0.8697 | 0.10422 | -0.1792 | -1.8555 | -1.8716
current_sweep | 8 | 1 | 8 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.982 | 0.4473 | 99.9035 | 0.053616 | -0.0023562 | 0.053668 | 119.767 | 0.446 | 0.053449 | -0.21499 | -0.0013 | -0.000167 | -0.1792 | -0.2905 | -0.3117
current_sweep | 8 | 1 | 9 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.982 | 0.4473 | 99.9035 | 0.053616 | -0.0023562 | 0.053668 | 119.767 | 0.447 | 0.053602 | -0.21499 | -0.0003 | -1.4e-05 | -0.1792 | -0.067 | -0.0263
current_sweep | 8 | 1 | 10 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.982 | 0.4473 | 99.9035 | 0.053616 | -0.0023562 | 0.053668 | 119.767 | 0.447 | 0.053581 | -0.21499 | -0.0003 | -3.5e-05 | -0.1792 | -0.067 | -0.0655
current_sweep | 8 | 1 | 11 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.982 | 0.4473 | 99.9035 | 0.053616 | -0.0023562 | 0.053668 | 119.767 | 0.447 | 0.053549 | -0.21499 | -0.0003 | -6.7e-05 | -0.1792 | -0.067 | -0.1252
current_sweep | 8 | 1 | 12 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.982 | 0.4473 | 99.9035 | 0.053616 | -0.0023562 | 0.053668 | 119.767 | 0.447 | 0.053568 | -0.21499 | -0.0003 | -4.8e-05 | -0.1792 | -0.067 | -0.0898
current_sweep | 8 | 1 | 13 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9895 | 0.44718 | 99.9005 | 0.053603 | -0.0023926 | 0.053657 | 119.774 | 0.447 | 0.053543 | -0.21548 | -0.000177 | -6e-05 | -0.1796 | -0.0397 | -0.1123
current_sweep | 8 | 1 | 14 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9895 | 0.44718 | 99.9005 | 0.053603 | -0.0023926 | 0.053657 | 119.774 | 0.446 | 0.053525 | -0.21548 | -0.001177 | -7.8e-05 | -0.1796 | -0.2633 | -0.1459
current_sweep | 8 | 1 | 15 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9895 | 0.44718 | 99.9005 | 0.053603 | -0.0023926 | 0.053657 | 119.774 | 0.447 | 0.053592 | -0.21548 | -0.000177 | -1.1e-05 | -0.1796 | -0.0397 | -0.0209
current_sweep | 8 | 1 | 16 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9895 | 0.44718 | 99.9005 | 0.053603 | -0.0023926 | 0.053657 | 119.774 | 0.446 | 0.053523 | -0.21548 | -0.001177 | -8e-05 | -0.1796 | -0.2633 | -0.1496
current_sweep | 8 | 1 | 17 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9895 | 0.44718 | 99.9005 | 0.053603 | -0.0023926 | 0.053657 | 119.774 | 0.447 | 0.053558 | -0.21548 | -0.000177 | -4.5e-05 | -0.1796 | -0.0397 | -0.0843
current_sweep | 8 | 1 | 18 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9895 | 0.44718 | 99.9005 | 0.053603 | -0.0023926 | 0.053657 | 119.774 | 0.447 | 0.053546 | -0.21548 | -0.000177 | -5.7e-05 | -0.1796 | -0.0397 | -0.1067
current_sweep | 8 | 2 | 1 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0089 | 0.44719 | 99.9006 | 0.053613 | -0.0023892 | 0.053666 | 119.716 | 1.305 | 0.15633 | -0.2929 | 0.85781 | 0.10271 | -0.2441 | -2.7252 | -2.8059
current_sweep | 8 | 2 | 2 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0089 | 0.44719 | 99.9006 | 0.053613 | -0.0023892 | 0.053666 | 119.716 | 0.447 | 0.053592 | -0.2929 | -0.000187 | -2.1e-05 | -0.2441 | -0.0417 | -0.0392
current_sweep | 8 | 2 | 3 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0089 | 0.44719 | 99.9006 | 0.053613 | -0.0023892 | 0.053666 | 119.716 | 0.447 | 0.053609 | -0.2929 | -0.000187 | -4e-06 | -0.2441 | -0.0417 | -0.0075
current_sweep | 8 | 2 | 4 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0089 | 0.44719 | 99.9006 | 0.053613 | -0.0023892 | 0.053666 | 119.716 | 0.447 | 0.05356 | -0.2929 | -0.000187 | -5.3e-05 | -0.2441 | -0.0417 | -0.0989
current_sweep | 8 | 2 | 5 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0089 | 0.44719 | 99.9006 | 0.053613 | -0.0023892 | 0.053666 | 119.716 | 0.447 | 0.053587 | -0.2929 | -0.000187 | -2.6e-05 | -0.2441 | -0.0417 | -0.0485
current_sweep | 8 | 2 | 6 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0089 | 0.44719 | 99.9006 | 0.053613 | -0.0023892 | 0.053666 | 119.716 | 0.446 | 0.053438 | -0.2929 | -0.001187 | -0.000175 | -0.2441 | -0.2654 | -0.3264
current_sweep | 8 | 2 | 7 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9809 | 0.44731 | 99.9035 | 0.053616 | -0.0023572 | 0.053668 | 119.767 | 1.317 | 0.15781 | -0.21392 | 0.86969 | 0.10419 | -0.1783 | -1.8567 | -1.8917
current_sweep | 8 | 2 | 8 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9809 | 0.44731 | 99.9035 | 0.053616 | -0.0023572 | 0.053668 | 119.767 | 0.446 | 0.053446 | -0.21392 | -0.001305 | -0.00017 | -0.1783 | -0.2918 | -0.3176
current_sweep | 8 | 2 | 9 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9809 | 0.44731 | 99.9035 | 0.053616 | -0.0023572 | 0.053668 | 119.767 | 0.447 | 0.053593 | -0.21392 | -0.000305 | -2.3e-05 | -0.1783 | -0.0682 | -0.0434
current_sweep | 8 | 2 | 10 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9809 | 0.44731 | 99.9035 | 0.053616 | -0.0023572 | 0.053668 | 119.767 | 0.447 | 0.053574 | -0.21392 | -0.000305 | -4.2e-05 | -0.1783 | -0.0682 | -0.0788
current_sweep | 8 | 2 | 11 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9809 | 0.44731 | 99.9035 | 0.053616 | -0.0023572 | 0.053668 | 119.767 | 0.447 | 0.053556 | -0.21392 | -0.000305 | -6e-05 | -0.1783 | -0.0682 | -0.1124
current_sweep | 8 | 2 | 12 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9809 | 0.44731 | 99.9035 | 0.053616 | -0.0023572 | 0.053668 | 119.767 | 0.447 | 0.053557 | -0.21392 | -0.000305 | -5.9e-05 | -0.1783 | -0.0682 | -0.1105
current_sweep | 8 | 2 | 13 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9879 | 0.44717 | 99.9006 | 0.053602 | -0.0023913 | 0.053655 | 119.782 | 0.447 | 0.053543 | -0.20591 | -0.000174 | -5.9e-05 | -0.1716 | -0.0389 | -0.1103
current_sweep | 8 | 2 | 14 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9879 | 0.44717 | 99.9006 | 0.053602 | -0.0023913 | 0.053655 | 119.782 | 0.446 | 0.053533 | -0.20591 | -0.001174 | -6.9e-05 | -0.1716 | -0.2625 | -0.129
current_sweep | 8 | 2 | 15 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9879 | 0.44717 | 99.9006 | 0.053602 | -0.0023913 | 0.053655 | 119.782 | 0.447 | 0.053597 | -0.20591 | -0.000174 | -5e-06 | -0.1716 | -0.0389 | -0.0096
current_sweep | 8 | 2 | 16 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9879 | 0.44717 | 99.9006 | 0.053602 | -0.0023913 | 0.053655 | 119.782 | 0.447 | 0.053541 | -0.20591 | -0.000174 | -6.1e-05 | -0.1716 | -0.0389 | -0.114
current_sweep | 8 | 2 | 17 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9879 | 0.44717 | 99.9006 | 0.053602 | -0.0023913 | 0.053655 | 119.782 | 0.447 | 0.053563 | -0.20591 | -0.000174 | -3.9e-05 | -0.1716 | -0.0389 | -0.073
current_sweep | 8 | 2 | 18 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9879 | 0.44717 | 99.9006 | 0.053602 | -0.0023913 | 0.053655 | 119.782 | 0.447 | 0.053567 | -0.20591 | -0.000174 | -3.5e-05 | -0.1716 | -0.0389 | -0.0655
current_sweep | 8 | 3 | 1 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0007 | 0.44718 | 99.9007 | 0.053609 | -0.0023875 | 0.053662 | 119.729 | 1.306 | 0.15641 | -0.27166 | 0.85882 | 0.10281 | -0.2264 | -2.6494 | -2.7428
current_sweep | 8 | 3 | 2 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0007 | 0.44718 | 99.9007 | 0.053609 | -0.0023875 | 0.053662 | 119.729 | 0.447 | 0.053584 | -0.27166 | -0.000181 | -2.5e-05 | -0.2264 | -0.0405 | -0.0461
current_sweep | 8 | 3 | 3 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0007 | 0.44718 | 99.9007 | 0.053609 | -0.0023875 | 0.053662 | 119.729 | 0.447 | 0.053611 | -0.27166 | -0.000181 | 2e-06 | -0.2264 | -0.0405 | 0.0042
current_sweep | 8 | 3 | 4 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0007 | 0.44718 | 99.9007 | 0.053609 | -0.0023875 | 0.053662 | 119.729 | 0.447 | 0.053563 | -0.27166 | -0.000181 | -4.6e-05 | -0.2264 | -0.0405 | -0.0853
current_sweep | 8 | 3 | 5 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0007 | 0.44718 | 99.9007 | 0.053609 | -0.0023875 | 0.053662 | 119.729 | 0.447 | 0.053582 | -0.27166 | -0.000181 | -2.7e-05 | -0.2264 | -0.0405 | -0.0499
current_sweep | 8 | 3 | 6 | 1 | 120 | 0.4472 | 0.999 | 60 | 120.0007 | 0.44718 | 99.9007 | 0.053609 | -0.0023875 | 0.053662 | 119.729 | 0.446 | 0.053458 | -0.27166 | -0.001181 | -0.000151 | -0.2264 | -0.2641 | -0.2812
current_sweep | 8 | 3 | 7 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9841 | 0.44731 | 99.9035 | 0.053618 | -0.0023574 | 0.05367 | 119.768 | 1.317 | 0.15781 | -0.21615 | 0.86969 | 0.10419 | -0.1801 | -1.857 | -1.8953
current_sweep | 8 | 3 | 8 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9841 | 0.44731 | 99.9035 | 0.053618 | -0.0023574 | 0.05367 | 119.768 | 0.446 | 0.05346 | -0.21615 | -0.001306 | -0.000158 | -0.1801 | -0.292 | -0.2944
current_sweep | 8 | 3 | 9 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9841 | 0.44731 | 99.9035 | 0.053618 | -0.0023574 | 0.05367 | 119.768 | 0.447 | 0.05358 | -0.21615 | -0.000306 | -3.8e-05 | -0.1801 | -0.0685 | -0.0706
current_sweep | 8 | 3 | 10 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9841 | 0.44731 | 99.9035 | 0.053618 | -0.0023574 | 0.05367 | 119.768 | 0.447 | 0.053574 | -0.21615 | -0.000306 | -4.4e-05 | -0.1801 | -0.0685 | -0.0818
current_sweep | 8 | 3 | 11 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9841 | 0.44731 | 99.9035 | 0.053618 | -0.0023574 | 0.05367 | 119.768 | 0.447 | 0.053545 | -0.21615 | -0.000306 | -7.3e-05 | -0.1801 | -0.0685 | -0.1359
current_sweep | 8 | 3 | 12 | 2 | 120 | 0.4472 | 0.999 | 60 | 119.9841 | 0.44731 | 99.9035 | 0.053618 | -0.0023574 | 0.05367 | 119.768 | 0.447 | 0.053571 | -0.21615 | -0.000306 | -4.7e-05 | -0.1801 | -0.0685 | -0.0874
current_sweep | 8 | 3 | 13 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9903 | 0.44718 | 99.9007 | 0.053603 | -0.0023909 | 0.053657 | 119.784 | 0.447 | 0.053543 | -0.20632 | -0.000176 | -6e-05 | -0.1719 | -0.0393 | -0.1128
current_sweep | 8 | 3 | 14 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9903 | 0.44718 | 99.9007 | 0.053603 | -0.0023909 | 0.053657 | 119.784 | 0.446 | 0.053531 | -0.20632 | -0.001176 | -7.2e-05 | -0.1719 | -0.2629 | -0.1351
current_sweep | 8 | 3 | 15 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9903 | 0.44718 | 99.9007 | 0.053603 | -0.0023909 | 0.053657 | 119.784 | 0.447 | 0.053594 | -0.20632 | -0.000176 | -9e-06 | -0.1719 | -0.0393 | -0.0176
current_sweep | 8 | 3 | 16 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9903 | 0.44718 | 99.9007 | 0.053603 | -0.0023909 | 0.053657 | 119.784 | 0.446 | 0.05354 | -0.20632 | -0.001176 | -6.3e-05 | -0.1719 | -0.2629 | -0.1184
current_sweep | 8 | 3 | 17 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9903 | 0.44718 | 99.9007 | 0.053603 | -0.0023909 | 0.053657 | 119.784 | 0.447 | 0.053576 | -0.20632 | -0.000176 | -2.7e-05 | -0.1719 | -0.0393 | -0.0512
current_sweep | 8 | 3 | 18 | 3 | 120 | 0.4472 | 0.999 | 60 | 119.9903 | 0.44718 | 99.9007 | 0.053603 | -0.0023909 | 0.053657 | 119.784 | 0.447 | 0.05357 | -0.20632 | -0.000176 | -3.3e-05 | -0.1719 | -0.0393 | -0.0624
current_sweep | 9 | 1 | 1 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0231 | 0.76968 | 99.9008 | 0.092287 | -0.0041112 | 0.092379 | 119.74 | 2.27 | 0.2719 | -0.28306 | 1.5003 | 0.17961 | -0.2358 | -1.6905 | -1.792
current_sweep | 9 | 1 | 2 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0231 | 0.76968 | 99.9008 | 0.092287 | -0.0041112 | 0.092379 | 119.74 | 0.77 | 0.092258 | -0.28306 | 0.000322 | -2.9e-05 | -0.2358 | 0.0419 | -0.032
current_sweep | 9 | 1 | 3 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0231 | 0.76968 | 99.9008 | 0.092287 | -0.0041112 | 0.092379 | 119.74 | 0.77 | 0.092295 | -0.28306 | 0.000322 | 8e-06 | -0.2358 | 0.0419 | 0.0081
current_sweep | 9 | 1 | 4 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0231 | 0.76968 | 99.9008 | 0.092287 | -0.0041112 | 0.092379 | 119.74 | 0.77 | 0.092221 | -0.28306 | 0.000322 | -6.6e-05 | -0.2358 | 0.0419 | -0.072
current_sweep | 9 | 1 | 5 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0231 | 0.76968 | 99.9008 | 0.092287 | -0.0041112 | 0.092379 | 119.74 | 0.77 | 0.09225 | -0.28306 | 0.000322 | -3.7e-05 | -0.2358 | 0.0419 | -0.0406
current_sweep | 9 | 1 | 6 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0231 | 0.76968 | 99.9008 | 0.092287 | -0.0041112 | 0.092379 | 119.74 | 0.768 | 0.092017 | -0.28306 | -0.001678 | -0.00027 | -0.2358 | -0.218 | -0.2931
current_sweep | 9 | 1 | 7 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9854 | 0.7698 | 99.9038 | 0.092275 | -0.0040502 | 0.092364 | 119.773 | 2.283 | 0.27353 | -0.21243 | 1.5132 | 0.18125 | -0.177 | -1.1426 | -1.1911
current_sweep | 9 | 1 | 8 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9854 | 0.7698 | 99.9038 | 0.092275 | -0.0040502 | 0.092364 | 119.773 | 0.769 | 0.092191 | -0.21243 | -0.000795 | -8.4e-05 | -0.177 | -0.1033 | -0.0915
current_sweep | 9 | 1 | 9 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9854 | 0.7698 | 99.9038 | 0.092275 | -0.0040502 | 0.092364 | 119.773 | 0.77 | 0.092271 | -0.21243 | 0.000205 | -4e-06 | -0.177 | 0.0266 | -0.0048
current_sweep | 9 | 1 | 10 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9854 | 0.7698 | 99.9038 | 0.092275 | -0.0040502 | 0.092364 | 119.773 | 0.77 | 0.092237 | -0.21243 | 0.000205 | -3.8e-05 | -0.177 | 0.0266 | -0.0416
current_sweep | 9 | 1 | 11 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9854 | 0.7698 | 99.9038 | 0.092275 | -0.0040502 | 0.092364 | 119.773 | 0.769 | 0.092185 | -0.21243 | -0.000795 | -9e-05 | -0.177 | -0.1033 | -0.098
current_sweep | 9 | 1 | 12 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9854 | 0.7698 | 99.9038 | 0.092275 | -0.0040502 | 0.092364 | 119.773 | 0.769 | 0.092212 | -0.21243 | -0.000795 | -6.3e-05 | -0.177 | -0.1033 | -0.0687
current_sweep | 9 | 1 | 13 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9887 | 0.76966 | 99.9007 | 0.092258 | -0.0041136 | 0.09235 | 119.766 | 0.769 | 0.092171 | -0.22269 | -0.000656 | -8.7e-05 | -0.1856 | -0.0852 | -0.0946
current_sweep | 9 | 1 | 14 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9887 | 0.76966 | 99.9007 | 0.092258 | -0.0041136 | 0.09235 | 119.766 | 0.769 | 0.092162 | -0.22269 | -0.000656 | -9.6e-05 | -0.1856 | -0.0852 | -0.1044
current_sweep | 9 | 1 | 15 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9887 | 0.76966 | 99.9007 | 0.092258 | -0.0041136 | 0.09235 | 119.766 | 0.77 | 0.092262 | -0.22269 | 0.000344 | 4e-06 | -0.1856 | 0.0447 | 0.004
current_sweep | 9 | 1 | 16 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9887 | 0.76966 | 99.9007 | 0.092258 | -0.0041136 | 0.09235 | 119.766 | 0.769 | 0.092135 | -0.22269 | -0.000656 | -0.000123 | -0.1856 | -0.0852 | -0.1337
current_sweep | 9 | 1 | 17 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9887 | 0.76966 | 99.9007 | 0.092258 | -0.0041136 | 0.09235 | 119.766 | 0.769 | 0.092201 | -0.22269 | -0.000656 | -5.7e-05 | -0.1856 | -0.0852 | -0.0621
current_sweep | 9 | 1 | 18 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9887 | 0.76966 | 99.9007 | 0.092258 | -0.0041136 | 0.09235 | 119.766 | 0.769 | 0.092183 | -0.22269 | -0.000656 | -7.5e-05 | -0.1856 | -0.0852 | -0.0816
current_sweep | 9 | 2 | 1 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0091 | 0.76968 | 99.9009 | 0.092278 | -0.00411 | 0.092369 | 119.748 | 2.271 | 0.27196 | -0.26108 | 1.5013 | 0.17969 | -0.2175 | -1.648 | -1.7587
current_sweep | 9 | 2 | 2 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0091 | 0.76968 | 99.9009 | 0.092278 | -0.00411 | 0.092369 | 119.748 | 0.77 | 0.092316 | -0.26108 | 0.000316 | 3.8e-05 | -0.2175 | 0.041 | 0.0416
current_sweep | 9 | 2 | 3 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0091 | 0.76968 | 99.9009 | 0.092278 | -0.00411 | 0.092369 | 119.748 | 0.77 | 0.092293 | -0.26108 | 0.000316 | 1.5e-05 | -0.2175 | 0.041 | 0.0167
current_sweep | 9 | 2 | 4 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0091 | 0.76968 | 99.9009 | 0.092278 | -0.00411 | 0.092369 | 119.748 | 0.77 | 0.092249 | -0.26108 | 0.000316 | -2.9e-05 | -0.2175 | 0.041 | -0.031
current_sweep | 9 | 2 | 5 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0091 | 0.76968 | 99.9009 | 0.092278 | -0.00411 | 0.092369 | 119.748 | 0.77 | 0.092288 | -0.26108 | 0.000316 | 1e-05 | -0.2175 | 0.041 | 0.0113
current_sweep | 9 | 2 | 6 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0091 | 0.76968 | 99.9009 | 0.092278 | -0.00411 | 0.092369 | 119.748 | 0.768 | 0.092043 | -0.26108 | -0.001684 | -0.000235 | -0.2175 | -0.2188 | -0.2542
current_sweep | 9 | 2 | 7 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9827 | 0.76982 | 99.9036 | 0.092276 | -0.0040534 | 0.092365 | 119.753 | 2.283 | 0.27353 | -0.22973 | 1.5132 | 0.18125 | -0.1915 | -1.1457 | -1.1922
current_sweep | 9 | 2 | 8 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9827 | 0.76982 | 99.9036 | 0.092276 | -0.0040534 | 0.092365 | 119.753 | 0.769 | 0.092142 | -0.22973 | -0.00082 | -0.000134 | -0.1915 | -0.1065 | -0.1453
current_sweep | 9 | 2 | 9 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9827 | 0.76982 | 99.9036 | 0.092276 | -0.0040534 | 0.092365 | 119.753 | 0.77 | 0.092281 | -0.22973 | 0.00018 | 5e-06 | -0.1915 | 0.0234 | 0.0053
current_sweep | 9 | 2 | 10 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9827 | 0.76982 | 99.9036 | 0.092276 | -0.0040534 | 0.092365 | 119.753 | 0.77 | 0.09222 | -0.22973 | 0.00018 | -5.6e-05 | -0.1915 | 0.0234 | -0.0608
current_sweep | 9 | 2 | 11 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9827 | 0.76982 | 99.9036 | 0.092276 | -0.0040534 | 0.092365 | 119.753 | 0.769 | 0.092162 | -0.22973 | -0.00082 | -0.000114 | -0.1915 | -0.1065 | -0.1237
current_sweep | 9 | 2 | 12 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9827 | 0.76982 | 99.9036 | 0.092276 | -0.0040534 | 0.092365 | 119.753 | 0.769 | 0.092222 | -0.22973 | -0.00082 | -5.4e-05 | -0.1915 | -0.1065 | -0.0586
current_sweep | 9 | 2 | 13 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9937 | 0.76968 | 99.9007 | 0.092265 | -0.0041142 | 0.092357 | 119.778 | 0.769 | 0.092179 | -0.21572 | -0.000682 | -8.6e-05 | -0.1798 | -0.0886 | -0.0935
current_sweep | 9 | 2 | 14 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9937 | 0.76968 | 99.9007 | 0.092265 | -0.0041142 | 0.092357 | 119.778 | 0.769 | 0.092152 | -0.21572 | -0.000682 | -0.000113 | -0.1798 | -0.0886 | -0.1228
current_sweep | 9 | 2 | 15 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9937 | 0.76968 | 99.9007 | 0.092265 | -0.0041142 | 0.092357 | 119.778 | 0.77 | 0.092244 | -0.21572 | 0.000318 | -2.1e-05 | -0.1798 | 0.0414 | -0.0231
current_sweep | 9 | 2 | 16 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9937 | 0.76968 | 99.9007 | 0.092265 | -0.0041142 | 0.092357 | 119.778 | 0.769 | 0.092144 | -0.21572 | -0.000682 | -0.000121 | -0.1798 | -0.0886 | -0.1315
current_sweep | 9 | 2 | 17 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9937 | 0.76968 | 99.9007 | 0.092265 | -0.0041142 | 0.092357 | 119.778 | 0.769 | 0.092204 | -0.21572 | -0.000682 | -6.1e-05 | -0.1798 | -0.0886 | -0.0664
current_sweep | 9 | 2 | 18 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9937 | 0.76968 | 99.9007 | 0.092265 | -0.0041142 | 0.092357 | 119.778 | 0.769 | 0.092166 | -0.21572 | -0.000682 | -9.9e-05 | -0.1798 | -0.0886 | -0.1076
current_sweep | 9 | 3 | 1 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0164 | 0.76967 | 99.9009 | 0.092281 | -0.0041082 | 0.092373 | 119.732 | 2.271 | 0.27193 | -0.28442 | 1.5013 | 0.17964 | -0.237 | -1.6458 | -1.7763
current_sweep | 9 | 3 | 2 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0164 | 0.76967 | 99.9009 | 0.092281 | -0.0041082 | 0.092373 | 119.732 | 0.77 | 0.092276 | -0.28442 | 0.000333 | -5e-06 | -0.237 | 0.0432 | -0.0056
current_sweep | 9 | 3 | 3 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0164 | 0.76967 | 99.9009 | 0.092281 | -0.0041082 | 0.092373 | 119.732 | 0.77 | 0.09233 | -0.28442 | 0.000333 | 4.9e-05 | -0.237 | 0.0432 | 0.0529
current_sweep | 9 | 3 | 4 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0164 | 0.76967 | 99.9009 | 0.092281 | -0.0041082 | 0.092373 | 119.732 | 0.77 | 0.092229 | -0.28442 | 0.000333 | -5.2e-05 | -0.237 | 0.0432 | -0.0566
current_sweep | 9 | 3 | 5 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0164 | 0.76967 | 99.9009 | 0.092281 | -0.0041082 | 0.092373 | 119.732 | 0.77 | 0.092304 | -0.28442 | 0.000333 | 2.3e-05 | -0.237 | 0.0432 | 0.0247
current_sweep | 9 | 3 | 6 | 1 | 120 | 0.7697 | 0.999 | 60 | 120.0164 | 0.76967 | 99.9009 | 0.092281 | -0.0041082 | 0.092373 | 119.732 | 0.768 | 0.092064 | -0.28442 | -0.001667 | -0.000217 | -0.237 | -0.2166 | -0.2354
current_sweep | 9 | 3 | 7 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9808 | 0.76983 | 99.9039 | 0.092276 | -0.0040474 | 0.092364 | 119.764 | 2.283 | 0.27351 | -0.21678 | 1.5132 | 0.18124 | -0.1807 | -1.1465 | -1.1967
current_sweep | 9 | 3 | 8 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9808 | 0.76983 | 99.9039 | 0.092276 | -0.0040474 | 0.092364 | 119.764 | 0.769 | 0.092142 | -0.21678 | -0.000826 | -0.000134 | -0.1807 | -0.1073 | -0.1448
current_sweep | 9 | 3 | 9 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9808 | 0.76983 | 99.9039 | 0.092276 | -0.0040474 | 0.092364 | 119.764 | 0.77 | 0.092262 | -0.21678 | 0.000174 | -1.4e-05 | -0.1807 | 0.0226 | -0.0148
current_sweep | 9 | 3 | 10 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9808 | 0.76983 | 99.9039 | 0.092276 | -0.0040474 | 0.092364 | 119.764 | 0.77 | 0.092223 | -0.21678 | 0.000174 | -5.3e-05 | -0.1807 | 0.0226 | -0.057
current_sweep | 9 | 3 | 11 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9808 | 0.76983 | 99.9039 | 0.092276 | -0.0040474 | 0.092364 | 119.764 | 0.769 | 0.092185 | -0.21678 | -0.000826 | -9.1e-05 | -0.1807 | -0.1073 | -0.0982
current_sweep | 9 | 3 | 12 | 2 | 120 | 0.7697 | 0.999 | 60 | 119.9808 | 0.76983 | 99.9039 | 0.092276 | -0.0040474 | 0.092364 | 119.764 | 0.769 | 0.092206 | -0.21678 | -0.000826 | -7e-05 | -0.1807 | -0.1073 | -0.0755
current_sweep | 9 | 3 | 13 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9894 | 0.76966 | 99.9006 | 0.092259 | -0.0041169 | 0.092351 | 119.782 | 0.769 | 0.092193 | -0.20741 | -0.00066 | -6.6e-05 | -0.1729 | -0.0858 | -0.0718
current_sweep | 9 | 3 | 14 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9894 | 0.76966 | 99.9006 | 0.092259 | -0.0041169 | 0.092351 | 119.782 | 0.769 | 0.092155 | -0.20741 | -0.00066 | -0.000104 | -0.1729 | -0.0858 | -0.113
current_sweep | 9 | 3 | 15 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9894 | 0.76966 | 99.9006 | 0.092259 | -0.0041169 | 0.092351 | 119.782 | 0.77 | 0.092251 | -0.20741 | 0.00034 | -8e-06 | -0.1729 | 0.0441 | -0.009
current_sweep | 9 | 3 | 16 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9894 | 0.76966 | 99.9006 | 0.092259 | -0.0041169 | 0.092351 | 119.782 | 0.769 | 0.092158 | -0.20741 | -0.00066 | -0.000101 | -0.1729 | -0.0858 | -0.1098
current_sweep | 9 | 3 | 17 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9894 | 0.76966 | 99.9006 | 0.092259 | -0.0041169 | 0.092351 | 119.782 | 0.769 | 0.092215 | -0.20741 | -0.00066 | -4.4e-05 | -0.1729 | -0.0858 | -0.048
current_sweep | 9 | 3 | 18 | 3 | 120 | 0.7697 | 0.999 | 60 | 119.9894 | 0.76966 | 99.9006 | 0.092259 | -0.0041169 | 0.092351 | 119.782 | 0.769 | 0.092175 | -0.20741 | -0.00066 | -8.4e-05 | -0.1729 | -0.0858 | -0.0913
current_sweep | 10 | 1 | 1 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0043 | 1.3244 | 99.901 | 0.15877 | -0.0070589 | 0.15893 | 119.751 | 3.934 | 0.47099 | -0.25334 | 2.6096 | 0.31222 | -0.2111 | -0.98464 | -1.1178
current_sweep | 10 | 1 | 2 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0043 | 1.3244 | 99.901 | 0.15877 | -0.0070589 | 0.15893 | 119.751 | 1.326 | 0.15875 | -0.25334 | 0.001626 | -2.6e-05 | -0.2111 | 0.1228 | -0.0165
current_sweep | 10 | 1 | 3 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0043 | 1.3244 | 99.901 | 0.15877 | -0.0070589 | 0.15893 | 119.751 | 1.326 | 0.15886 | -0.25334 | 0.001626 | 8.3e-05 | -0.2111 | 0.1228 | 0.0522
current_sweep | 10 | 1 | 4 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0043 | 1.3244 | 99.901 | 0.15877 | -0.0070589 | 0.15893 | 119.751 | 1.325 | 0.15868 | -0.25334 | 0.000626 | -9.1e-05 | -0.2111 | 0.0473 | -0.0574
current_sweep | 10 | 1 | 5 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0043 | 1.3244 | 99.901 | 0.15877 | -0.0070589 | 0.15893 | 119.751 | 1.326 | 0.15875 | -0.25334 | 0.001626 | -1.9e-05 | -0.2111 | 0.1228 | -0.0121
current_sweep | 10 | 1 | 6 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0043 | 1.3244 | 99.901 | 0.15877 | -0.0070589 | 0.15893 | 119.751 | 1.323 | 0.15844 | -0.25334 | -0.001374 | -0.000331 | -0.2111 | -0.1037 | -0.2086
current_sweep | 10 | 1 | 7 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9786 | 1.3247 | 99.9037 | 0.15878 | -0.0069707 | 0.15894 | 119.768 | 3.948 | 0.47289 | -0.21064 | 2.6233 | 0.31411 | -0.1756 | -0.65755 | -0.72608
current_sweep | 10 | 1 | 8 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9786 | 1.3247 | 99.9037 | 0.15878 | -0.0069707 | 0.15894 | 119.768 | 1.325 | 0.15875 | -0.21064 | 0.000289 | -3e-05 | -0.1756 | 0.0218 | -0.0188
current_sweep | 10 | 1 | 9 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9786 | 1.3247 | 99.9037 | 0.15878 | -0.0069707 | 0.15894 | 119.768 | 1.325 | 0.15881 | -0.21064 | 0.000289 | 2.4e-05 | -0.1756 | 0.0218 | 0.0152
current_sweep | 10 | 1 | 10 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9786 | 1.3247 | 99.9037 | 0.15878 | -0.0069707 | 0.15894 | 119.768 | 1.325 | 0.15873 | -0.21064 | 0.000289 | -5.3e-05 | -0.1756 | 0.0218 | -0.0333
current_sweep | 10 | 1 | 11 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9786 | 1.3247 | 99.9037 | 0.15878 | -0.0069707 | 0.15894 | 119.768 | 1.324 | 0.15867 | -0.21064 | -0.000711 | -0.000113 | -0.1756 | -0.0536 | -0.0711
current_sweep | 10 | 1 | 12 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9786 | 1.3247 | 99.9037 | 0.15878 | -0.0069707 | 0.15894 | 119.768 | 1.324 | 0.15868 | -0.21064 | -0.000711 | -0.0001 | -0.1756 | -0.0536 | -0.0629
current_sweep | 10 | 1 | 13 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9907 | 1.3244 | 99.9006 | 0.15876 | -0.0070823 | 0.15892 | 119.767 | 1.324 | 0.15866 | -0.22375 | -0.000425 | -0.000105 | -0.1865 | -0.0321 | -0.066
current_sweep | 10 | 1 | 14 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9907 | 1.3244 | 99.9006 | 0.15876 | -0.0070823 | 0.15892 | 119.767 | 1.323 | 0.15857 | -0.22375 | -0.001425 | -0.000196 | -0.1865 | -0.1076 | -0.1233
current_sweep | 10 | 1 | 15 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9907 | 1.3244 | 99.9006 | 0.15876 | -0.0070823 | 0.15892 | 119.767 | 1.325 | 0.15877 | -0.22375 | 0.000575 | 6e-06 | -0.1865 | 0.0434 | 0.0039
current_sweep | 10 | 1 | 16 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9907 | 1.3244 | 99.9006 | 0.15876 | -0.0070823 | 0.15892 | 119.767 | 1.324 | 0.1586 | -0.22375 | -0.000425 | -0.000162 | -0.1865 | -0.0321 | -0.1019
current_sweep | 10 | 1 | 17 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9907 | 1.3244 | 99.9006 | 0.15876 | -0.0070823 | 0.15892 | 119.767 | 1.324 | 0.15865 | -0.22375 | -0.000425 | -0.000114 | -0.1865 | -0.0321 | -0.0716
current_sweep | 10 | 1 | 18 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9907 | 1.3244 | 99.9006 | 0.15876 | -0.0070823 | 0.15892 | 119.767 | 1.324 | 0.15863 | -0.22375 | -0.000425 | -0.00013 | -0.1865 | -0.0321 | -0.0817
current_sweep | 10 | 2 | 1 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0009 | 1.3244 | 99.9006 | 0.15877 | -0.0070676 | 0.15893 | 119.782 | 3.934 | 0.47132 | -0.2189 | 2.6096 | 0.31255 | -0.1824 | -0.98654 | -1.0483
current_sweep | 10 | 2 | 2 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0009 | 1.3244 | 99.9006 | 0.15877 | -0.0070676 | 0.15893 | 119.782 | 1.326 | 0.15881 | -0.2189 | 0.001601 | 4e-05 | -0.1824 | 0.1209 | 0.0251
current_sweep | 10 | 2 | 3 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0009 | 1.3244 | 99.9006 | 0.15877 | -0.0070676 | 0.15893 | 119.782 | 1.326 | 0.15888 | -0.2189 | 0.001601 | 0.000112 | -0.1824 | 0.1209 | 0.0705
current_sweep | 10 | 2 | 4 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0009 | 1.3244 | 99.9006 | 0.15877 | -0.0070676 | 0.15893 | 119.782 | 1.325 | 0.15878 | -0.2189 | 0.000601 | 1.3e-05 | -0.1824 | 0.0454 | 0.0081
current_sweep | 10 | 2 | 5 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0009 | 1.3244 | 99.9006 | 0.15877 | -0.0070676 | 0.15893 | 119.782 | 1.326 | 0.15882 | -0.2189 | 0.001601 | 5e-05 | -0.1824 | 0.1209 | 0.0314
current_sweep | 10 | 2 | 6 | 1 | 120 | 1.3246 | 0.999 | 60 | 120.0009 | 1.3244 | 99.9006 | 0.15877 | -0.0070676 | 0.15893 | 119.782 | 1.323 | 0.15843 | -0.2189 | -0.001399 | -0.000338 | -0.1824 | -0.1056 | -0.2129
current_sweep | 10 | 2 | 7 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9853 | 1.3247 | 99.9035 | 0.15879 | -0.0069797 | 0.15895 | 119.768 | 3.948 | 0.47287 | -0.21727 | 2.6233 | 0.31408 | -0.1811 | -0.65732 | -0.73491
current_sweep | 10 | 2 | 8 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9853 | 1.3247 | 99.9035 | 0.15879 | -0.0069797 | 0.15895 | 119.768 | 1.325 | 0.15876 | -0.21727 | 0.000292 | -3.5e-05 | -0.1811 | 0.0221 | -0.022
current_sweep | 10 | 2 | 9 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9853 | 1.3247 | 99.9035 | 0.15879 | -0.0069797 | 0.15895 | 119.768 | 1.325 | 0.15877 | -0.21727 | 0.000292 | -2.2e-05 | -0.1811 | 0.0221 | -0.0138
current_sweep | 10 | 2 | 10 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9853 | 1.3247 | 99.9035 | 0.15879 | -0.0069797 | 0.15895 | 119.768 | 1.325 | 0.15871 | -0.21727 | 0.000292 | -8.2e-05 | -0.1811 | 0.0221 | -0.0516
current_sweep | 10 | 2 | 11 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9853 | 1.3247 | 99.9035 | 0.15879 | -0.0069797 | 0.15895 | 119.768 | 1.324 | 0.15866 | -0.21727 | -0.000708 | -0.000128 | -0.1811 | -0.0534 | -0.0806
current_sweep | 10 | 2 | 12 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9853 | 1.3247 | 99.9035 | 0.15879 | -0.0069797 | 0.15895 | 119.768 | 1.324 | 0.15867 | -0.21727 | -0.000708 | -0.000126 | -0.1811 | -0.0534 | -0.0793
current_sweep | 10 | 2 | 13 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9886 | 1.3244 | 99.9008 | 0.15875 | -0.007074 | 0.15891 | 119.783 | 1.324 | 0.15868 | -0.20564 | -0.000376 | -7.6e-05 | -0.1714 | -0.0284 | -0.0481
current_sweep | 10 | 2 | 14 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9886 | 1.3244 | 99.9008 | 0.15875 | -0.007074 | 0.15891 | 119.783 | 1.323 | 0.15854 | -0.20564 | -0.001376 | -0.000212 | -0.1714 | -0.1039 | -0.1338
current_sweep | 10 | 2 | 15 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9886 | 1.3244 | 99.9008 | 0.15875 | -0.007074 | 0.15891 | 119.783 | 1.325 | 0.15878 | -0.20564 | 0.000624 | 2.7e-05 | -0.1714 | 0.0471 | 0.0168
current_sweep | 10 | 2 | 16 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9886 | 1.3244 | 99.9008 | 0.15875 | -0.007074 | 0.15891 | 119.783 | 1.324 | 0.15861 | -0.20564 | -0.000376 | -0.000139 | -0.1714 | -0.0284 | -0.0878
current_sweep | 10 | 2 | 17 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9886 | 1.3244 | 99.9008 | 0.15875 | -0.007074 | 0.15891 | 119.783 | 1.324 | 0.15862 | -0.20564 | -0.000376 | -0.000127 | -0.1714 | -0.0284 | -0.0802
current_sweep | 10 | 2 | 18 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9886 | 1.3244 | 99.9008 | 0.15875 | -0.007074 | 0.15891 | 119.783 | 1.324 | 0.15864 | -0.20564 | -0.000376 | -0.000114 | -0.1714 | -0.0284 | -0.0721
current_sweep | 10 | 3 | 1 | 1 | 120 | 1.3246 | 0.999 | 60 | 119.9834 | 1.3244 | 99.9007 | 0.15874 | -0.0070627 | 0.1589 | 119.694 | 3.934 | 0.47093 | -0.28935 | 2.6096 | 0.31219 | -0.2412 | -0.98315 | -1.1114
current_sweep | 10 | 3 | 2 | 1 | 120 | 1.3246 | 0.999 | 60 | 119.9834 | 1.3244 | 99.9007 | 0.15874 | -0.0070627 | 0.1589 | 119.694 | 1.326 | 0.15883 | -0.28935 | 0.001646 | 8.6e-05 | -0.2412 | 0.1243 | 0.0545
current_sweep | 10 | 3 | 3 | 1 | 120 | 1.3246 | 0.999 | 60 | 119.9834 | 1.3244 | 99.9007 | 0.15874 | -0.0070627 | 0.1589 | 119.694 | 1.326 | 0.15886 | -0.28935 | 0.001646 | 0.000119 | -0.2412 | 0.1243 | 0.0753
current_sweep | 10 | 3 | 4 | 1 | 120 | 1.3246 | 0.999 | 60 | 119.9834 | 1.3244 | 99.9007 | 0.15874 | -0.0070627 | 0.1589 | 119.694 | 1.325 | 0.15865 | -0.28935 | 0.000646 | -8.9e-05 | -0.2412 | 0.0488 | -0.0558
current_sweep | 10 | 3 | 5 | 1 | 120 | 1.3246 | 0.999 | 60 | 119.9834 | 1.3244 | 99.9007 | 0.15874 | -0.0070627 | 0.1589 | 119.694 | 1.326 | 0.15891 | -0.28935 | 0.001646 | 0.000166 | -0.2412 | 0.1243 | 0.1049
current_sweep | 10 | 3 | 6 | 1 | 120 | 1.3246 | 0.999 | 60 | 119.9834 | 1.3244 | 99.9007 | 0.15874 | -0.0070627 | 0.1589 | 119.694 | 1.323 | 0.15844 | -0.28935 | -0.001354 | -0.000302 | -0.2412 | -0.1022 | -0.19
current_sweep | 10 | 3 | 7 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9806 | 1.3247 | 99.9034 | 0.15879 | -0.0069802 | 0.15894 | 119.766 | 3.948 | 0.4729 | -0.21464 | 2.6233 | 0.31412 | -0.1789 | -0.65737 | -0.7254
current_sweep | 10 | 3 | 8 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9806 | 1.3247 | 99.9034 | 0.15879 | -0.0069802 | 0.15894 | 119.766 | 1.325 | 0.15874 | -0.21464 | 0.000292 | -5e-05 | -0.1789 | 0.022 | -0.0314
current_sweep | 10 | 3 | 9 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9806 | 1.3247 | 99.9034 | 0.15879 | -0.0069802 | 0.15894 | 119.766 | 1.325 | 0.15877 | -0.21464 | 0.000292 | -1.9e-05 | -0.1789 | 0.022 | -0.0119
current_sweep | 10 | 3 | 10 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9806 | 1.3247 | 99.9034 | 0.15879 | -0.0069802 | 0.15894 | 119.766 | 1.325 | 0.15873 | -0.21464 | 0.000292 | -5.3e-05 | -0.1789 | 0.022 | -0.0333
current_sweep | 10 | 3 | 11 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9806 | 1.3247 | 99.9034 | 0.15879 | -0.0069802 | 0.15894 | 119.766 | 1.324 | 0.15864 | -0.21464 | -0.000708 | -0.00015 | -0.1789 | -0.0535 | -0.0944
current_sweep | 10 | 3 | 12 | 2 | 120 | 1.3246 | 0.999 | 60 | 119.9806 | 1.3247 | 99.9034 | 0.15879 | -0.0069802 | 0.15894 | 119.766 | 1.324 | 0.15866 | -0.21464 | -0.000708 | -0.000128 | -0.1789 | -0.0535 | -0.0805
current_sweep | 10 | 3 | 13 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9932 | 1.3244 | 99.9006 | 0.15876 | -0.0070806 | 0.15892 | 119.78 | 1.324 | 0.15866 | -0.21319 | -0.000371 | -9.9e-05 | -0.1777 | -0.028 | -0.0621
current_sweep | 10 | 3 | 14 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9932 | 1.3244 | 99.9006 | 0.15876 | -0.0070806 | 0.15892 | 119.78 | 1.324 | 0.15858 | -0.21319 | -0.000371 | -0.000174 | -0.1777 | -0.028 | -0.1093
current_sweep | 10 | 3 | 15 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9932 | 1.3244 | 99.9006 | 0.15876 | -0.0070806 | 0.15892 | 119.78 | 1.325 | 0.15881 | -0.21319 | 0.000629 | 5e-05 | -0.1777 | 0.0475 | 0.0318
current_sweep | 10 | 3 | 16 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9932 | 1.3244 | 99.9006 | 0.15876 | -0.0070806 | 0.15892 | 119.78 | 1.324 | 0.15861 | -0.21319 | -0.000371 | -0.000146 | -0.1777 | -0.028 | -0.0917
current_sweep | 10 | 3 | 17 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9932 | 1.3244 | 99.9006 | 0.15876 | -0.0070806 | 0.15892 | 119.78 | 1.324 | 0.15867 | -0.21319 | -0.000371 | -8.5e-05 | -0.1777 | -0.028 | -0.0532
current_sweep | 10 | 3 | 18 | 3 | 120 | 1.3246 | 0.999 | 60 | 119.9932 | 1.3244 | 99.9006 | 0.15876 | -0.0070806 | 0.15892 | 119.78 | 1.324 | 0.15867 | -0.21319 | -0.000371 | -9.2e-05 | -0.1777 | -0.028 | -0.0576
current_sweep | 11 | 1 | 1 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9943 | 2.2794 | 99.9007 | 0.27324 | -0.01217 | 0.27351 | 119.771 | 6.803 | 0.81461 | -0.22327 | 4.5236 | 0.54136 | -0.1861 | -0.51382 | -0.62418
current_sweep | 11 | 1 | 2 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9943 | 2.2794 | 99.9007 | 0.27324 | -0.01217 | 0.27351 | 119.771 | 2.283 | 0.27346 | -0.22327 | 0.003621 | 0.000223 | -0.1861 | 0.1589 | 0.0817
current_sweep | 11 | 1 | 3 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9943 | 2.2794 | 99.9007 | 0.27324 | -0.01217 | 0.27351 | 119.771 | 2.283 | 0.2733 | -0.22327 | 0.003621 | 5.9e-05 | -0.1861 | 0.1589 | 0.0217
current_sweep | 11 | 1 | 4 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9943 | 2.2794 | 99.9007 | 0.27324 | -0.01217 | 0.27351 | 119.771 | 2.281 | 0.27319 | -0.22327 | 0.001621 | -5.4e-05 | -0.1861 | 0.0711 | -0.0197
current_sweep | 11 | 1 | 5 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9943 | 2.2794 | 99.9007 | 0.27324 | -0.01217 | 0.27351 | 119.771 | 2.283 | 0.27346 | -0.22327 | 0.003621 | 0.000218 | -0.1861 | 0.1589 | 0.0798
current_sweep | 11 | 1 | 6 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9943 | 2.2794 | 99.9007 | 0.27324 | -0.01217 | 0.27351 | 119.771 | 2.277 | 0.27258 | -0.22327 | -0.002379 | -0.000661 | -0.1861 | -0.1044 | -0.2419
current_sweep | 11 | 1 | 7 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9866 | 2.28 | 99.904 | 0.2733 | -0.011983 | 0.27357 | 119.752 | 6.816 | 0.81632 | -0.23458 | 4.536 | 0.54302 | -0.1955 | -0.34989 | -0.43762
current_sweep | 11 | 1 | 8 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9866 | 2.28 | 99.904 | 0.2733 | -0.011983 | 0.27357 | 119.752 | 2.281 | 0.27321 | -0.23458 | 0.001023 | -9.4e-05 | -0.1955 | 0.0449 | -0.0344
current_sweep | 11 | 1 | 9 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9866 | 2.28 | 99.904 | 0.2733 | -0.011983 | 0.27357 | 119.752 | 2.282 | 0.27336 | -0.23458 | 0.002023 | 6e-05 | -0.1955 | 0.0887 | 0.0219
current_sweep | 11 | 1 | 10 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9866 | 2.28 | 99.904 | 0.2733 | -0.011983 | 0.27357 | 119.752 | 2.28 | 0.27317 | -0.23458 | 2.3e-05 | -0.000136 | -0.1955 | 0.001 | -0.0498
current_sweep | 11 | 1 | 11 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9866 | 2.28 | 99.904 | 0.2733 | -0.011983 | 0.27357 | 119.752 | 2.279 | 0.27298 | -0.23458 | -0.000977 | -0.000327 | -0.1955 | -0.0429 | -0.1197
current_sweep | 11 | 1 | 12 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9866 | 2.28 | 99.904 | 0.2733 | -0.011983 | 0.27357 | 119.752 | 2.28 | 0.27318 | -0.23458 | 2.3e-05 | -0.000121 | -0.1955 | 0.001 | -0.0443
current_sweep | 11 | 1 | 13 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9903 | 2.2795 | 99.9007 | 0.27325 | -0.012186 | 0.27352 | 119.779 | 2.279 | 0.27305 | -0.21126 | -0.000521 | -0.000203 | -0.1761 | -0.0229 | -0.0742
current_sweep | 11 | 1 | 14 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9903 | 2.2795 | 99.9007 | 0.27325 | -0.012186 | 0.27352 | 119.779 | 2.278 | 0.27297 | -0.21126 | -0.001521 | -0.000282 | -0.1761 | -0.0667 | -0.1031
current_sweep | 11 | 1 | 15 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9903 | 2.2795 | 99.9007 | 0.27325 | -0.012186 | 0.27352 | 119.779 | 2.281 | 0.27331 | -0.21126 | 0.001479 | 5.9e-05 | -0.1761 | 0.0649 | 0.0217
current_sweep | 11 | 1 | 16 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9903 | 2.2795 | 99.9007 | 0.27325 | -0.012186 | 0.27352 | 119.779 | 2.279 | 0.27296 | -0.21126 | -0.000521 | -0.000289 | -0.1761 | -0.0229 | -0.1057
current_sweep | 11 | 1 | 17 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9903 | 2.2795 | 99.9007 | 0.27325 | -0.012186 | 0.27352 | 119.779 | 2.28 | 0.27313 | -0.21126 | 0.000479 | -0.00012 | -0.1761 | 0.021 | -0.0438
current_sweep | 11 | 1 | 18 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9903 | 2.2795 | 99.9007 | 0.27325 | -0.012186 | 0.27352 | 119.779 | 2.279 | 0.27305 | -0.21126 | -0.000521 | -0.000202 | -0.1761 | -0.0229 | -0.0738
current_sweep | 11 | 2 | 1 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9988 | 2.2794 | 99.901 | 0.27326 | -0.012159 | 0.27353 | 119.706 | 6.803 | 0.8147 | -0.29285 | 4.5236 | 0.54145 | -0.244 | -0.51534 | -0.61808
current_sweep | 11 | 2 | 2 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9988 | 2.2794 | 99.901 | 0.27326 | -0.012159 | 0.27353 | 119.706 | 2.283 | 0.27333 | -0.29285 | 0.003587 | 7.7e-05 | -0.244 | 0.1573 | 0.0281
current_sweep | 11 | 2 | 3 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9988 | 2.2794 | 99.901 | 0.27326 | -0.012159 | 0.27353 | 119.706 | 2.283 | 0.27333 | -0.29285 | 0.003587 | 7.4e-05 | -0.244 | 0.1573 | 0.027
current_sweep | 11 | 2 | 4 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9988 | 2.2794 | 99.901 | 0.27326 | -0.012159 | 0.27353 | 119.706 | 2.281 | 0.27321 | -0.29285 | 0.001587 | -4.1e-05 | -0.244 | 0.0696 | -0.0151
current_sweep | 11 | 2 | 5 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9988 | 2.2794 | 99.901 | 0.27326 | -0.012159 | 0.27353 | 119.706 | 2.283 | 0.27334 | -0.29285 | 0.003587 | 8.4e-05 | -0.244 | 0.1573 | 0.0306
current_sweep | 11 | 2 | 6 | 1 | 120 | 2.2797 | 0.999 | 60 | 119.9988 | 2.2794 | 99.901 | 0.27326 | -0.012159 | 0.27353 | 119.706 | 2.276 | 0.27252 | -0.29285 | -0.003413 | -0.000741 | -0.244 | -0.1498 | -0.2713
current_sweep | 11 | 2 | 7 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9792 | 2.2799 | 99.9041 | 0.27328 | -0.011971 | 0.27354 | 119.758 | 6.815 | 0.8161 | -0.22116 | 4.5351 | 0.54282 | -0.1843 | -0.36063 | -0.45554
current_sweep | 11 | 2 | 8 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9792 | 2.2799 | 99.9041 | 0.27328 | -0.011971 | 0.27354 | 119.758 | 2.281 | 0.27323 | -0.22116 | 0.001111 | -4.5e-05 | -0.1843 | 0.0487 | -0.0164
current_sweep | 11 | 2 | 9 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9792 | 2.2799 | 99.9041 | 0.27328 | -0.011971 | 0.27354 | 119.758 | 2.281 | 0.27328 | -0.22116 | 0.001111 | 7e-06 | -0.1843 | 0.0487 | 0.0026
current_sweep | 11 | 2 | 10 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9792 | 2.2799 | 99.9041 | 0.27328 | -0.011971 | 0.27354 | 119.758 | 2.28 | 0.2731 | -0.22116 | 0.000111 | -0.000174 | -0.1843 | 0.0049 | -0.0636
current_sweep | 11 | 2 | 11 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9792 | 2.2799 | 99.9041 | 0.27328 | -0.011971 | 0.27354 | 119.758 | 2.279 | 0.27299 | -0.22116 | -0.000889 | -0.000287 | -0.1843 | -0.039 | -0.105
current_sweep | 11 | 2 | 12 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9792 | 2.2799 | 99.9041 | 0.27328 | -0.011971 | 0.27354 | 119.758 | 2.28 | 0.2731 | -0.22116 | 0.000111 | -0.000178 | -0.1843 | 0.0049 | -0.0651
current_sweep | 11 | 2 | 13 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9924 | 2.2795 | 99.9005 | 0.27325 | -0.012193 | 0.27352 | 119.781 | 2.279 | 0.27308 | -0.21137 | -0.000512 | -0.000175 | -0.1762 | -0.0225 | -0.064
current_sweep | 11 | 2 | 14 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9924 | 2.2795 | 99.9005 | 0.27325 | -0.012193 | 0.27352 | 119.781 | 2.279 | 0.27301 | -0.21137 | -0.000512 | -0.00024 | -0.1762 | -0.0225 | -0.0878
current_sweep | 11 | 2 | 15 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9924 | 2.2795 | 99.9005 | 0.27325 | -0.012193 | 0.27352 | 119.781 | 2.282 | 0.27336 | -0.21137 | 0.002488 | 0.000111 | -0.1762 | 0.1091 | 0.0406
current_sweep | 11 | 2 | 16 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9924 | 2.2795 | 99.9005 | 0.27325 | -0.012193 | 0.27352 | 119.781 | 2.279 | 0.273 | -0.21137 | -0.000512 | -0.000251 | -0.1762 | -0.0225 | -0.0918
current_sweep | 11 | 2 | 17 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9924 | 2.2795 | 99.9005 | 0.27325 | -0.012193 | 0.27352 | 119.781 | 2.28 | 0.27317 | -0.21137 | 0.000488 | -8e-05 | -0.1762 | 0.0214 | -0.0293
current_sweep | 11 | 2 | 18 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9924 | 2.2795 | 99.9005 | 0.27325 | -0.012193 | 0.27352 | 119.781 | 2.28 | 0.27311 | -0.21137 | 0.000488 | -0.000145 | -0.1762 | 0.0214 | -0.0531
current_sweep | 11 | 3 | 1 | 1 | 120 | 2.2797 | 0.999 | 60 | 120.0251 | 2.2794 | 99.901 | 0.27332 | -0.012155 | 0.27359 | 119.755 | 6.803 | 0.8147 | -0.27005 | 4.5236 | 0.54139 | -0.225 | -0.51511 | -0.63945
current_sweep | 11 | 3 | 2 | 1 | 120 | 2.2797 | 0.999 | 60 | 120.0251 | 2.2794 | 99.901 | 0.27332 | -0.012155 | 0.27359 | 119.755 | 2.282 | 0.27337 | -0.27005 | 0.002592 | 5.8e-05 | -0.225 | 0.1137 | 0.0211
current_sweep | 11 | 3 | 3 | 1 | 120 | 2.2797 | 0.999 | 60 | 120.0251 | 2.2794 | 99.901 | 0.27332 | -0.012155 | 0.27359 | 119.755 | 2.283 | 0.27344 | -0.27005 | 0.003592 | 0.000122 | -0.225 | 0.1576 | 0.0445
current_sweep | 11 | 3 | 4 | 1 | 120 | 2.2797 | 0.999 | 60 | 120.0251 | 2.2794 | 99.901 | 0.27332 | -0.012155 | 0.27359 | 119.755 | 2.281 | 0.27322 | -0.27005 | 0.001592 | -9.8e-05 | -0.225 | 0.0698 | -0.036
current_sweep | 11 | 3 | 5 | 1 | 120 | 2.2797 | 0.999 | 60 | 120.0251 | 2.2794 | 99.901 | 0.27332 | -0.012155 | 0.27359 | 119.755 | 2.283 | 0.27341 | -0.27005 | 0.003592 | 9.6e-05 | -0.225 | 0.1576 | 0.035
current_sweep | 11 | 3 | 6 | 1 | 120 | 2.2797 | 0.999 | 60 | 120.0251 | 2.2794 | 99.901 | 0.27332 | -0.012155 | 0.27359 | 119.755 | 2.276 | 0.27268 | -0.27005 | -0.003408 | -0.000634 | -0.225 | -0.1495 | -0.2321
current_sweep | 11 | 3 | 7 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9898 | 2.28 | 99.904 | 0.27331 | -0.011982 | 0.27357 | 119.767 | 6.815 | 0.81627 | -0.22282 | 4.535 | 0.54296 | -0.1857 | -0.36399 | -0.4465
current_sweep | 11 | 3 | 8 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9898 | 2.28 | 99.904 | 0.27331 | -0.011982 | 0.27357 | 119.767 | 2.281 | 0.27329 | -0.22282 | 0.001035 | -1.7e-05 | -0.1857 | 0.0454 | -0.0062
current_sweep | 11 | 3 | 9 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9898 | 2.28 | 99.904 | 0.27331 | -0.011982 | 0.27357 | 119.767 | 2.281 | 0.27329 | -0.22282 | 0.001035 | -2e-05 | -0.1857 | 0.0454 | -0.0073
current_sweep | 11 | 3 | 10 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9898 | 2.28 | 99.904 | 0.27331 | -0.011982 | 0.27357 | 119.767 | 2.28 | 0.27316 | -0.22282 | 3.5e-05 | -0.000148 | -0.1857 | 0.0015 | -0.0542
current_sweep | 11 | 3 | 11 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9898 | 2.28 | 99.904 | 0.27331 | -0.011982 | 0.27357 | 119.767 | 2.279 | 0.27305 | -0.22282 | -0.000965 | -0.000258 | -0.1857 | -0.0423 | -0.0944
current_sweep | 11 | 3 | 12 | 2 | 120 | 2.2797 | 0.999 | 60 | 119.9898 | 2.28 | 99.904 | 0.27331 | -0.011982 | 0.27357 | 119.767 | 2.28 | 0.27311 | -0.22282 | 3.5e-05 | -0.000202 | -0.1857 | 0.0015 | -0.0739
current_sweep | 11 | 3 | 13 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9895 | 2.2795 | 99.9005 | 0.27325 | -0.012198 | 0.27352 | 119.784 | 2.28 | 0.2731 | -0.2055 | 0.000461 | -0.000149 | -0.1713 | 0.0202 | -0.0544
current_sweep | 11 | 3 | 14 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9895 | 2.2795 | 99.9005 | 0.27325 | -0.012198 | 0.27352 | 119.784 | 2.279 | 0.273 | -0.2055 | -0.000539 | -0.000246 | -0.1713 | -0.0236 | -0.0899
current_sweep | 11 | 3 | 15 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9895 | 2.2795 | 99.9005 | 0.27325 | -0.012198 | 0.27352 | 119.784 | 2.282 | 0.27338 | -0.2055 | 0.002461 | 0.000135 | -0.1713 | 0.108 | 0.0496
current_sweep | 11 | 3 | 16 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9895 | 2.2795 | 99.9005 | 0.27325 | -0.012198 | 0.27352 | 119.784 | 2.279 | 0.27301 | -0.2055 | -0.000539 | -0.00024 | -0.1713 | -0.0236 | -0.0877
current_sweep | 11 | 3 | 17 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9895 | 2.2795 | 99.9005 | 0.27325 | -0.012198 | 0.27352 | 119.784 | 2.28 | 0.27315 | -0.2055 | 0.000461 | -9.4e-05 | -0.1713 | 0.0202 | -0.0342
current_sweep | 11 | 3 | 18 | 3 | 120 | 2.2797 | 0.999 | 60 | 119.9895 | 2.2795 | 99.9005 | 0.27325 | -0.012198 | 0.27352 | 119.784 | 2.28 | 0.27311 | -0.2055 | 0.000461 | -0.000138 | -0.1713 | 0.0202 | -0.0503
current_sweep | 12 | 1 | 1 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0182 | 3.923 | 99.9011 | 0.47037 | -0.020926 | 0.47084 | 119.745 | 11.741 | 1.406 | -0.27319 | 7.818 | 0.93559 | -0.2276 | -0.23878 | -0.36489
current_sweep | 12 | 1 | 2 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0182 | 3.923 | 99.9011 | 0.47037 | -0.020926 | 0.47084 | 119.745 | 3.928 | 0.47038 | -0.27319 | 0.004966 | 6e-06 | -0.2276 | 0.1266 | 0.0013
current_sweep | 12 | 1 | 3 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0182 | 3.923 | 99.9011 | 0.47037 | -0.020926 | 0.47084 | 119.745 | 3.93 | 0.47049 | -0.27319 | 0.006966 | 0.000124 | -0.2276 | 0.1776 | 0.0264
current_sweep | 12 | 1 | 4 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0182 | 3.923 | 99.9011 | 0.47037 | -0.020926 | 0.47084 | 119.745 | 3.926 | 0.47016 | -0.27319 | 0.002966 | -0.000211 | -0.2276 | 0.0756 | -0.0448
current_sweep | 12 | 1 | 5 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0182 | 3.923 | 99.9011 | 0.47037 | -0.020926 | 0.47084 | 119.745 | 3.929 | 0.47045 | -0.27319 | 0.005966 | 7.9e-05 | -0.2276 | 0.1521 | 0.0169
current_sweep | 12 | 1 | 6 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0182 | 3.923 | 99.9011 | 0.47037 | -0.020926 | 0.47084 | 119.745 | 3.919 | 0.46918 | -0.27319 | -0.004034 | -0.001189 | -0.2276 | -0.1028 | -0.2527
current_sweep | 12 | 1 | 7 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0077 | 3.9243 | 99.9036 | 0.4705 | -0.020667 | 0.47095 | 119.8 | 11.759 | 1.4088 | -0.2077 | 7.8347 | 0.93828 | -0.1731 | -0.11886 | -0.19209
current_sweep | 12 | 1 | 8 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0077 | 3.9243 | 99.9036 | 0.4705 | -0.020667 | 0.47095 | 119.8 | 3.928 | 0.47044 | -0.2077 | 0.003669 | -5.6e-05 | -0.1731 | 0.0935 | -0.0119
current_sweep | 12 | 1 | 9 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0077 | 3.9243 | 99.9036 | 0.4705 | -0.020667 | 0.47095 | 119.8 | 3.927 | 0.47049 | -0.2077 | 0.002669 | -9e-06 | -0.1731 | 0.068 | -0.0019
current_sweep | 12 | 1 | 10 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0077 | 3.9243 | 99.9036 | 0.4705 | -0.020667 | 0.47095 | 119.8 | 3.926 | 0.47036 | -0.2077 | 0.001669 | -0.000139 | -0.1731 | 0.0425 | -0.0295
current_sweep | 12 | 1 | 11 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0077 | 3.9243 | 99.9036 | 0.4705 | -0.020667 | 0.47095 | 119.8 | 3.924 | 0.46995 | -0.2077 | -0.000331 | -0.000549 | -0.1731 | -0.0084 | -0.1166
current_sweep | 12 | 1 | 12 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0077 | 3.9243 | 99.9036 | 0.4705 | -0.020667 | 0.47095 | 119.8 | 3.925 | 0.47018 | -0.2077 | 0.000669 | -0.000314 | -0.1731 | 0.017 | -0.0667
current_sweep | 12 | 1 | 13 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.991 | 3.923 | 99.9004 | 0.47025 | -0.021006 | 0.47072 | 119.78 | 3.923 | 0.47 | -0.211 | 3.4e-05 | -0.000251 | -0.1758 | 0.0009 | -0.0533
current_sweep | 12 | 1 | 14 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.991 | 3.923 | 99.9004 | 0.47025 | -0.021006 | 0.47072 | 119.78 | 3.922 | 0.46986 | -0.211 | -0.000966 | -0.000389 | -0.1758 | -0.0246 | -0.0826
current_sweep | 12 | 1 | 15 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.991 | 3.923 | 99.9004 | 0.47025 | -0.021006 | 0.47072 | 119.78 | 3.928 | 0.47056 | -0.211 | 0.005034 | 0.000308 | -0.1758 | 0.1283 | 0.0656
current_sweep | 12 | 1 | 16 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.991 | 3.923 | 99.9004 | 0.47025 | -0.021006 | 0.47072 | 119.78 | 3.922 | 0.46986 | -0.211 | -0.000966 | -0.000395 | -0.1758 | -0.0246 | -0.0839
current_sweep | 12 | 1 | 17 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.991 | 3.923 | 99.9004 | 0.47025 | -0.021006 | 0.47072 | 119.78 | 3.924 | 0.47011 | -0.211 | 0.001034 | -0.000141 | -0.1758 | 0.0264 | -0.0299
current_sweep | 12 | 1 | 18 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.991 | 3.923 | 99.9004 | 0.47025 | -0.021006 | 0.47072 | 119.78 | 3.924 | 0.47004 | -0.211 | 0.001034 | -0.000208 | -0.1758 | 0.0264 | -0.0441
current_sweep | 12 | 2 | 1 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0176 | 3.923 | 99.9008 | 0.47037 | -0.020937 | 0.47083 | 119.751 | 11.741 | 1.4061 | -0.26656 | 7.818 | 0.93569 | -0.2221 | -0.23909 | -0.35736
current_sweep | 12 | 2 | 2 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0176 | 3.923 | 99.9008 | 0.47037 | -0.020937 | 0.47083 | 119.751 | 3.929 | 0.47039 | -0.26656 | 0.005954 | 2.2e-05 | -0.2221 | 0.1518 | 0.0048
current_sweep | 12 | 2 | 3 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0176 | 3.923 | 99.9008 | 0.47037 | -0.020937 | 0.47083 | 119.751 | 3.929 | 0.47051 | -0.26656 | 0.005954 | 0.00014 | -0.2221 | 0.1518 | 0.0299
current_sweep | 12 | 2 | 4 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0176 | 3.923 | 99.9008 | 0.47037 | -0.020937 | 0.47083 | 119.751 | 3.926 | 0.47019 | -0.26656 | 0.002954 | -0.000177 | -0.2221 | 0.0753 | -0.0375
current_sweep | 12 | 2 | 5 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0176 | 3.923 | 99.9008 | 0.47037 | -0.020937 | 0.47083 | 119.751 | 3.929 | 0.47043 | -0.26656 | 0.005954 | 6.5e-05 | -0.2221 | 0.1518 | 0.0139
current_sweep | 12 | 2 | 6 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0176 | 3.923 | 99.9008 | 0.47037 | -0.020937 | 0.47083 | 119.751 | 3.918 | 0.46919 | -0.26656 | -0.005046 | -0.001182 | -0.2221 | -0.1286 | -0.2512
current_sweep | 12 | 2 | 7 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0108 | 3.9243 | 99.9035 | 0.47051 | -0.020677 | 0.47096 | 119.802 | 11.759 | 1.4088 | -0.20884 | 7.8347 | 0.93828 | -0.174 | -0.1186 | -0.19333
current_sweep | 12 | 2 | 8 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0108 | 3.9243 | 99.9035 | 0.47051 | -0.020677 | 0.47096 | 119.802 | 3.928 | 0.47053 | -0.20884 | 0.003679 | 2.1e-05 | -0.174 | 0.0938 | 0.0045
current_sweep | 12 | 2 | 9 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0108 | 3.9243 | 99.9035 | 0.47051 | -0.020677 | 0.47096 | 119.802 | 3.927 | 0.47054 | -0.20884 | 0.002679 | 2.9e-05 | -0.174 | 0.0683 | 0.0062
current_sweep | 12 | 2 | 10 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0108 | 3.9243 | 99.9035 | 0.47051 | -0.020677 | 0.47096 | 119.802 | 3.926 | 0.47038 | -0.20884 | 0.001679 | -0.000124 | -0.174 | 0.0428 | -0.0263
current_sweep | 12 | 2 | 11 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0108 | 3.9243 | 99.9035 | 0.47051 | -0.020677 | 0.47096 | 119.802 | 3.923 | 0.47003 | -0.20884 | -0.001321 | -0.000481 | -0.174 | -0.0337 | -0.1022
current_sweep | 12 | 2 | 12 | 2 | 120 | 3.9234 | 0.999 | 60 | 120.0108 | 3.9243 | 99.9035 | 0.47051 | -0.020677 | 0.47096 | 119.802 | 3.925 | 0.47025 | -0.20884 | 0.000679 | -0.000262 | -0.174 | 0.0173 | -0.0556
current_sweep | 12 | 2 | 13 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9933 | 3.9229 | 99.9004 | 0.47025 | -0.021005 | 0.47072 | 119.773 | 3.923 | 0.46995 | -0.22032 | 8.5e-05 | -0.00031 | -0.1836 | 0.0022 | -0.0658
current_sweep | 12 | 2 | 14 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9933 | 3.9229 | 99.9004 | 0.47025 | -0.021005 | 0.47072 | 119.773 | 3.922 | 0.46984 | -0.22032 | -0.000915 | -0.000418 | -0.1836 | -0.0233 | -0.0888
current_sweep | 12 | 2 | 15 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9933 | 3.9229 | 99.9004 | 0.47025 | -0.021005 | 0.47072 | 119.773 | 3.928 | 0.47055 | -0.22032 | 0.005085 | 0.000298 | -0.1836 | 0.1296 | 0.0634
current_sweep | 12 | 2 | 16 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9933 | 3.9229 | 99.9004 | 0.47025 | -0.021005 | 0.47072 | 119.773 | 3.922 | 0.4698 | -0.22032 | -0.000915 | -0.000457 | -0.1836 | -0.0233 | -0.0971
current_sweep | 12 | 2 | 17 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9933 | 3.9229 | 99.9004 | 0.47025 | -0.021005 | 0.47072 | 119.773 | 3.924 | 0.47008 | -0.22032 | 0.001085 | -0.000173 | -0.1836 | 0.0277 | -0.0367
current_sweep | 12 | 2 | 18 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9933 | 3.9229 | 99.9004 | 0.47025 | -0.021005 | 0.47072 | 119.773 | 3.923 | 0.47003 | -0.22032 | 8.5e-05 | -0.000221 | -0.1836 | 0.0022 | -0.0469
current_sweep | 12 | 3 | 1 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0142 | 3.923 | 99.9011 | 0.47035 | -0.020919 | 0.47082 | 119.771 | 11.742 | 1.4062 | -0.24321 | 7.819 | 0.9358 | -0.2027 | -0.2303 | -0.34757
current_sweep | 12 | 3 | 2 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0142 | 3.923 | 99.9011 | 0.47035 | -0.020919 | 0.47082 | 119.771 | 3.929 | 0.47066 | -0.24321 | 0.005965 | 0.000303 | -0.2027 | 0.1521 | 0.0643
current_sweep | 12 | 3 | 3 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0142 | 3.923 | 99.9011 | 0.47035 | -0.020919 | 0.47082 | 119.771 | 3.929 | 0.47057 | -0.24321 | 0.005965 | 0.000214 | -0.2027 | 0.1521 | 0.0454
current_sweep | 12 | 3 | 4 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0142 | 3.923 | 99.9011 | 0.47035 | -0.020919 | 0.47082 | 119.771 | 3.926 | 0.47024 | -0.24321 | 0.002965 | -0.000118 | -0.2027 | 0.0756 | -0.0252
current_sweep | 12 | 3 | 5 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0142 | 3.923 | 99.9011 | 0.47035 | -0.020919 | 0.47082 | 119.771 | 3.93 | 0.4707 | -0.24321 | 0.006965 | 0.00035 | -0.2027 | 0.1775 | 0.0743
current_sweep | 12 | 3 | 6 | 1 | 120 | 3.9234 | 0.999 | 60 | 120.0142 | 3.923 | 99.9011 | 0.47035 | -0.020919 | 0.47082 | 119.771 | 3.918 | 0.46923 | -0.24321 | -0.005035 | -0.001127 | -0.2027 | -0.1283 | -0.2397
current_sweep | 12 | 3 | 7 | 2 | 120 | 3.9234 | 0.999 | 60 | 119.9566 | 3.9242 | 99.9038 | 0.47028 | -0.02064 | 0.47073 | 119.748 | 11.76 | 1.4087 | -0.20857 | 7.8358 | 0.93837 | -0.1739 | -0.10697 | -0.15499
current_sweep | 12 | 3 | 8 | 2 | 120 | 3.9234 | 0.999 | 60 | 119.9566 | 3.9242 | 99.9038 | 0.47028 | -0.02064 | 0.47073 | 119.748 | 3.928 | 0.47039 | -0.20857 | 0.003802 | 0.000111 | -0.1739 | 0.0969 | 0.0236
current_sweep | 12 | 3 | 9 | 2 | 120 | 3.9234 | 0.999 | 60 | 119.9566 | 3.9242 | 99.9038 | 0.47028 | -0.02064 | 0.47073 | 119.748 | 3.928 | 0.47057 | -0.20857 | 0.003802 | 0.000295 | -0.1739 | 0.0969 | 0.0627
current_sweep | 12 | 3 | 10 | 2 | 120 | 3.9234 | 0.999 | 60 | 119.9566 | 3.9242 | 99.9038 | 0.47028 | -0.02064 | 0.47073 | 119.748 | 3.926 | 0.47032 | -0.20857 | 0.001802 | 3.9e-05 | -0.1739 | 0.0459 | 0.0083
current_sweep | 12 | 3 | 11 | 2 | 120 | 3.9234 | 0.999 | 60 | 119.9566 | 3.9242 | 99.9038 | 0.47028 | -0.02064 | 0.47073 | 119.748 | 3.923 | 0.46989 | -0.20857 | -0.001198 | -0.000392 | -0.1739 | -0.0305 | -0.0834
current_sweep | 12 | 3 | 12 | 2 | 120 | 3.9234 | 0.999 | 60 | 119.9566 | 3.9242 | 99.9038 | 0.47028 | -0.02064 | 0.47073 | 119.748 | 3.925 | 0.47027 | -0.20857 | 0.000802 | -1.1e-05 | -0.1739 | 0.0204 | -0.0024
current_sweep | 12 | 3 | 13 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9939 | 3.9229 | 99.9003 | 0.47026 | -0.021015 | 0.47073 | 119.766 | 3.923 | 0.46994 | -0.22793 | 7.4e-05 | -0.000318 | -0.1899 | 0.0019 | -0.0676
current_sweep | 12 | 3 | 14 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9939 | 3.9229 | 99.9003 | 0.47026 | -0.021015 | 0.47073 | 119.766 | 3.921 | 0.4697 | -0.22793 | -0.001926 | -0.000556 | -0.1899 | -0.0491 | -0.1182
current_sweep | 12 | 3 | 15 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9939 | 3.9229 | 99.9003 | 0.47026 | -0.021015 | 0.47073 | 119.766 | 3.928 | 0.4705 | -0.22793 | 0.005074 | 0.000245 | -0.1899 | 0.1293 | 0.0521
current_sweep | 12 | 3 | 16 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9939 | 3.9229 | 99.9003 | 0.47026 | -0.021015 | 0.47073 | 119.766 | 3.922 | 0.46979 | -0.22793 | -0.000926 | -0.000464 | -0.1899 | -0.0236 | -0.0986
current_sweep | 12 | 3 | 17 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9939 | 3.9229 | 99.9003 | 0.47026 | -0.021015 | 0.47073 | 119.766 | 3.924 | 0.46996 | -0.22793 | 0.001074 | -0.000293 | -0.1899 | 0.0274 | -0.0623
current_sweep | 12 | 3 | 18 | 3 | 120 | 3.9234 | 0.999 | 60 | 119.9939 | 3.9229 | 99.9003 | 0.47026 | -0.021015 | 0.47073 | 119.766 | 3.923 | 0.47001 | -0.22793 | 7.4e-05 | -0.000245 | -0.1899 | 0.0019 | -0.0521
current_sweep | 13 | 1 | 1 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0089 | 6.7513 | 99.9011 | 0.80941 | -0.035981 | 0.81021 | 119.731 | 20.246 | 2.4247 | -0.27791 | 13.4947 | 1.6153 | -0.2316 | -0.038685 | -0.14368
current_sweep | 13 | 1 | 2 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0089 | 6.7513 | 99.9011 | 0.80941 | -0.035981 | 0.81021 | 119.731 | 6.763 | 0.80989 | -0.27791 | 0.011722 | 0.000476 | -0.2316 | 0.1736 | 0.0588
current_sweep | 13 | 1 | 3 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0089 | 6.7513 | 99.9011 | 0.80941 | -0.035981 | 0.81021 | 119.731 | 6.763 | 0.80978 | -0.27791 | 0.011722 | 0.000366 | -0.2316 | 0.1736 | 0.0452
current_sweep | 13 | 1 | 4 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0089 | 6.7513 | 99.9011 | 0.80941 | -0.035981 | 0.81021 | 119.731 | 6.758 | 0.80939 | -0.27791 | 0.006722 | -1.9e-05 | -0.2316 | 0.0996 | -0.0023
current_sweep | 13 | 1 | 5 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0089 | 6.7513 | 99.9011 | 0.80941 | -0.035981 | 0.81021 | 119.731 | 6.763 | 0.80991 | -0.27791 | 0.011722 | 0.000495 | -0.2316 | 0.1736 | 0.0612
current_sweep | 13 | 1 | 6 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0089 | 6.7513 | 99.9011 | 0.80941 | -0.035981 | 0.81021 | 119.731 | 6.744 | 0.80757 | -0.27791 | -0.007278 | -0.001838 | -0.2316 | -0.1078 | -0.2271
current_sweep | 13 | 1 | 7 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9867 | 6.7539 | 99.9029 | 0.80959 | -0.035652 | 0.81038 | 119.767 | 20.269 | 2.4273 | -0.21967 | 13.5151 | 1.6177 | -0.1831 | 0.036091 | -0.061935
current_sweep | 13 | 1 | 8 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9867 | 6.7539 | 99.9029 | 0.80959 | -0.035652 | 0.81038 | 119.767 | 6.76 | 0.8096 | -0.21967 | 0.006104 | 5e-06 | -0.1831 | 0.0904 | 0.0006
current_sweep | 13 | 1 | 9 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9867 | 6.7539 | 99.9029 | 0.80959 | -0.035652 | 0.81038 | 119.767 | 6.761 | 0.80975 | -0.21967 | 0.007104 | 0.000162 | -0.1831 | 0.1052 | 0.02
current_sweep | 13 | 1 | 10 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9867 | 6.7539 | 99.9029 | 0.80959 | -0.035652 | 0.81038 | 119.767 | 6.756 | 0.8091 | -0.21967 | 0.002104 | -0.000492 | -0.1831 | 0.0312 | -0.0607
current_sweep | 13 | 1 | 11 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9867 | 6.7539 | 99.9029 | 0.80959 | -0.035652 | 0.81038 | 119.767 | 6.753 | 0.80872 | -0.21967 | -0.000896 | -0.00087 | -0.1831 | -0.0133 | -0.1074
current_sweep | 13 | 1 | 12 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9867 | 6.7539 | 99.9029 | 0.80959 | -0.035652 | 0.81038 | 119.767 | 6.756 | 0.80926 | -0.21967 | 0.002104 | -0.000333 | -0.1831 | 0.0312 | -0.0411
current_sweep | 13 | 1 | 13 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.9916 | 6.7517 | 99.9006 | 0.80934 | -0.036114 | 0.81015 | 119.779 | 6.753 | 0.80888 | -0.2126 | 0.001313 | -0.000457 | -0.1772 | 0.0194 | -0.0565
current_sweep | 13 | 1 | 14 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.9916 | 6.7517 | 99.9006 | 0.80934 | -0.036114 | 0.81015 | 119.779 | 6.751 | 0.8087 | -0.2126 | -0.000687 | -0.000638 | -0.1772 | -0.0102 | -0.0789
current_sweep | 13 | 1 | 15 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.9916 | 6.7517 | 99.9006 | 0.80934 | -0.036114 | 0.81015 | 119.779 | 6.761 | 0.8099 | -0.2126 | 0.009313 | 0.000558 | -0.1772 | 0.1379 | 0.0689
current_sweep | 13 | 1 | 16 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.9916 | 6.7517 | 99.9006 | 0.80934 | -0.036114 | 0.81015 | 119.779 | 6.751 | 0.80864 | -0.2126 | -0.000687 | -0.000703 | -0.1772 | -0.0102 | -0.0869
current_sweep | 13 | 1 | 17 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.9916 | 6.7517 | 99.9006 | 0.80934 | -0.036114 | 0.81015 | 119.779 | 6.754 | 0.80913 | -0.2126 | 0.002313 | -0.000205 | -0.1772 | 0.0343 | -0.0254
current_sweep | 13 | 1 | 18 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.9916 | 6.7517 | 99.9006 | 0.80934 | -0.036114 | 0.81015 | 119.779 | 6.753 | 0.80895 | -0.2126 | 0.001313 | -0.000389 | -0.1772 | 0.0194 | -0.0481
current_sweep | 13 | 2 | 1 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.012 | 6.7514 | 99.9009 | 0.80945 | -0.036041 | 0.81025 | 119.746 | 20.245 | 2.4247 | -0.26601 | 13.4936 | 1.6152 | -0.2217 | -0.045401 | -0.15075
current_sweep | 13 | 2 | 2 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.012 | 6.7514 | 99.9009 | 0.80945 | -0.036041 | 0.81025 | 119.746 | 6.763 | 0.80986 | -0.26601 | 0.011601 | 0.000413 | -0.2217 | 0.1718 | 0.051
current_sweep | 13 | 2 | 3 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.012 | 6.7514 | 99.9009 | 0.80945 | -0.036041 | 0.81025 | 119.746 | 6.763 | 0.80991 | -0.26601 | 0.011601 | 0.000468 | -0.2217 | 0.1718 | 0.0578
current_sweep | 13 | 2 | 4 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.012 | 6.7514 | 99.9009 | 0.80945 | -0.036041 | 0.81025 | 119.746 | 6.758 | 0.80939 | -0.26601 | 0.006601 | -5.3e-05 | -0.2217 | 0.0978 | -0.0065
current_sweep | 13 | 2 | 5 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.012 | 6.7514 | 99.9009 | 0.80945 | -0.036041 | 0.81025 | 119.746 | 6.763 | 0.80988 | -0.26601 | 0.011601 | 0.000436 | -0.2217 | 0.1718 | 0.0539
current_sweep | 13 | 2 | 6 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.012 | 6.7514 | 99.9009 | 0.80945 | -0.036041 | 0.81025 | 119.746 | 6.745 | 0.80771 | -0.26601 | -0.006399 | -0.001739 | -0.2217 | -0.0948 | -0.2148
current_sweep | 13 | 2 | 7 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9853 | 6.7539 | 99.9029 | 0.80958 | -0.035653 | 0.81036 | 119.773 | 20.271 | 2.428 | -0.21231 | 13.5171 | 1.6184 | -0.1769 | 0.046379 | -0.031554
current_sweep | 13 | 2 | 8 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9853 | 6.7539 | 99.9029 | 0.80958 | -0.035653 | 0.81036 | 119.773 | 6.761 | 0.80966 | -0.21231 | 0.007132 | 7.9e-05 | -0.1769 | 0.1056 | 0.0097
current_sweep | 13 | 2 | 9 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9853 | 6.7539 | 99.9029 | 0.80958 | -0.035653 | 0.81036 | 119.773 | 6.76 | 0.80973 | -0.21231 | 0.006132 | 0.000151 | -0.1769 | 0.0908 | 0.0186
current_sweep | 13 | 2 | 10 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9853 | 6.7539 | 99.9029 | 0.80958 | -0.035653 | 0.81036 | 119.773 | 6.757 | 0.80933 | -0.21231 | 0.003132 | -0.000249 | -0.1769 | 0.0464 | -0.0308
current_sweep | 13 | 2 | 11 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9853 | 6.7539 | 99.9029 | 0.80958 | -0.035653 | 0.81036 | 119.773 | 6.753 | 0.80876 | -0.21231 | -0.000868 | -0.00082 | -0.1769 | -0.0128 | -0.1013
current_sweep | 13 | 2 | 12 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.9853 | 6.7539 | 99.9029 | 0.80958 | -0.035653 | 0.81036 | 119.773 | 6.756 | 0.80923 | -0.21231 | 0.002132 | -0.000347 | -0.1769 | 0.0316 | -0.0429
current_sweep | 13 | 2 | 13 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.99 | 6.7517 | 99.9006 | 0.80933 | -0.036115 | 0.81013 | 119.783 | 6.752 | 0.80877 | -0.20704 | 0.000325 | -0.000557 | -0.1725 | 0.0048 | -0.0688
current_sweep | 13 | 2 | 14 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.99 | 6.7517 | 99.9006 | 0.80933 | -0.036115 | 0.81013 | 119.783 | 6.751 | 0.80863 | -0.20704 | -0.000675 | -0.000702 | -0.1725 | -0.01 | -0.0868
current_sweep | 13 | 2 | 15 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.99 | 6.7517 | 99.9006 | 0.80933 | -0.036115 | 0.81013 | 119.783 | 6.761 | 0.80995 | -0.20704 | 0.009325 | 0.000627 | -0.1725 | 0.1381 | 0.0775
current_sweep | 13 | 2 | 16 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.99 | 6.7517 | 99.9006 | 0.80933 | -0.036115 | 0.81013 | 119.783 | 6.75 | 0.80853 | -0.20704 | -0.001675 | -0.0008 | -0.1725 | -0.0248 | -0.0989
current_sweep | 13 | 2 | 17 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.99 | 6.7517 | 99.9006 | 0.80933 | -0.036115 | 0.81013 | 119.783 | 6.754 | 0.80905 | -0.20704 | 0.002325 | -0.000274 | -0.1725 | 0.0344 | -0.0339
current_sweep | 13 | 2 | 18 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.99 | 6.7517 | 99.9006 | 0.80933 | -0.036115 | 0.81013 | 119.783 | 6.753 | 0.809 | -0.20704 | 0.001325 | -0.000328 | -0.1725 | 0.0196 | -0.0405
current_sweep | 13 | 3 | 1 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0156 | 6.7513 | 99.9013 | 0.80944 | -0.035993 | 0.81024 | 119.735 | 20.242 | 2.4237 | -0.28059 | 13.4907 | 1.6143 | -0.2338 | -0.058173 | -0.18954
current_sweep | 13 | 3 | 2 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0156 | 6.7513 | 99.9013 | 0.80944 | -0.035993 | 0.81024 | 119.735 | 6.763 | 0.80989 | -0.28059 | 0.011739 | 0.000453 | -0.2338 | 0.1739 | 0.056
current_sweep | 13 | 3 | 3 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0156 | 6.7513 | 99.9013 | 0.80944 | -0.035993 | 0.81024 | 119.735 | 6.763 | 0.80978 | -0.28059 | 0.011739 | 0.000348 | -0.2338 | 0.1739 | 0.043
current_sweep | 13 | 3 | 4 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0156 | 6.7513 | 99.9013 | 0.80944 | -0.035993 | 0.81024 | 119.735 | 6.757 | 0.80907 | -0.28059 | 0.005739 | -0.000365 | -0.2338 | 0.085 | -0.045
current_sweep | 13 | 3 | 5 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0156 | 6.7513 | 99.9013 | 0.80944 | -0.035993 | 0.81024 | 119.735 | 6.763 | 0.80992 | -0.28059 | 0.011739 | 0.000483 | -0.2338 | 0.1739 | 0.0597
current_sweep | 13 | 3 | 6 | 1 | 120 | 6.7523 | 0.999 | 60 | 120.0156 | 6.7513 | 99.9013 | 0.80944 | -0.035993 | 0.81024 | 119.735 | 6.745 | 0.80756 | -0.28059 | -0.006261 | -0.001878 | -0.2338 | -0.0927 | -0.232
current_sweep | 13 | 3 | 7 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.984 | 6.7538 | 99.9026 | 0.80956 | -0.03571 | 0.81035 | 119.777 | 20.273 | 2.4283 | -0.20696 | 13.5192 | 1.6187 | -0.1725 | 0.056561 | -0.018037
current_sweep | 13 | 3 | 8 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.984 | 6.7538 | 99.9026 | 0.80956 | -0.03571 | 0.81035 | 119.777 | 6.76 | 0.80969 | -0.20696 | 0.006153 | 0.000122 | -0.1725 | 0.0911 | 0.015
current_sweep | 13 | 3 | 9 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.984 | 6.7538 | 99.9026 | 0.80956 | -0.03571 | 0.81035 | 119.777 | 6.76 | 0.80971 | -0.20696 | 0.006153 | 0.00015 | -0.1725 | 0.0911 | 0.0185
current_sweep | 13 | 3 | 10 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.984 | 6.7538 | 99.9026 | 0.80956 | -0.03571 | 0.81035 | 119.777 | 6.757 | 0.80942 | -0.20696 | 0.003153 | -0.000144 | -0.1725 | 0.0467 | -0.0178
current_sweep | 13 | 3 | 11 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.984 | 6.7538 | 99.9026 | 0.80956 | -0.03571 | 0.81035 | 119.777 | 6.753 | 0.80879 | -0.20696 | -0.000847 | -0.000777 | -0.1725 | -0.0125 | -0.096
current_sweep | 13 | 3 | 12 | 2 | 120 | 6.7523 | 0.999 | 60 | 119.984 | 6.7538 | 99.9026 | 0.80956 | -0.03571 | 0.81035 | 119.777 | 6.756 | 0.80924 | -0.20696 | 0.002153 | -0.000328 | -0.1725 | 0.0319 | -0.0406
current_sweep | 13 | 3 | 13 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.994 | 6.7517 | 99.9007 | 0.80936 | -0.036089 | 0.81016 | 119.779 | 6.753 | 0.80893 | -0.21499 | 0.001308 | -0.000423 | -0.1792 | 0.0194 | -0.0523
current_sweep | 13 | 3 | 14 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.994 | 6.7517 | 99.9007 | 0.80936 | -0.036089 | 0.81016 | 119.779 | 6.751 | 0.80867 | -0.21499 | -0.000692 | -0.000688 | -0.1792 | -0.0102 | -0.085
current_sweep | 13 | 3 | 15 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.994 | 6.7517 | 99.9007 | 0.80936 | -0.036089 | 0.81016 | 119.779 | 6.761 | 0.80995 | -0.21499 | 0.009308 | 0.000594 | -0.1792 | 0.1379 | 0.0734
current_sweep | 13 | 3 | 16 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.994 | 6.7517 | 99.9007 | 0.80936 | -0.036089 | 0.81016 | 119.779 | 6.751 | 0.80869 | -0.21499 | -0.000692 | -0.000672 | -0.1792 | -0.0102 | -0.083
current_sweep | 13 | 3 | 17 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.994 | 6.7517 | 99.9007 | 0.80936 | -0.036089 | 0.81016 | 119.779 | 6.754 | 0.80911 | -0.21499 | 0.002308 | -0.000252 | -0.1792 | 0.0342 | -0.0311
current_sweep | 13 | 3 | 18 | 3 | 120 | 6.7523 | 0.999 | 60 | 119.994 | 6.7517 | 99.9007 | 0.80936 | -0.036089 | 0.81016 | 119.779 | 6.753 | 0.80901 | -0.21499 | 0.001308 | -0.000346 | -0.1792 | 0.0194 | -0.0427
current_sweep | 14 | 1 | 1 | 1 | 120 | 11.621 | 0.999 | 60 | 119.9989 | 11.6276 | 99.9051 | 1.394 | -0.060708 | 1.3953 | 119.745 | 34.905 | 4.1793 | -0.25389 | 23.2774 | 2.7854 | -0.2116 | 0.064038 | -0.061728
current_sweep | 14 | 1 | 2 | 1 | 120 | 11.621 | 0.999 | 60 | 119.9989 | 11.6276 | 99.9051 | 1.394 | -0.060708 | 1.3953 | 119.745 | 11.648 | 1.3948 | -0.25389 | 0.020446 | 0.000819 | -0.2116 | 0.1758 | 0.0587
current_sweep | 14 | 1 | 3 | 1 | 120 | 11.621 | 0.999 | 60 | 119.9989 | 11.6276 | 99.9051 | 1.394 | -0.060708 | 1.3953 | 119.745 | 11.647 | 1.3945 | -0.25389 | 0.019446 | 0.000484 | -0.2116 | 0.1672 | 0.0347
current_sweep | 14 | 1 | 4 | 1 | 120 | 11.621 | 0.999 | 60 | 119.9989 | 11.6276 | 99.9051 | 1.394 | -0.060708 | 1.3953 | 119.745 | 11.637 | 1.3934 | -0.25389 | 0.009446 | -0.000549 | -0.2116 | 0.0812 | -0.0394
current_sweep | 14 | 1 | 5 | 1 | 120 | 11.621 | 0.999 | 60 | 119.9989 | 11.6276 | 99.9051 | 1.394 | -0.060708 | 1.3953 | 119.745 | 11.648 | 1.3948 | -0.25389 | 0.020446 | 0.000879 | -0.2116 | 0.1758 | 0.063
current_sweep | 14 | 1 | 6 | 1 | 120 | 11.621 | 0.999 | 60 | 119.9989 | 11.6276 | 99.9051 | 1.394 | -0.060708 | 1.3953 | 119.745 | 11.617 | 1.3909 | -0.25389 | -0.010554 | -0.003111 | -0.2116 | -0.0908 | -0.2232
current_sweep | 14 | 1 | 7 | 2 | 120 | 11.621 | 0.999 | 60 | 119.986 | 11.6195 | 99.9036 | 1.3928 | -0.061111 | 1.3942 | 119.773 | 34.926 | 4.1833 | -0.21302 | 23.3065 | 2.7905 | -0.1775 | 0.1933 | 0.11444
current_sweep | 14 | 1 | 8 | 2 | 120 | 11.621 | 0.999 | 60 | 119.986 | 11.6195 | 99.9036 | 1.3928 | -0.061111 | 1.3942 | 119.773 | 11.634 | 1.3935 | -0.21302 | 0.014461 | 0.000616 | -0.1775 | 0.1245 | 0.0442
current_sweep | 14 | 1 | 9 | 2 | 120 | 11.621 | 0.999 | 60 | 119.986 | 11.6195 | 99.9036 | 1.3928 | -0.061111 | 1.3942 | 119.773 | 11.632 | 1.3931 | -0.21302 | 0.012461 | 0.000309 | -0.1775 | 0.1072 | 0.0222
current_sweep | 14 | 1 | 10 | 2 | 120 | 11.621 | 0.999 | 60 | 119.986 | 11.6195 | 99.9036 | 1.3928 | -0.061111 | 1.3942 | 119.773 | 11.626 | 1.3926 | -0.21302 | 0.006461 | -0.000251 | -0.1775 | 0.0556 | -0.018
current_sweep | 14 | 1 | 11 | 2 | 120 | 11.621 | 0.999 | 60 | 119.986 | 11.6195 | 99.9036 | 1.3928 | -0.061111 | 1.3942 | 119.773 | 11.621 | 1.3919 | -0.21302 | 0.001461 | -0.000898 | -0.1775 | 0.0126 | -0.0645
current_sweep | 14 | 1 | 12 | 2 | 120 | 11.621 | 0.999 | 60 | 119.986 | 11.6195 | 99.9036 | 1.3928 | -0.061111 | 1.3942 | 119.773 | 11.625 | 1.3924 | -0.21302 | 0.005461 | -0.000486 | -0.1775 | 0.047 | -0.0349
current_sweep | 14 | 1 | 13 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9909 | 11.6177 | 99.9054 | 1.3927 | -0.060613 | 1.394 | 119.764 | 11.62 | 1.3918 | -0.22687 | 0.002297 | -0.000938 | -0.1891 | 0.0198 | -0.0673
current_sweep | 14 | 1 | 14 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9909 | 11.6177 | 99.9054 | 1.3927 | -0.060613 | 1.394 | 119.764 | 11.616 | 1.3912 | -0.22687 | -0.001703 | -0.00146 | -0.1891 | -0.0147 | -0.1048
current_sweep | 14 | 1 | 15 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9909 | 11.6177 | 99.9054 | 1.3927 | -0.060613 | 1.394 | 119.764 | 11.635 | 1.3937 | -0.22687 | 0.017297 | 0.000981 | -0.1891 | 0.1489 | 0.0705
current_sweep | 14 | 1 | 16 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9909 | 11.6177 | 99.9054 | 1.3927 | -0.060613 | 1.394 | 119.764 | 11.616 | 1.3913 | -0.22687 | -0.001703 | -0.001355 | -0.1891 | -0.0147 | -0.0973
current_sweep | 14 | 1 | 17 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9909 | 11.6177 | 99.9054 | 1.3927 | -0.060613 | 1.394 | 119.764 | 11.622 | 1.3919 | -0.22687 | 0.004297 | -0.000766 | -0.1891 | 0.037 | -0.055
current_sweep | 14 | 1 | 18 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9909 | 11.6177 | 99.9054 | 1.3927 | -0.060613 | 1.394 | 119.764 | 11.62 | 1.392 | -0.22687 | 0.002297 | -0.000738 | -0.1891 | 0.0198 | -0.053
current_sweep | 14 | 2 | 1 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0083 | 11.6279 | 99.9047 | 1.3941 | -0.060876 | 1.3954 | 119.747 | 34.909 | 4.1791 | -0.26135 | 23.2811 | 2.785 | -0.2178 | 0.072173 | -0.07829
current_sweep | 14 | 2 | 2 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0083 | 11.6279 | 99.9047 | 1.3941 | -0.060876 | 1.3954 | 119.747 | 11.648 | 1.3949 | -0.26135 | 0.020059 | 0.000789 | -0.2178 | 0.1725 | 0.0566
current_sweep | 14 | 2 | 3 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0083 | 11.6279 | 99.9047 | 1.3941 | -0.060876 | 1.3954 | 119.747 | 11.648 | 1.3947 | -0.26135 | 0.020059 | 0.000553 | -0.2178 | 0.1725 | 0.0396
current_sweep | 14 | 2 | 4 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0083 | 11.6279 | 99.9047 | 1.3941 | -0.060876 | 1.3954 | 119.747 | 11.639 | 1.3934 | -0.26135 | 0.011059 | -0.000768 | -0.2178 | 0.0951 | -0.0551
current_sweep | 14 | 2 | 5 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0083 | 11.6279 | 99.9047 | 1.3941 | -0.060876 | 1.3954 | 119.747 | 11.649 | 1.395 | -0.26135 | 0.021059 | 0.000857 | -0.2178 | 0.1811 | 0.0614
current_sweep | 14 | 2 | 6 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0083 | 11.6279 | 99.9047 | 1.3941 | -0.060876 | 1.3954 | 119.747 | 11.617 | 1.391 | -0.26135 | -0.010941 | -0.003088 | -0.2178 | -0.0941 | -0.2215
current_sweep | 14 | 2 | 7 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9813 | 11.6203 | 99.9034 | 1.3929 | -0.061178 | 1.3942 | 119.76 | 34.927 | 4.1834 | -0.2213 | 23.3067 | 2.7905 | -0.1844 | 0.18948 | 0.11423
current_sweep | 14 | 2 | 8 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9813 | 11.6203 | 99.9034 | 1.3929 | -0.061178 | 1.3942 | 119.76 | 11.634 | 1.3933 | -0.2213 | 0.013684 | 0.000434 | -0.1844 | 0.1178 | 0.0312
current_sweep | 14 | 2 | 9 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9813 | 11.6203 | 99.9034 | 1.3929 | -0.061178 | 1.3942 | 119.76 | 11.632 | 1.3934 | -0.2213 | 0.011684 | 0.00054 | -0.1844 | 0.1006 | 0.0388
current_sweep | 14 | 2 | 10 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9813 | 11.6203 | 99.9034 | 1.3929 | -0.061178 | 1.3942 | 119.76 | 11.626 | 1.3926 | -0.2213 | 0.005684 | -0.000262 | -0.1844 | 0.0489 | -0.0188
current_sweep | 14 | 2 | 11 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9813 | 11.6203 | 99.9034 | 1.3929 | -0.061178 | 1.3942 | 119.76 | 11.621 | 1.3918 | -0.2213 | 0.000684 | -0.001078 | -0.1844 | 0.0059 | -0.0774
current_sweep | 14 | 2 | 12 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9813 | 11.6203 | 99.9034 | 1.3929 | -0.061178 | 1.3942 | 119.76 | 11.626 | 1.3926 | -0.2213 | 0.005684 | -0.000253 | -0.1844 | 0.0489 | -0.0181
current_sweep | 14 | 2 | 13 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9926 | 11.6168 | 99.9048 | 1.3926 | -0.060781 | 1.3939 | 119.767 | 11.617 | 1.3914 | -0.22562 | 0.000241 | -0.001224 | -0.188 | 0.0021 | -0.0879
current_sweep | 14 | 2 | 14 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9926 | 11.6168 | 99.9048 | 1.3926 | -0.060781 | 1.3939 | 119.767 | 11.615 | 1.3912 | -0.22562 | -0.001759 | -0.001435 | -0.188 | -0.0151 | -0.103
current_sweep | 14 | 2 | 15 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9926 | 11.6168 | 99.9048 | 1.3926 | -0.060781 | 1.3939 | 119.767 | 11.633 | 1.3934 | -0.22562 | 0.016241 | 0.000763 | -0.188 | 0.1398 | 0.0548
current_sweep | 14 | 2 | 16 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9926 | 11.6168 | 99.9048 | 1.3926 | -0.060781 | 1.3939 | 119.767 | 11.614 | 1.391 | -0.22562 | -0.002759 | -0.001638 | -0.188 | -0.0238 | -0.1176
current_sweep | 14 | 2 | 17 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9926 | 11.6168 | 99.9048 | 1.3926 | -0.060781 | 1.3939 | 119.767 | 11.621 | 1.3919 | -0.22562 | 0.004241 | -0.000744 | -0.188 | 0.0365 | -0.0534
current_sweep | 14 | 2 | 18 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9926 | 11.6168 | 99.9048 | 1.3926 | -0.060781 | 1.3939 | 119.767 | 11.618 | 1.3916 | -0.22562 | 0.001241 | -0.000961 | -0.188 | 0.0107 | -0.069
current_sweep | 14 | 3 | 1 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0259 | 11.628 | 99.9055 | 1.3943 | -0.060603 | 1.3957 | 119.737 | 34.911 | 4.182 | -0.28888 | 23.283 | 2.7877 | -0.2407 | 0.077315 | -0.023597
current_sweep | 14 | 3 | 2 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0259 | 11.628 | 99.9055 | 1.3943 | -0.060603 | 1.3957 | 119.737 | 11.648 | 1.3949 | -0.28888 | 0.01999 | 0.000566 | -0.2407 | 0.1719 | 0.0406
current_sweep | 14 | 3 | 3 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0259 | 11.628 | 99.9055 | 1.3943 | -0.060603 | 1.3957 | 119.737 | 11.649 | 1.3949 | -0.28888 | 0.02099 | 0.000571 | -0.2407 | 0.1805 | 0.0409
current_sweep | 14 | 3 | 4 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0259 | 11.628 | 99.9055 | 1.3943 | -0.060603 | 1.3957 | 119.737 | 11.639 | 1.3943 | -0.28888 | 0.01099 | -2.8e-05 | -0.2407 | 0.0945 | -0.002
current_sweep | 14 | 3 | 5 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0259 | 11.628 | 99.9055 | 1.3943 | -0.060603 | 1.3957 | 119.737 | 11.649 | 1.395 | -0.28888 | 0.02099 | 0.000622 | -0.2407 | 0.1805 | 0.0446
current_sweep | 14 | 3 | 6 | 1 | 120 | 11.621 | 0.999 | 60 | 120.0259 | 11.628 | 99.9055 | 1.3943 | -0.060603 | 1.3957 | 119.737 | 11.619 | 1.3913 | -0.28888 | -0.00901 | -0.003072 | -0.2407 | -0.0775 | -0.2203
current_sweep | 14 | 3 | 7 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9803 | 11.6197 | 99.9032 | 1.3928 | -0.061238 | 1.3941 | 119.773 | 34.926 | 4.1828 | -0.20734 | 23.3063 | 2.79 | -0.1728 | 0.19223 | 0.10727
current_sweep | 14 | 3 | 8 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9803 | 11.6197 | 99.9032 | 1.3928 | -0.061238 | 1.3941 | 119.773 | 11.633 | 1.3934 | -0.20734 | 0.013337 | 0.000599 | -0.1728 | 0.1148 | 0.043
current_sweep | 14 | 3 | 9 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9803 | 11.6197 | 99.9032 | 1.3928 | -0.061238 | 1.3941 | 119.773 | 11.631 | 1.3932 | -0.20734 | 0.011337 | 0.000392 | -0.1728 | 0.0976 | 0.0281
current_sweep | 14 | 3 | 10 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9803 | 11.6197 | 99.9032 | 1.3928 | -0.061238 | 1.3941 | 119.773 | 11.626 | 1.3924 | -0.20734 | 0.006337 | -0.000355 | -0.1728 | 0.0545 | -0.0255
current_sweep | 14 | 3 | 11 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9803 | 11.6197 | 99.9032 | 1.3928 | -0.061238 | 1.3941 | 119.773 | 11.621 | 1.3919 | -0.20734 | 0.001337 | -0.000898 | -0.1728 | 0.0115 | -0.0645
current_sweep | 14 | 3 | 12 | 2 | 120 | 11.621 | 0.999 | 60 | 119.9803 | 11.6197 | 99.9032 | 1.3928 | -0.061238 | 1.3941 | 119.773 | 11.625 | 1.3924 | -0.20734 | 0.005337 | -0.000403 | -0.1728 | 0.0459 | -0.0289
current_sweep | 14 | 3 | 13 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9894 | 11.6172 | 99.9052 | 1.3926 | -0.060657 | 1.3939 | 119.782 | 11.62 | 1.3917 | -0.20739 | 0.002837 | -0.000885 | -0.1728 | 0.0244 | -0.0635
current_sweep | 14 | 3 | 14 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9894 | 11.6172 | 99.9052 | 1.3926 | -0.060657 | 1.3939 | 119.782 | 11.615 | 1.3913 | -0.20739 | -0.002163 | -0.001319 | -0.1728 | -0.0186 | -0.0947
current_sweep | 14 | 3 | 15 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9894 | 11.6172 | 99.9052 | 1.3926 | -0.060657 | 1.3939 | 119.782 | 11.636 | 1.3938 | -0.20739 | 0.018837 | 0.001191 | -0.1728 | 0.1622 | 0.0855
current_sweep | 14 | 3 | 16 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9894 | 11.6172 | 99.9052 | 1.3926 | -0.060657 | 1.3939 | 119.782 | 11.616 | 1.3913 | -0.20739 | -0.001163 | -0.001291 | -0.1728 | -0.01 | -0.0927
current_sweep | 14 | 3 | 17 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9894 | 11.6172 | 99.9052 | 1.3926 | -0.060657 | 1.3939 | 119.782 | 11.621 | 1.392 | -0.20739 | 0.003837 | -0.000646 | -0.1728 | 0.033 | -0.0464
current_sweep | 14 | 3 | 18 | 3 | 120 | 11.621 | 0.999 | 60 | 119.9894 | 11.6172 | 99.9052 | 1.3926 | -0.060657 | 1.3939 | 119.782 | 11.621 | 1.3921 | -0.20739 | 0.003837 | -0.000528 | -0.1728 | 0.033 | -0.0379
current_sweep | 15 | 1 | 1 | 1 | 120 | 20 | 0.999 | 60 | 120.0134 | 19.9777 | 99.9021 | 2.3952 | -0.10461 | 2.3976 | 119.741 | 60.045 | 7.19 | -0.27241 | 40.0673 | 4.7947 | -0.227 | 0.18665 | 0.058796
current_sweep | 15 | 1 | 2 | 1 | 120 | 20 | 0.999 | 60 | 120.0134 | 19.9777 | 99.9021 | 2.3952 | -0.10461 | 2.3976 | 119.741 | 20.011 | 2.3962 | -0.27241 | 0.033289 | 0.000938 | -0.227 | 0.1666 | 0.0392
current_sweep | 15 | 1 | 3 | 1 | 120 | 20 | 0.999 | 60 | 120.0134 | 19.9777 | 99.9021 | 2.3952 | -0.10461 | 2.3976 | 119.741 | 20.013 | 2.3958 | -0.27241 | 0.035289 | 0.000529 | -0.227 | 0.1766 | 0.0221
current_sweep | 15 | 1 | 4 | 1 | 120 | 20 | 0.999 | 60 | 120.0134 | 19.9777 | 99.9021 | 2.3952 | -0.10461 | 2.3976 | 119.741 | 20 | 2.3948 | -0.27241 | 0.022289 | -0.000398 | -0.227 | 0.1116 | -0.0166
current_sweep | 15 | 1 | 5 | 1 | 120 | 20 | 0.999 | 60 | 120.0134 | 19.9777 | 99.9021 | 2.3952 | -0.10461 | 2.3976 | 119.741 | 20.012 | 2.3963 | -0.27241 | 0.034289 | 0.00104 | -0.227 | 0.1716 | 0.0434
current_sweep | 15 | 1 | 6 | 1 | 120 | 20 | 0.999 | 60 | 120.0134 | 19.9777 | 99.9021 | 2.3952 | -0.10461 | 2.3976 | 119.741 | 19.964 | 2.3899 | -0.27241 | -0.013711 | -0.005383 | -0.227 | -0.0686 | -0.2247
current_sweep | 15 | 1 | 7 | 2 | 120 | 20 | 0.999 | 60 | 119.9865 | 20.0095 | 99.9052 | 2.3986 | -0.1044 | 2.4009 | 119.757 | 60.213 | 7.211 | -0.22945 | 40.2035 | 4.8124 | -0.1912 | 0.30723 | 0.21126
current_sweep | 15 | 1 | 8 | 2 | 120 | 20 | 0.999 | 60 | 119.9865 | 20.0095 | 99.9052 | 2.3986 | -0.1044 | 2.4009 | 119.757 | 20.034 | 2.3995 | -0.22945 | 0.024475 | 0.000888 | -0.1912 | 0.1223 | 0.037
current_sweep | 15 | 1 | 9 | 2 | 120 | 20 | 0.999 | 60 | 119.9865 | 20.0095 | 99.9052 | 2.3986 | -0.1044 | 2.4009 | 119.757 | 20.033 | 2.3995 | -0.22945 | 0.023475 | 0.000928 | -0.1912 | 0.1173 | 0.0387
current_sweep | 15 | 1 | 10 | 2 | 120 | 20 | 0.999 | 60 | 119.9865 | 20.0095 | 99.9052 | 2.3986 | -0.1044 | 2.4009 | 119.757 | 20.02 | 2.3977 | -0.22945 | 0.010475 | -0.000944 | -0.1912 | 0.0523 | -0.0394
current_sweep | 15 | 1 | 11 | 2 | 120 | 20 | 0.999 | 60 | 119.9865 | 20.0095 | 99.9052 | 2.3986 | -0.1044 | 2.4009 | 119.757 | 20.013 | 2.397 | -0.22945 | 0.003475 | -0.001633 | -0.1912 | 0.0174 | -0.0681
current_sweep | 15 | 1 | 12 | 2 | 120 | 20 | 0.999 | 60 | 119.9865 | 20.0095 | 99.9052 | 2.3986 | -0.1044 | 2.4009 | 119.757 | 20.022 | 2.3983 | -0.22945 | 0.012475 | -0.000345 | -0.1912 | 0.0623 | -0.0144
current_sweep | 15 | 1 | 13 | 3 | 120 | 20 | 0.999 | 60 | 119.9897 | 20.0055 | 99.905 | 2.3982 | -0.1046 | 2.4005 | 119.764 | 20.011 | 2.3966 | -0.2257 | 0.005484 | -0.00154 | -0.1881 | 0.0274 | -0.0642
current_sweep | 15 | 1 | 14 | 3 | 120 | 20 | 0.999 | 60 | 119.9897 | 20.0055 | 99.905 | 2.3982 | -0.1046 | 2.4005 | 119.764 | 20.004 | 2.3959 | -0.2257 | -0.001516 | -0.002244 | -0.1881 | -0.0076 | -0.0936
current_sweep | 15 | 1 | 15 | 3 | 120 | 20 | 0.999 | 60 | 119.9897 | 20.0055 | 99.905 | 2.3982 | -0.1046 | 2.4005 | 119.764 | 20.035 | 2.3995 | -0.2257 | 0.029484 | 0.00128 | -0.1881 | 0.1474 | 0.0534
current_sweep | 15 | 1 | 16 | 3 | 120 | 20 | 0.999 | 60 | 119.9897 | 20.0055 | 99.905 | 2.3982 | -0.1046 | 2.4005 | 119.764 | 20.005 | 2.3959 | -0.2257 | -0.000516 | -0.002251 | -0.1881 | -0.0026 | -0.0939
current_sweep | 15 | 1 | 17 | 3 | 120 | 20 | 0.999 | 60 | 119.9897 | 20.0055 | 99.905 | 2.3982 | -0.1046 | 2.4005 | 119.764 | 20.014 | 2.3971 | -0.2257 | 0.008484 | -0.001075 | -0.1881 | 0.0424 | -0.0448
current_sweep | 15 | 1 | 18 | 3 | 120 | 20 | 0.999 | 60 | 119.9897 | 20.0055 | 99.905 | 2.3982 | -0.1046 | 2.4005 | 119.764 | 20.008 | 2.3963 | -0.2257 | 0.002484 | -0.001861 | -0.1881 | 0.0124 | -0.0776
current_sweep | 15 | 2 | 1 | 1 | 120 | 20 | 0.999 | 60 | 119.9964 | 19.9697 | 99.9016 | 2.3939 | -0.10459 | 2.3963 | 119.739 | 60.016 | 7.1863 | -0.2574 | 40.0463 | 4.7924 | -0.2145 | 0.17868 | 0.063514
current_sweep | 15 | 2 | 2 | 1 | 120 | 20 | 0.999 | 60 | 119.9964 | 19.9697 | 99.9016 | 2.3939 | -0.10459 | 2.3963 | 119.739 | 20.007 | 2.3957 | -0.2574 | 0.037348 | 0.001762 | -0.2145 | 0.187 | 0.0736
current_sweep | 15 | 2 | 3 | 1 | 120 | 20 | 0.999 | 60 | 119.9964 | 19.9697 | 99.9016 | 2.3939 | -0.10459 | 2.3963 | 119.739 | 20.011 | 2.3967 | -0.2574 | 0.041348 | 0.002814 | -0.2145 | 0.2071 | 0.1176
current_sweep | 15 | 2 | 4 | 1 | 120 | 20 | 0.999 | 60 | 119.9964 | 19.9697 | 99.9016 | 2.3939 | -0.10459 | 2.3963 | 119.739 | 19.99 | 2.3937 | -0.2574 | 0.020348 | -0.000277 | -0.2145 | 0.1019 | -0.0116
current_sweep | 15 | 2 | 5 | 1 | 120 | 20 | 0.999 | 60 | 119.9964 | 19.9697 | 99.9016 | 2.3939 | -0.10459 | 2.3963 | 119.739 | 20.008 | 2.3958 | -0.2574 | 0.038348 | 0.001866 | -0.2145 | 0.192 | 0.078
current_sweep | 15 | 2 | 6 | 1 | 120 | 20 | 0.999 | 60 | 119.9964 | 19.9697 | 99.9016 | 2.3939 | -0.10459 | 2.3963 | 119.739 | 19.962 | 2.3908 | -0.2574 | -0.007652 | -0.003087 | -0.2145 | -0.0383 | -0.1289
current_sweep | 15 | 2 | 7 | 2 | 120 | 20 | 0.999 | 60 | 119.989 | 20.0096 | 99.905 | 2.3987 | -0.1045 | 2.4009 | 119.76 | 60.206 | 7.2104 | -0.22903 | 40.1964 | 4.8117 | -0.1909 | 0.29513 | 0.20009
current_sweep | 15 | 2 | 8 | 2 | 120 | 20 | 0.999 | 60 | 119.989 | 20.0096 | 99.905 | 2.3987 | -0.1045 | 2.4009 | 119.76 | 20.035 | 2.3997 | -0.22903 | 0.025387 | 0.001081 | -0.1909 | 0.1269 | 0.0451
current_sweep | 15 | 2 | 9 | 2 | 120 | 20 | 0.999 | 60 | 119.989 | 20.0096 | 99.905 | 2.3987 | -0.1045 | 2.4009 | 119.76 | 20.031 | 2.3989 | -0.22903 | 0.021387 | 0.00023 | -0.1909 | 0.1069 | 0.0096
current_sweep | 15 | 2 | 10 | 2 | 120 | 20 | 0.999 | 60 | 119.989 | 20.0096 | 99.905 | 2.3987 | -0.1045 | 2.4009 | 119.76 | 20.018 | 2.3975 | -0.22903 | 0.008387 | -0.001184 | -0.1909 | 0.0419 | -0.0494
current_sweep | 15 | 2 | 11 | 2 | 120 | 20 | 0.999 | 60 | 119.989 | 20.0096 | 99.905 | 2.3987 | -0.1045 | 2.4009 | 119.76 | 20.014 | 2.3972 | -0.22903 | 0.004387 | -0.001422 | -0.1909 | 0.0219 | -0.0593
current_sweep | 15 | 2 | 12 | 2 | 120 | 20 | 0.999 | 60 | 119.989 | 20.0096 | 99.905 | 2.3987 | -0.1045 | 2.4009 | 119.76 | 20.02 | 2.3976 | -0.22903 | 0.010387 | -0.001034 | -0.1909 | 0.0519 | -0.0431
current_sweep | 15 | 2 | 13 | 3 | 120 | 20 | 0.999 | 60 | 119.9929 | 20.0042 | 99.9049 | 2.3981 | -0.10464 | 2.4004 | 119.776 | 20.013 | 2.3972 | -0.21692 | 0.00884 | -0.000892 | -0.1808 | 0.0442 | -0.0372
current_sweep | 15 | 2 | 14 | 3 | 120 | 20 | 0.999 | 60 | 119.9929 | 20.0042 | 99.9049 | 2.3981 | -0.10464 | 2.4004 | 119.776 | 20.002 | 2.3956 | -0.21692 | -0.00216 | -0.002488 | -0.1808 | -0.0108 | -0.1038
current_sweep | 15 | 2 | 15 | 3 | 120 | 20 | 0.999 | 60 | 119.9929 | 20.0042 | 99.9049 | 2.3981 | -0.10464 | 2.4004 | 119.776 | 20.037 | 2.3998 | -0.21692 | 0.03284 | 0.001736 | -0.1808 | 0.1642 | 0.0724
current_sweep | 15 | 2 | 16 | 3 | 120 | 20 | 0.999 | 60 | 119.9929 | 20.0042 | 99.9049 | 2.3981 | -0.10464 | 2.4004 | 119.776 | 20.008 | 2.3965 | -0.21692 | 0.00384 | -0.001583 | -0.1808 | 0.0192 | -0.066
current_sweep | 15 | 2 | 17 | 3 | 120 | 20 | 0.999 | 60 | 119.9929 | 20.0042 | 99.9049 | 2.3981 | -0.10464 | 2.4004 | 119.776 | 20.012 | 2.3968 | -0.21692 | 0.00784 | -0.001317 | -0.1808 | 0.0392 | -0.0549
current_sweep | 15 | 2 | 18 | 3 | 120 | 20 | 0.999 | 60 | 119.9929 | 20.0042 | 99.9049 | 2.3981 | -0.10464 | 2.4004 | 119.776 | 20.011 | 2.3967 | -0.21692 | 0.00684 | -0.001391 | -0.1808 | 0.0342 | -0.058
current_sweep | 15 | 3 | 1 | 1 | 120 | 20 | 0.999 | 60 | 120.0116 | 19.968 | 99.9011 | 2.394 | -0.10479 | 2.3964 | 119.736 | 59.999 | 7.1834 | -0.27564 | 40.031 | 4.7893 | -0.2297 | 0.15842 | 0.017761
current_sweep | 15 | 3 | 2 | 1 | 120 | 20 | 0.999 | 60 | 120.0116 | 19.968 | 99.9011 | 2.394 | -0.10479 | 2.3964 | 119.736 | 20.002 | 2.395 | -0.27564 | 0.033967 | 0.000953 | -0.2297 | 0.1701 | 0.0398
current_sweep | 15 | 3 | 3 | 1 | 120 | 20 | 0.999 | 60 | 120.0116 | 19.968 | 99.9011 | 2.394 | -0.10479 | 2.3964 | 119.736 | 20.004 | 2.3956 | -0.27564 | 0.035967 | 0.001619 | -0.2297 | 0.1801 | 0.0676
current_sweep | 15 | 3 | 4 | 1 | 120 | 20 | 0.999 | 60 | 120.0116 | 19.968 | 99.9011 | 2.394 | -0.10479 | 2.3964 | 119.736 | 19.984 | 2.3926 | -0.27564 | 0.015967 | -0.001386 | -0.2297 | 0.08 | -0.0579
current_sweep | 15 | 3 | 5 | 1 | 120 | 20 | 0.999 | 60 | 120.0116 | 19.968 | 99.9011 | 2.394 | -0.10479 | 2.3964 | 119.736 | 20.003 | 2.3951 | -0.27564 | 0.034967 | 0.001056 | -0.2297 | 0.1751 | 0.0441
current_sweep | 15 | 3 | 6 | 1 | 120 | 20 | 0.999 | 60 | 120.0116 | 19.968 | 99.9011 | 2.394 | -0.10479 | 2.3964 | 119.736 | 19.955 | 2.3897 | -0.27564 | -0.013033 | -0.004291 | -0.2297 | -0.0653 | -0.1793
current_sweep | 15 | 3 | 7 | 2 | 120 | 20 | 0.999 | 60 | 119.985 | 20.0092 | 99.905 | 2.3985 | -0.10446 | 2.4008 | 119.778 | 60.217 | 7.212 | -0.20701 | 40.2078 | 4.8135 | -0.1725 | 0.31574 | 0.22836
current_sweep | 15 | 3 | 8 | 2 | 120 | 20 | 0.999 | 60 | 119.985 | 20.0092 | 99.905 | 2.3985 | -0.10446 | 2.4008 | 119.778 | 20.034 | 2.3997 | -0.20701 | 0.024843 | 0.00119 | -0.1725 | 0.1242 | 0.0496
current_sweep | 15 | 3 | 9 | 2 | 120 | 20 | 0.999 | 60 | 119.985 | 20.0092 | 99.905 | 2.3985 | -0.10446 | 2.4008 | 119.778 | 20.031 | 2.3989 | -0.20701 | 0.021843 | 0.000383 | -0.1725 | 0.1092 | 0.016
current_sweep | 15 | 3 | 10 | 2 | 120 | 20 | 0.999 | 60 | 119.985 | 20.0092 | 99.905 | 2.3985 | -0.10446 | 2.4008 | 119.778 | 20.022 | 2.398 | -0.20701 | 0.012843 | -0.000503 | -0.1725 | 0.0642 | -0.021
current_sweep | 15 | 3 | 11 | 2 | 120 | 20 | 0.999 | 60 | 119.985 | 20.0092 | 99.905 | 2.3985 | -0.10446 | 2.4008 | 119.778 | 20.013 | 2.3972 | -0.20701 | 0.003843 | -0.001335 | -0.1725 | 0.0192 | -0.0557
current_sweep | 15 | 3 | 12 | 2 | 120 | 20 | 0.999 | 60 | 119.985 | 20.0092 | 99.905 | 2.3985 | -0.10446 | 2.4008 | 119.778 | 20.02 | 2.3976 | -0.20701 | 0.010843 | -0.000874 | -0.1725 | 0.0542 | -0.0365
current_sweep | 15 | 3 | 13 | 3 | 120 | 20 | 0.999 | 60 | 119.9949 | 20.0055 | 99.9052 | 2.3983 | -0.10449 | 2.4006 | 119.766 | 20.014 | 2.3972 | -0.22887 | 0.008507 | -0.001054 | -0.1907 | 0.0425 | -0.044
current_sweep | 15 | 3 | 14 | 3 | 120 | 20 | 0.999 | 60 | 119.9949 | 20.0055 | 99.9052 | 2.3983 | -0.10449 | 2.4006 | 119.766 | 20.005 | 2.396 | -0.22887 | -0.000493 | -0.002323 | -0.1907 | -0.0025 | -0.0969
current_sweep | 15 | 3 | 15 | 3 | 120 | 20 | 0.999 | 60 | 119.9949 | 20.0055 | 99.9052 | 2.3983 | -0.10449 | 2.4006 | 119.766 | 20.038 | 2.4002 | -0.22887 | 0.032507 | 0.001904 | -0.1907 | 0.1625 | 0.0794
current_sweep | 15 | 3 | 16 | 3 | 120 | 20 | 0.999 | 60 | 119.9949 | 20.0055 | 99.9052 | 2.3983 | -0.10449 | 2.4006 | 119.766 | 20.008 | 2.3965 | -0.22887 | 0.002507 | -0.001748 | -0.1907 | 0.0125 | -0.0729
current_sweep | 15 | 3 | 17 | 3 | 120 | 20 | 0.999 | 60 | 119.9949 | 20.0055 | 99.9052 | 2.3983 | -0.10449 | 2.4006 | 119.766 | 20.014 | 2.3971 | -0.22887 | 0.008507 | -0.00117 | -0.1907 | 0.0425 | -0.0488
current_sweep | 15 | 3 | 18 | 3 | 120 | 20 | 0.999 | 60 | 119.9949 | 20.0055 | 99.9052 | 2.3983 | -0.10449 | 2.4006 | 119.766 | 20.012 | 2.3971 | -0.22887 | 0.006507 | -0.001222 | -0.1907 | 0.0325 | -0.051

## Voltage_Sweep
sweep | point | sample | channel | phase | source_vln | source_i | source_pf | source_freq | ref_vln | ref_i | ref_pf | ref_kW | ref_kVAR | ref_kVA | dut_vln | dut_i | dut_kW | err_vln | err_i | err_kW | pct_err_vln | pct_err_i | pct_err_kW
voltage_sweep | 1 | 1 | 1 | 1 | 100 | 1 | 0.999 | 60 | 99.9971 | 0.99998 | 99.8999 | 0.099895 | -0.0044572 | 0.099995 | 99.713 | 2.97 | 0.2963 | -0.28415 | 1.97 | 0.19641 | -0.2842 | -0.99801 | -1.1283
voltage_sweep | 1 | 1 | 2 | 1 | 100 | 1 | 0.999 | 60 | 99.9971 | 0.99998 | 99.8999 | 0.099895 | -0.0044572 | 0.099995 | 99.713 | 1.001 | 0.099874 | -0.28415 | 0.00102 | -2.1e-05 | -0.2842 | 0.102 | -0.0211
voltage_sweep | 1 | 1 | 3 | 1 | 100 | 1 | 0.999 | 60 | 99.9971 | 0.99998 | 99.8999 | 0.099895 | -0.0044572 | 0.099995 | 99.713 | 1.001 | 0.099959 | -0.28415 | 0.00102 | 6.4e-05 | -0.2842 | 0.102 | 0.064
voltage_sweep | 1 | 1 | 4 | 1 | 100 | 1 | 0.999 | 60 | 99.9971 | 0.99998 | 99.8999 | 0.099895 | -0.0044572 | 0.099995 | 99.713 | 1 | 0.099834 | -0.28415 | 2e-05 | -6.1e-05 | -0.2842 | 0.002 | -0.0612
voltage_sweep | 1 | 1 | 5 | 1 | 100 | 1 | 0.999 | 60 | 99.9971 | 0.99998 | 99.8999 | 0.099895 | -0.0044572 | 0.099995 | 99.713 | 1.001 | 0.099872 | -0.28415 | 0.00102 | -2.3e-05 | -0.2842 | 0.102 | -0.0231
voltage_sweep | 1 | 1 | 6 | 1 | 100 | 1 | 0.999 | 60 | 99.9971 | 0.99998 | 99.8999 | 0.099895 | -0.0044572 | 0.099995 | 99.713 | 0.998 | 0.099667 | -0.28415 | -0.00198 | -0.000228 | -0.2842 | -0.198 | -0.2283
voltage_sweep | 1 | 1 | 7 | 2 | 100 | 1 | 0.999 | 60 | 99.9844 | 1.0002 | 99.9036 | 0.099905 | -0.0043899 | 0.1 | 99.793 | 2.979 | 0.29736 | -0.19137 | 1.9788 | 0.19745 | -0.1914 | -0.71652 | -0.78632
voltage_sweep | 1 | 1 | 8 | 2 | 100 | 1 | 0.999 | 60 | 99.9844 | 1.0002 | 99.9036 | 0.099905 | -0.0043899 | 0.1 | 99.793 | 1 | 0.099854 | -0.19137 | -0.000166 | -5.1e-05 | -0.1914 | -0.0166 | -0.0506
voltage_sweep | 1 | 1 | 9 | 2 | 100 | 1 | 0.999 | 60 | 99.9844 | 1.0002 | 99.9036 | 0.099905 | -0.0043899 | 0.1 | 99.793 | 1.001 | 0.099912 | -0.19137 | 0.000834 | 7e-06 | -0.1914 | 0.0833 | 0.0074
voltage_sweep | 1 | 1 | 10 | 2 | 100 | 1 | 0.999 | 60 | 99.9844 | 1.0002 | 99.9036 | 0.099905 | -0.0043899 | 0.1 | 99.793 | 1 | 0.099859 | -0.19137 | -0.000166 | -4.6e-05 | -0.1914 | -0.0166 | -0.0456
voltage_sweep | 1 | 1 | 11 | 2 | 100 | 1 | 0.999 | 60 | 99.9844 | 1.0002 | 99.9036 | 0.099905 | -0.0043899 | 0.1 | 99.793 | 1 | 0.099799 | -0.19137 | -0.000166 | -0.000106 | -0.1914 | -0.0166 | -0.1057
voltage_sweep | 1 | 1 | 12 | 2 | 100 | 1 | 0.999 | 60 | 99.9844 | 1.0002 | 99.9036 | 0.099905 | -0.0043899 | 0.1 | 99.793 | 1 | 0.099855 | -0.19137 | -0.000166 | -5e-05 | -0.1914 | -0.0166 | -0.0496
voltage_sweep | 1 | 1 | 13 | 3 | 100 | 1 | 0.999 | 60 | 99.9948 | 0.99994 | 99.9006 | 0.09989 | -0.0044567 | 0.099989 | 99.789 | 0.999 | 0.099787 | -0.20581 | -0.000943 | -0.000103 | -0.2058 | -0.0943 | -0.1029
voltage_sweep | 1 | 1 | 14 | 3 | 100 | 1 | 0.999 | 60 | 99.9948 | 0.99994 | 99.9006 | 0.09989 | -0.0044567 | 0.099989 | 99.789 | 0.999 | 0.099744 | -0.20581 | -0.000943 | -0.000146 | -0.2058 | -0.0943 | -0.1459
voltage_sweep | 1 | 1 | 15 | 3 | 100 | 1 | 0.999 | 60 | 99.9948 | 0.99994 | 99.9006 | 0.09989 | -0.0044567 | 0.099989 | 99.789 | 1 | 0.099886 | -0.20581 | 5.7e-05 | -4e-06 | -0.2058 | 0.0057 | -0.0038
voltage_sweep | 1 | 1 | 16 | 3 | 100 | 1 | 0.999 | 60 | 99.9948 | 0.99994 | 99.9006 | 0.09989 | -0.0044567 | 0.099989 | 99.789 | 0.999 | 0.099757 | -0.20581 | -0.000943 | -0.000133 | -0.2058 | -0.0943 | -0.1329
voltage_sweep | 1 | 1 | 17 | 3 | 100 | 1 | 0.999 | 60 | 99.9948 | 0.99994 | 99.9006 | 0.09989 | -0.0044567 | 0.099989 | 99.789 | 1 | 0.099802 | -0.20581 | 5.7e-05 | -8.8e-05 | -0.2058 | 0.0057 | -0.0879
voltage_sweep | 1 | 1 | 18 | 3 | 100 | 1 | 0.999 | 60 | 99.9948 | 0.99994 | 99.9006 | 0.09989 | -0.0044567 | 0.099989 | 99.789 | 1 | 0.099806 | -0.20581 | 5.7e-05 | -8.4e-05 | -0.2058 | 0.0057 | -0.0839
voltage_sweep | 1 | 2 | 1 | 1 | 100 | 1 | 0.999 | 60 | 99.9935 | 0.99995 | 99.899 | 0.099887 | -0.0044586 | 0.099988 | 99.696 | 2.97 | 0.29625 | -0.29745 | 1.9701 | 0.19636 | -0.2975 | -0.99493 | -1.1387
voltage_sweep | 1 | 2 | 2 | 1 | 100 | 1 | 0.999 | 60 | 99.9935 | 0.99995 | 99.899 | 0.099887 | -0.0044586 | 0.099988 | 99.696 | 1.001 | 0.099863 | -0.29745 | 0.001051 | -2.4e-05 | -0.2975 | 0.1051 | -0.0244
voltage_sweep | 1 | 2 | 3 | 1 | 100 | 1 | 0.999 | 60 | 99.9935 | 0.99995 | 99.899 | 0.099887 | -0.0044586 | 0.099988 | 99.696 | 1.001 | 0.099909 | -0.29745 | 0.001051 | 2.2e-05 | -0.2975 | 0.1051 | 0.0216
voltage_sweep | 1 | 2 | 4 | 1 | 100 | 1 | 0.999 | 60 | 99.9935 | 0.99995 | 99.899 | 0.099887 | -0.0044586 | 0.099988 | 99.696 | 1 | 0.099831 | -0.29745 | 5.1e-05 | -5.6e-05 | -0.2975 | 0.0051 | -0.0564
voltage_sweep | 1 | 2 | 5 | 1 | 100 | 1 | 0.999 | 60 | 99.9935 | 0.99995 | 99.899 | 0.099887 | -0.0044586 | 0.099988 | 99.696 | 1.001 | 0.099834 | -0.29745 | 0.001051 | -5.3e-05 | -0.2975 | 0.1051 | -0.0534
voltage_sweep | 1 | 2 | 6 | 1 | 100 | 1 | 0.999 | 60 | 99.9935 | 0.99995 | 99.899 | 0.099887 | -0.0044586 | 0.099988 | 99.696 | 0.998 | 0.099635 | -0.29745 | -0.001949 | -0.000252 | -0.2975 | -0.1949 | -0.2527
voltage_sweep | 1 | 2 | 7 | 2 | 100 | 1 | 0.999 | 60 | 99.9835 | 1.0002 | 99.9037 | 0.099904 | -0.0043877 | 0.1 | 99.788 | 2.979 | 0.29732 | -0.1955 | 1.9788 | 0.19742 | -0.1955 | -0.71652 | -0.79659
voltage_sweep | 1 | 2 | 8 | 2 | 100 | 1 | 0.999 | 60 | 99.9835 | 1.0002 | 99.9037 | 0.099904 | -0.0043877 | 0.1 | 99.788 | 1 | 0.099839 | -0.1955 | -0.000166 | -6.5e-05 | -0.1955 | -0.0166 | -0.0649
voltage_sweep | 1 | 2 | 9 | 2 | 100 | 1 | 0.999 | 60 | 99.9835 | 1.0002 | 99.9037 | 0.099904 | -0.0043877 | 0.1 | 99.788 | 1.001 | 0.09992 | -0.1955 | 0.000834 | 1.6e-05 | -0.1955 | 0.0833 | 0.0162
voltage_sweep | 1 | 2 | 10 | 2 | 100 | 1 | 0.999 | 60 | 99.9835 | 1.0002 | 99.9037 | 0.099904 | -0.0043877 | 0.1 | 99.788 | 1 | 0.099863 | -0.1955 | -0.000166 | -4.1e-05 | -0.1955 | -0.0166 | -0.0409
voltage_sweep | 1 | 2 | 11 | 2 | 100 | 1 | 0.999 | 60 | 99.9835 | 1.0002 | 99.9037 | 0.099904 | -0.0043877 | 0.1 | 99.788 | 1 | 0.099797 | -0.1955 | -0.000166 | -0.000107 | -0.1955 | -0.0166 | -0.1069
voltage_sweep | 1 | 2 | 12 | 2 | 100 | 1 | 0.999 | 60 | 99.9835 | 1.0002 | 99.9037 | 0.099904 | -0.0043877 | 0.1 | 99.788 | 1 | 0.099862 | -0.1955 | -0.000166 | -4.2e-05 | -0.1955 | -0.0166 | -0.0419
voltage_sweep | 1 | 2 | 13 | 3 | 100 | 1 | 0.999 | 60 | 99.9906 | 0.99992 | 99.9004 | 0.099883 | -0.0044605 | 0.099983 | 99.798 | 0.999 | 0.099789 | -0.19263 | -0.00092 | -9.4e-05 | -0.1926 | -0.092 | -0.0941
voltage_sweep | 1 | 2 | 14 | 3 | 100 | 1 | 0.999 | 60 | 99.9906 | 0.99992 | 99.9004 | 0.099883 | -0.0044605 | 0.099983 | 99.798 | 0.999 | 0.099765 | -0.19263 | -0.00092 | -0.000118 | -0.1926 | -0.092 | -0.1182
voltage_sweep | 1 | 2 | 15 | 3 | 100 | 1 | 0.999 | 60 | 99.9906 | 0.99992 | 99.9004 | 0.099883 | -0.0044605 | 0.099983 | 99.798 | 1 | 0.099887 | -0.19263 | 8e-05 | 4e-06 | -0.1926 | 0.008 | 0.004
voltage_sweep | 1 | 2 | 16 | 3 | 100 | 1 | 0.999 | 60 | 99.9906 | 0.99992 | 99.9004 | 0.099883 | -0.0044605 | 0.099983 | 99.798 | 0.999 | 0.099763 | -0.19263 | -0.00092 | -0.00012 | -0.1926 | -0.092 | -0.1202
voltage_sweep | 1 | 2 | 17 | 3 | 100 | 1 | 0.999 | 60 | 99.9906 | 0.99992 | 99.9004 | 0.099883 | -0.0044605 | 0.099983 | 99.798 | 1 | 0.099825 | -0.19263 | 8e-05 | -5.8e-05 | -0.1926 | 0.008 | -0.0581
voltage_sweep | 1 | 2 | 18 | 3 | 100 | 1 | 0.999 | 60 | 99.9906 | 0.99992 | 99.9004 | 0.099883 | -0.0044605 | 0.099983 | 99.798 | 1 | 0.099798 | -0.19263 | 8e-05 | -8.5e-05 | -0.1926 | 0.008 | -0.0851
voltage_sweep | 1 | 3 | 1 | 1 | 100 | 1 | 0.999 | 60 | 100.007 | 1 | 99.9005 | 0.099907 | -0.0044536 | 0.10001 | 99.735 | 2.97 | 0.29628 | -0.27196 | 1.97 | 0.19637 | -0.2719 | -0.99995 | -1.1481
voltage_sweep | 1 | 3 | 2 | 1 | 100 | 1 | 0.999 | 60 | 100.007 | 1 | 99.9005 | 0.099907 | -0.0044536 | 0.10001 | 99.735 | 1.001 | 0.099898 | -0.27196 | 0.001001 | -9e-06 | -0.2719 | 0.1001 | -0.0094
voltage_sweep | 1 | 3 | 3 | 1 | 100 | 1 | 0.999 | 60 | 100.007 | 1 | 99.9005 | 0.099907 | -0.0044536 | 0.10001 | 99.735 | 1.001 | 0.099928 | -0.27196 | 0.001001 | 2.1e-05 | -0.2719 | 0.1001 | 0.0206
voltage_sweep | 1 | 3 | 4 | 1 | 100 | 1 | 0.999 | 60 | 100.007 | 1 | 99.9005 | 0.099907 | -0.0044536 | 0.10001 | 99.735 | 1 | 0.099826 | -0.27196 | 1e-06 | -8.1e-05 | -0.2719 | 0.0001 | -0.0815
voltage_sweep | 1 | 3 | 5 | 1 | 100 | 1 | 0.999 | 60 | 100.007 | 1 | 99.9005 | 0.099907 | -0.0044536 | 0.10001 | 99.735 | 1.001 | 0.099888 | -0.27196 | 0.001001 | -1.9e-05 | -0.2719 | 0.1001 | -0.0194
voltage_sweep | 1 | 3 | 6 | 1 | 100 | 1 | 0.999 | 60 | 100.007 | 1 | 99.9005 | 0.099907 | -0.0044536 | 0.10001 | 99.735 | 0.998 | 0.099598 | -0.27196 | -0.001999 | -0.000309 | -0.2719 | -0.1999 | -0.3097
voltage_sweep | 1 | 3 | 7 | 2 | 100 | 1 | 0.999 | 60 | 99.9891 | 1.0002 | 99.9035 | 0.099908 | -0.0043926 | 0.1 | 99.795 | 2.979 | 0.29737 | -0.19408 | 1.9788 | 0.19746 | -0.1941 | -0.71536 | -0.78606
voltage_sweep | 1 | 3 | 8 | 2 | 100 | 1 | 0.999 | 60 | 99.9891 | 1.0002 | 99.9035 | 0.099908 | -0.0043926 | 0.1 | 99.795 | 1 | 0.099851 | -0.19408 | -0.000155 | -5.7e-05 | -0.1941 | -0.0155 | -0.0571
voltage_sweep | 1 | 3 | 9 | 2 | 100 | 1 | 0.999 | 60 | 99.9891 | 1.0002 | 99.9035 | 0.099908 | -0.0043926 | 0.1 | 99.795 | 1.001 | 0.099918 | -0.19408 | 0.000845 | 1e-05 | -0.1941 | 0.0845 | 0.01
voltage_sweep | 1 | 3 | 10 | 2 | 100 | 1 | 0.999 | 60 | 99.9891 | 1.0002 | 99.9035 | 0.099908 | -0.0043926 | 0.1 | 99.795 | 1 | 0.099886 | -0.19408 | -0.000155 | -2.2e-05 | -0.1941 | -0.0155 | -0.022
voltage_sweep | 1 | 3 | 11 | 2 | 100 | 1 | 0.999 | 60 | 99.9891 | 1.0002 | 99.9035 | 0.099908 | -0.0043926 | 0.1 | 99.795 | 1 | 0.099815 | -0.19408 | -0.000155 | -9.3e-05 | -0.1941 | -0.0155 | -0.0931
voltage_sweep | 1 | 3 | 12 | 2 | 100 | 1 | 0.999 | 60 | 99.9891 | 1.0002 | 99.9035 | 0.099908 | -0.0043926 | 0.1 | 99.795 | 1 | 0.099851 | -0.19408 | -0.000155 | -5.7e-05 | -0.1941 | -0.0155 | -0.0571
voltage_sweep | 1 | 3 | 13 | 3 | 100 | 1 | 0.999 | 60 | 99.9934 | 0.99994 | 99.9007 | 0.099888 | -0.0044538 | 0.099987 | 99.792 | 0.999 | 0.099787 | -0.20136 | -0.000935 | -0.000101 | -0.2014 | -0.0935 | -0.1007
voltage_sweep | 1 | 3 | 14 | 3 | 100 | 1 | 0.999 | 60 | 99.9934 | 0.99994 | 99.9007 | 0.099888 | -0.0044538 | 0.099987 | 99.792 | 0.999 | 0.099747 | -0.20136 | -0.000935 | -0.000141 | -0.2014 | -0.0935 | -0.1408
voltage_sweep | 1 | 3 | 15 | 3 | 100 | 1 | 0.999 | 60 | 99.9934 | 0.99994 | 99.9007 | 0.099888 | -0.0044538 | 0.099987 | 99.792 | 1 | 0.099885 | -0.20136 | 6.5e-05 | -3e-06 | -0.2014 | 0.0065 | -0.0026
voltage_sweep | 1 | 3 | 16 | 3 | 100 | 1 | 0.999 | 60 | 99.9934 | 0.99994 | 99.9007 | 0.099888 | -0.0044538 | 0.099987 | 99.792 | 0.999 | 0.099765 | -0.20136 | -0.000935 | -0.000123 | -0.2014 | -0.0935 | -0.1227
voltage_sweep | 1 | 3 | 17 | 3 | 100 | 1 | 0.999 | 60 | 99.9934 | 0.99994 | 99.9007 | 0.099888 | -0.0044538 | 0.099987 | 99.792 | 1 | 0.0998 | -0.20136 | 6.5e-05 | -8.8e-05 | -0.2014 | 0.0065 | -0.0877
voltage_sweep | 1 | 3 | 18 | 3 | 100 | 1 | 0.999 | 60 | 99.9934 | 0.99994 | 99.9007 | 0.099888 | -0.0044538 | 0.099987 | 99.792 | 1 | 0.099812 | -0.20136 | 6.5e-05 | -7.6e-05 | -0.2014 | 0.0065 | -0.0757
voltage_sweep | 2 | 1 | 1 | 1 | 115.6 | 1 | 0.999 | 60 | 115.602 | 0.99999 | 99.9004 | 0.11549 | -0.0051444 | 0.1156 | 115.361 | 2.969 | 0.34257 | -0.24099 | 1.969 | 0.22709 | -0.2085 | -1.0322 | -1.1207
voltage_sweep | 2 | 1 | 2 | 1 | 115.6 | 1 | 0.999 | 60 | 115.602 | 0.99999 | 99.9004 | 0.11549 | -0.0051444 | 0.1156 | 115.361 | 1.001 | 0.11552 | -0.24099 | 0.001011 | 3.4e-05 | -0.2085 | 0.1011 | 0.0298
voltage_sweep | 2 | 1 | 3 | 1 | 115.6 | 1 | 0.999 | 60 | 115.602 | 0.99999 | 99.9004 | 0.11549 | -0.0051444 | 0.1156 | 115.361 | 1.001 | 0.11554 | -0.24099 | 0.001011 | 5.1e-05 | -0.2085 | 0.1011 | 0.0446
voltage_sweep | 2 | 1 | 4 | 1 | 115.6 | 1 | 0.999 | 60 | 115.602 | 0.99999 | 99.9004 | 0.11549 | -0.0051444 | 0.1156 | 115.361 | 1 | 0.11542 | -0.24099 | 1.1e-05 | -6.2e-05 | -0.2085 | 0.0011 | -0.0533
voltage_sweep | 2 | 1 | 5 | 1 | 115.6 | 1 | 0.999 | 60 | 115.602 | 0.99999 | 99.9004 | 0.11549 | -0.0051444 | 0.1156 | 115.361 | 1.001 | 0.11551 | -0.24099 | 0.001011 | 2e-05 | -0.2085 | 0.1011 | 0.0177
voltage_sweep | 2 | 1 | 6 | 1 | 115.6 | 1 | 0.999 | 60 | 115.602 | 0.99999 | 99.9004 | 0.11549 | -0.0051444 | 0.1156 | 115.361 | 0.998 | 0.11523 | -0.24099 | -0.001989 | -0.000257 | -0.2085 | -0.1989 | -0.2221
voltage_sweep | 2 | 1 | 7 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5831 | 1.0002 | 99.9036 | 0.11549 | -0.0050738 | 0.1156 | 115.385 | 2.978 | 0.34368 | -0.19808 | 1.9778 | 0.22819 | -0.1714 | -0.74955 | -0.80633
voltage_sweep | 2 | 1 | 8 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5831 | 1.0002 | 99.9036 | 0.11549 | -0.0050738 | 0.1156 | 115.385 | 1 | 0.11542 | -0.19808 | -0.000163 | -7.3e-05 | -0.1714 | -0.0163 | -0.0628
voltage_sweep | 2 | 1 | 9 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5831 | 1.0002 | 99.9036 | 0.11549 | -0.0050738 | 0.1156 | 115.385 | 1 | 0.1155 | -0.19808 | -0.000163 | 8e-06 | -0.1714 | -0.0163 | 0.0073
voltage_sweep | 2 | 1 | 10 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5831 | 1.0002 | 99.9036 | 0.11549 | -0.0050738 | 0.1156 | 115.385 | 1 | 0.11543 | -0.19808 | -0.000163 | -6e-05 | -0.1714 | -0.0163 | -0.0516
voltage_sweep | 2 | 1 | 11 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5831 | 1.0002 | 99.9036 | 0.11549 | -0.0050738 | 0.1156 | 115.385 | 1 | 0.11539 | -0.19808 | -0.000163 | -0.0001 | -0.1714 | -0.0163 | -0.0862
voltage_sweep | 2 | 1 | 12 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5831 | 1.0002 | 99.9036 | 0.11549 | -0.0050738 | 0.1156 | 115.385 | 1 | 0.11544 | -0.19808 | -0.000163 | -4.6e-05 | -0.1714 | -0.0163 | -0.0395
voltage_sweep | 2 | 1 | 13 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5924 | 0.99993 | 99.9005 | 0.11547 | -0.0051532 | 0.11558 | 115.382 | 1 | 0.11541 | -0.2104 | 7.3e-05 | -6.3e-05 | -0.182 | 0.0073 | -0.0545
voltage_sweep | 2 | 1 | 14 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5924 | 0.99993 | 99.9005 | 0.11547 | -0.0051532 | 0.11558 | 115.382 | 0.999 | 0.11536 | -0.2104 | -0.000927 | -0.000112 | -0.182 | -0.0927 | -0.097
voltage_sweep | 2 | 1 | 15 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5924 | 0.99993 | 99.9005 | 0.11547 | -0.0051532 | 0.11558 | 115.382 | 1 | 0.11549 | -0.2104 | 7.3e-05 | 1.9e-05 | -0.182 | 0.0073 | 0.0165
voltage_sweep | 2 | 1 | 16 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5924 | 0.99993 | 99.9005 | 0.11547 | -0.0051532 | 0.11558 | 115.382 | 0.999 | 0.11536 | -0.2104 | -0.000927 | -0.000107 | -0.182 | -0.0927 | -0.0926
voltage_sweep | 2 | 1 | 17 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5924 | 0.99993 | 99.9005 | 0.11547 | -0.0051532 | 0.11558 | 115.382 | 1 | 0.11542 | -0.2104 | 7.3e-05 | -4.7e-05 | -0.182 | 0.0073 | -0.0407
voltage_sweep | 2 | 1 | 18 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5924 | 0.99993 | 99.9005 | 0.11547 | -0.0051532 | 0.11558 | 115.382 | 1 | 0.11539 | -0.2104 | 7.3e-05 | -7.5e-05 | -0.182 | 0.0073 | -0.0649
voltage_sweep | 2 | 2 | 1 | 1 | 115.6 | 1 | 0.999 | 60 | 115.6059 | 0.99998 | 99.9007 | 0.11549 | -0.005149 | 0.1156 | 115.369 | 2.97 | 0.34266 | -0.23686 | 1.97 | 0.22717 | -0.2049 | -0.99824 | -1.0991
voltage_sweep | 2 | 2 | 2 | 1 | 115.6 | 1 | 0.999 | 60 | 115.6059 | 0.99998 | 99.9007 | 0.11549 | -0.005149 | 0.1156 | 115.369 | 1.001 | 0.1155 | -0.23686 | 0.001018 | 1.3e-05 | -0.2049 | 0.1018 | 0.0112
voltage_sweep | 2 | 2 | 3 | 1 | 115.6 | 1 | 0.999 | 60 | 115.6059 | 0.99998 | 99.9007 | 0.11549 | -0.005149 | 0.1156 | 115.369 | 1.001 | 0.1155 | -0.23686 | 0.001018 | 1.6e-05 | -0.2049 | 0.1018 | 0.0138
voltage_sweep | 2 | 2 | 4 | 1 | 115.6 | 1 | 0.999 | 60 | 115.6059 | 0.99998 | 99.9007 | 0.11549 | -0.005149 | 0.1156 | 115.369 | 1 | 0.11546 | -0.23686 | 1.8e-05 | -2.5e-05 | -0.2049 | 0.0018 | -0.0217
voltage_sweep | 2 | 2 | 5 | 1 | 115.6 | 1 | 0.999 | 60 | 115.6059 | 0.99998 | 99.9007 | 0.11549 | -0.005149 | 0.1156 | 115.369 | 1.001 | 0.11547 | -0.23686 | 0.001018 | -1.5e-05 | -0.2049 | 0.1018 | -0.013
voltage_sweep | 2 | 2 | 6 | 1 | 115.6 | 1 | 0.999 | 60 | 115.6059 | 0.99998 | 99.9007 | 0.11549 | -0.005149 | 0.1156 | 115.369 | 0.998 | 0.1152 | -0.23686 | -0.001982 | -0.000291 | -0.2049 | -0.1982 | -0.252
voltage_sweep | 2 | 2 | 7 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5832 | 1.0002 | 99.9038 | 0.11549 | -0.0050684 | 0.1156 | 115.377 | 2.979 | 0.34372 | -0.20623 | 1.9788 | 0.22823 | -0.1784 | -0.71739 | -0.79537
voltage_sweep | 2 | 2 | 8 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5832 | 1.0002 | 99.9038 | 0.11549 | -0.0050684 | 0.1156 | 115.377 | 1 | 0.11544 | -0.20623 | -0.000175 | -4.8e-05 | -0.1784 | -0.0175 | -0.0418
voltage_sweep | 2 | 2 | 9 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5832 | 1.0002 | 99.9038 | 0.11549 | -0.0050684 | 0.1156 | 115.377 | 1.001 | 0.1155 | -0.20623 | 0.000825 | 1.2e-05 | -0.1784 | 0.0825 | 0.0102
voltage_sweep | 2 | 2 | 10 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5832 | 1.0002 | 99.9038 | 0.11549 | -0.0050684 | 0.1156 | 115.377 | 1 | 0.11545 | -0.20623 | -0.000175 | -3.8e-05 | -0.1784 | -0.0175 | -0.0331
voltage_sweep | 2 | 2 | 11 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5832 | 1.0002 | 99.9038 | 0.11549 | -0.0050684 | 0.1156 | 115.377 | 1 | 0.1154 | -0.20623 | -0.000175 | -8.7e-05 | -0.1784 | -0.0175 | -0.0756
voltage_sweep | 2 | 2 | 12 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5832 | 1.0002 | 99.9038 | 0.11549 | -0.0050684 | 0.1156 | 115.377 | 1 | 0.11546 | -0.20623 | -0.000175 | -3e-05 | -0.1784 | -0.0175 | -0.0262
voltage_sweep | 2 | 2 | 13 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5882 | 0.99991 | 99.9006 | 0.11546 | -0.0051518 | 0.11558 | 115.372 | 0.999 | 0.11536 | -0.21621 | -0.000907 | -0.000101 | -0.1871 | -0.0907 | -0.0872
voltage_sweep | 2 | 2 | 14 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5882 | 0.99991 | 99.9006 | 0.11546 | -0.0051518 | 0.11558 | 115.372 | 0.999 | 0.11535 | -0.21621 | -0.000907 | -0.000115 | -0.1871 | -0.0907 | -0.0993
voltage_sweep | 2 | 2 | 15 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5882 | 0.99991 | 99.9006 | 0.11546 | -0.0051518 | 0.11558 | 115.372 | 1 | 0.11548 | -0.21621 | 9.3e-05 | 1.5e-05 | -0.1871 | 0.0093 | 0.0133
voltage_sweep | 2 | 2 | 16 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5882 | 0.99991 | 99.9006 | 0.11546 | -0.0051518 | 0.11558 | 115.372 | 0.999 | 0.11533 | -0.21621 | -0.000907 | -0.000128 | -0.1871 | -0.0907 | -0.1105
voltage_sweep | 2 | 2 | 17 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5882 | 0.99991 | 99.9006 | 0.11546 | -0.0051518 | 0.11558 | 115.372 | 1 | 0.1154 | -0.21621 | 9.3e-05 | -6.3e-05 | -0.1871 | 0.0093 | -0.0542
voltage_sweep | 2 | 2 | 18 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5882 | 0.99991 | 99.9006 | 0.11546 | -0.0051518 | 0.11558 | 115.372 | 1 | 0.11537 | -0.21621 | 9.3e-05 | -8.9e-05 | -0.1871 | 0.0093 | -0.0768
voltage_sweep | 2 | 3 | 1 | 1 | 115.6 | 1 | 0.999 | 60 | 115.5993 | 0.99999 | 99.9004 | 0.11548 | -0.0051493 | 0.1156 | 115.325 | 2.969 | 0.34251 | -0.27427 | 1.969 | 0.22703 | -0.2373 | -1.0323 | -1.1375
voltage_sweep | 2 | 3 | 2 | 1 | 115.6 | 1 | 0.999 | 60 | 115.5993 | 0.99999 | 99.9004 | 0.11548 | -0.0051493 | 0.1156 | 115.325 | 1.001 | 0.11549 | -0.27427 | 0.00101 | 2e-06 | -0.2373 | 0.101 | 0.0018
voltage_sweep | 2 | 3 | 3 | 1 | 115.6 | 1 | 0.999 | 60 | 115.5993 | 0.99999 | 99.9004 | 0.11548 | -0.0051493 | 0.1156 | 115.325 | 1.001 | 0.11553 | -0.27427 | 0.00101 | 4.5e-05 | -0.2373 | 0.101 | 0.039
voltage_sweep | 2 | 3 | 4 | 1 | 115.6 | 1 | 0.999 | 60 | 115.5993 | 0.99999 | 99.9004 | 0.11548 | -0.0051493 | 0.1156 | 115.325 | 1 | 0.11542 | -0.27427 | 1e-05 | -5.9e-05 | -0.2373 | 0.001 | -0.051
voltage_sweep | 2 | 3 | 5 | 1 | 115.6 | 1 | 0.999 | 60 | 115.5993 | 0.99999 | 99.9004 | 0.11548 | -0.0051493 | 0.1156 | 115.325 | 1.001 | 0.11552 | -0.27427 | 0.00101 | 3.2e-05 | -0.2373 | 0.101 | 0.0278
voltage_sweep | 2 | 3 | 6 | 1 | 115.6 | 1 | 0.999 | 60 | 115.5993 | 0.99999 | 99.9004 | 0.11548 | -0.0051493 | 0.1156 | 115.325 | 0.998 | 0.11521 | -0.27427 | -0.00199 | -0.000272 | -0.2373 | -0.199 | -0.2355
voltage_sweep | 2 | 3 | 7 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5868 | 1.0002 | 99.9036 | 0.11549 | -0.0050739 | 0.11561 | 115.383 | 2.979 | 0.34375 | -0.20379 | 1.9788 | 0.22825 | -0.1763 | -0.71588 | -0.78924
voltage_sweep | 2 | 3 | 8 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5868 | 1.0002 | 99.9036 | 0.11549 | -0.0050739 | 0.11561 | 115.383 | 1 | 0.11543 | -0.20379 | -0.00016 | -6.2e-05 | -0.1763 | -0.016 | -0.0536
voltage_sweep | 2 | 3 | 9 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5868 | 1.0002 | 99.9036 | 0.11549 | -0.0050739 | 0.11561 | 115.383 | 1.001 | 0.11552 | -0.20379 | 0.00084 | 2.5e-05 | -0.1763 | 0.084 | 0.0218
voltage_sweep | 2 | 3 | 10 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5868 | 1.0002 | 99.9036 | 0.11549 | -0.0050739 | 0.11561 | 115.383 | 1 | 0.11546 | -0.20379 | -0.00016 | -2.9e-05 | -0.1763 | -0.016 | -0.025
voltage_sweep | 2 | 3 | 11 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5868 | 1.0002 | 99.9036 | 0.11549 | -0.0050739 | 0.11561 | 115.383 | 1 | 0.11539 | -0.20379 | -0.00016 | -0.000103 | -0.1763 | -0.016 | -0.0891
voltage_sweep | 2 | 3 | 12 | 2 | 115.6 | 1 | 0.999 | 60 | 115.5868 | 1.0002 | 99.9036 | 0.11549 | -0.0050739 | 0.11561 | 115.383 | 1 | 0.11544 | -0.20379 | -0.00016 | -5.2e-05 | -0.1763 | -0.016 | -0.0449
voltage_sweep | 2 | 3 | 13 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5919 | 0.99993 | 99.9006 | 0.11547 | -0.0051498 | 0.11558 | 115.369 | 0.999 | 0.11537 | -0.22291 | -0.000928 | -0.000103 | -0.1928 | -0.0928 | -0.089
voltage_sweep | 2 | 3 | 14 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5919 | 0.99993 | 99.9006 | 0.11547 | -0.0051498 | 0.11558 | 115.369 | 0.999 | 0.11536 | -0.22291 | -0.000928 | -0.000113 | -0.1928 | -0.0928 | -0.0977
voltage_sweep | 2 | 3 | 15 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5919 | 0.99993 | 99.9006 | 0.11547 | -0.0051498 | 0.11558 | 115.369 | 1 | 0.11548 | -0.22291 | 7.2e-05 | 1.2e-05 | -0.1928 | 0.0072 | 0.0106
voltage_sweep | 2 | 3 | 16 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5919 | 0.99993 | 99.9006 | 0.11547 | -0.0051498 | 0.11558 | 115.369 | 0.999 | 0.11534 | -0.22291 | -0.000928 | -0.000127 | -0.1928 | -0.0928 | -0.1098
voltage_sweep | 2 | 3 | 17 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5919 | 0.99993 | 99.9006 | 0.11547 | -0.0051498 | 0.11558 | 115.369 | 1 | 0.11541 | -0.22291 | 7.2e-05 | -5.5e-05 | -0.1928 | 0.0072 | -0.0474
voltage_sweep | 2 | 3 | 18 | 3 | 115.6 | 1 | 0.999 | 60 | 115.5919 | 0.99993 | 99.9006 | 0.11547 | -0.0051498 | 0.11558 | 115.369 | 1 | 0.11539 | -0.22291 | 7.2e-05 | -8.2e-05 | -0.1928 | 0.0072 | -0.0708
voltage_sweep | 3 | 1 | 1 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1011 | 1 | 99.9004 | 0.13097 | -0.0058419 | 0.1311 | 130.834 | 2.97 | 0.38862 | -0.26714 | 1.97 | 0.25765 | -0.2038 | -1.0004 | -1.0927
voltage_sweep | 3 | 1 | 2 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1011 | 1 | 99.9004 | 0.13097 | -0.0058419 | 0.1311 | 130.834 | 1.001 | 0.1311 | -0.26714 | 0.000996 | 0.00013 | -0.2038 | 0.0996 | 0.0992
voltage_sweep | 3 | 1 | 3 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1011 | 1 | 99.9004 | 0.13097 | -0.0058419 | 0.1311 | 130.834 | 1.001 | 0.13098 | -0.26714 | 0.000996 | 1e-05 | -0.2038 | 0.0996 | 0.0076
voltage_sweep | 3 | 1 | 4 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1011 | 1 | 99.9004 | 0.13097 | -0.0058419 | 0.1311 | 130.834 | 1 | 0.13093 | -0.26714 | -4e-06 | -4.3e-05 | -0.2038 | -0.0004 | -0.0329
voltage_sweep | 3 | 1 | 5 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1011 | 1 | 99.9004 | 0.13097 | -0.0058419 | 0.1311 | 130.834 | 1.001 | 0.13106 | -0.26714 | 0.000996 | 9.2e-05 | -0.2038 | 0.0996 | 0.0702
voltage_sweep | 3 | 1 | 6 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1011 | 1 | 99.9004 | 0.13097 | -0.0058419 | 0.1311 | 130.834 | 0.998 | 0.13057 | -0.26714 | -0.002004 | -0.0004 | -0.2038 | -0.2004 | -0.3055
voltage_sweep | 3 | 1 | 7 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0832 | 1.0002 | 99.9037 | 0.13098 | -0.0057518 | 0.1311 | 130.864 | 2.979 | 0.3899 | -0.21916 | 1.9788 | 0.25893 | -0.1672 | -0.71586 | -0.77071
voltage_sweep | 3 | 1 | 8 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0832 | 1.0002 | 99.9037 | 0.13098 | -0.0057518 | 0.1311 | 130.864 | 1 | 0.13086 | -0.21916 | -0.00016 | -0.000114 | -0.1672 | -0.016 | -0.0869
voltage_sweep | 3 | 1 | 9 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0832 | 1.0002 | 99.9037 | 0.13098 | -0.0057518 | 0.1311 | 130.864 | 1 | 0.13098 | -0.21916 | -0.00016 | -1e-06 | -0.1672 | -0.016 | -0.0006
voltage_sweep | 3 | 1 | 10 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0832 | 1.0002 | 99.9037 | 0.13098 | -0.0057518 | 0.1311 | 130.864 | 1 | 0.13097 | -0.21916 | -0.00016 | -1.1e-05 | -0.1672 | -0.016 | -0.0082
voltage_sweep | 3 | 1 | 11 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0832 | 1.0002 | 99.9037 | 0.13098 | -0.0057518 | 0.1311 | 130.864 | 1 | 0.13085 | -0.21916 | -0.00016 | -0.000126 | -0.1672 | -0.016 | -0.096
voltage_sweep | 3 | 1 | 12 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0832 | 1.0002 | 99.9037 | 0.13098 | -0.0057518 | 0.1311 | 130.864 | 1 | 0.1309 | -0.21916 | -0.00016 | -7.8e-05 | -0.1672 | -0.016 | -0.0594
voltage_sweep | 3 | 1 | 13 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0886 | 0.99992 | 99.9007 | 0.13095 | -0.0058393 | 0.13108 | 130.866 | 1 | 0.13089 | -0.22261 | 8.5e-05 | -6.1e-05 | -0.1698 | 0.0085 | -0.0468
voltage_sweep | 3 | 1 | 14 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0886 | 0.99992 | 99.9007 | 0.13095 | -0.0058393 | 0.13108 | 130.866 | 0.999 | 0.13081 | -0.22261 | -0.000915 | -0.000136 | -0.1698 | -0.0915 | -0.1041
voltage_sweep | 3 | 1 | 15 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0886 | 0.99992 | 99.9007 | 0.13095 | -0.0058393 | 0.13108 | 130.866 | 1.001 | 0.131 | -0.22261 | 0.001085 | 5.6e-05 | -0.1698 | 0.1085 | 0.0425
voltage_sweep | 3 | 1 | 16 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0886 | 0.99992 | 99.9007 | 0.13095 | -0.0058393 | 0.13108 | 130.866 | 0.999 | 0.13083 | -0.22261 | -0.000915 | -0.000113 | -0.1698 | -0.0915 | -0.0866
voltage_sweep | 3 | 1 | 17 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0886 | 0.99992 | 99.9007 | 0.13095 | -0.0058393 | 0.13108 | 130.866 | 1 | 0.13088 | -0.22261 | 8.5e-05 | -6.3e-05 | -0.1698 | 0.0085 | -0.0484
voltage_sweep | 3 | 1 | 18 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0886 | 0.99992 | 99.9007 | 0.13095 | -0.0058393 | 0.13108 | 130.866 | 1 | 0.13089 | -0.22261 | 8.5e-05 | -6.2e-05 | -0.1698 | 0.0085 | -0.0476
voltage_sweep | 3 | 2 | 1 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1118 | 0.99996 | 99.9004 | 0.13098 | -0.0058475 | 0.13111 | 130.828 | 2.97 | 0.38863 | -0.28383 | 1.97 | 0.25765 | -0.2165 | -0.99646 | -1.0943
voltage_sweep | 3 | 2 | 2 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1118 | 0.99996 | 99.9004 | 0.13098 | -0.0058475 | 0.13111 | 130.828 | 1.001 | 0.13095 | -0.28383 | 0.001036 | -2.6e-05 | -0.2165 | 0.1036 | -0.0195
voltage_sweep | 3 | 2 | 3 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1118 | 0.99996 | 99.9004 | 0.13098 | -0.0058475 | 0.13111 | 130.828 | 1.001 | 0.13111 | -0.28383 | 0.001036 | 0.00013 | -0.2165 | 0.1036 | 0.0996
voltage_sweep | 3 | 2 | 4 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1118 | 0.99996 | 99.9004 | 0.13098 | -0.0058475 | 0.13111 | 130.828 | 1 | 0.13093 | -0.28383 | 3.6e-05 | -4.2e-05 | -0.2165 | 0.0036 | -0.0317
voltage_sweep | 3 | 2 | 5 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1118 | 0.99996 | 99.9004 | 0.13098 | -0.0058475 | 0.13111 | 130.828 | 1.001 | 0.13093 | -0.28383 | 0.001036 | -4.2e-05 | -0.2165 | 0.1036 | -0.0317
voltage_sweep | 3 | 2 | 6 | 1 | 131.1 | 1 | 0.999 | 60 | 131.1118 | 0.99996 | 99.9004 | 0.13098 | -0.0058475 | 0.13111 | 130.828 | 0.998 | 0.13071 | -0.28383 | -0.001964 | -0.000263 | -0.2165 | -0.1964 | -0.2005
voltage_sweep | 3 | 2 | 7 | 2 | 131.1 | 1 | 0.999 | 60 | 131.084 | 1.0001 | 99.9037 | 0.13098 | -0.0057525 | 0.1311 | 130.865 | 2.979 | 0.38989 | -0.21895 | 1.9789 | 0.25892 | -0.167 | -0.71484 | -0.77309
voltage_sweep | 3 | 2 | 8 | 2 | 131.1 | 1 | 0.999 | 60 | 131.084 | 1.0001 | 99.9037 | 0.13098 | -0.0057525 | 0.1311 | 130.865 | 1 | 0.13087 | -0.21895 | -0.000149 | -0.000103 | -0.167 | -0.0149 | -0.0788
voltage_sweep | 3 | 2 | 9 | 2 | 131.1 | 1 | 0.999 | 60 | 131.084 | 1.0001 | 99.9037 | 0.13098 | -0.0057525 | 0.1311 | 130.865 | 1.001 | 0.13101 | -0.21895 | 0.000851 | 3.3e-05 | -0.167 | 0.085 | 0.025
voltage_sweep | 3 | 2 | 10 | 2 | 131.1 | 1 | 0.999 | 60 | 131.084 | 1.0001 | 99.9037 | 0.13098 | -0.0057525 | 0.1311 | 130.865 | 1 | 0.13096 | -0.21895 | -0.000149 | -1.5e-05 | -0.167 | -0.0149 | -0.0116
voltage_sweep | 3 | 2 | 11 | 2 | 131.1 | 1 | 0.999 | 60 | 131.084 | 1.0001 | 99.9037 | 0.13098 | -0.0057525 | 0.1311 | 130.865 | 1 | 0.13085 | -0.21895 | -0.000149 | -0.000122 | -0.167 | -0.0149 | -0.0933
voltage_sweep | 3 | 2 | 12 | 2 | 131.1 | 1 | 0.999 | 60 | 131.084 | 1.0001 | 99.9037 | 0.13098 | -0.0057525 | 0.1311 | 130.865 | 1 | 0.13092 | -0.21895 | -0.000149 | -5.4e-05 | -0.167 | -0.0149 | -0.0414
voltage_sweep | 3 | 2 | 13 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0877 | 0.99991 | 99.9008 | 0.13095 | -0.005836 | 0.13108 | 130.858 | 1 | 0.13086 | -0.22972 | 8.6e-05 | -8.2e-05 | -0.1752 | 0.0086 | -0.0629
voltage_sweep | 3 | 2 | 14 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0877 | 0.99991 | 99.9008 | 0.13095 | -0.005836 | 0.13108 | 130.858 | 0.999 | 0.13083 | -0.22972 | -0.000914 | -0.00012 | -0.1752 | -0.0914 | -0.0919
voltage_sweep | 3 | 2 | 15 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0877 | 0.99991 | 99.9008 | 0.13095 | -0.005836 | 0.13108 | 130.858 | 1.001 | 0.13101 | -0.22972 | 0.001086 | 6e-05 | -0.1752 | 0.1086 | 0.0455
voltage_sweep | 3 | 2 | 16 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0877 | 0.99991 | 99.9008 | 0.13095 | -0.005836 | 0.13108 | 130.858 | 0.999 | 0.13083 | -0.22972 | -0.000914 | -0.000118 | -0.1752 | -0.0914 | -0.0904
voltage_sweep | 3 | 2 | 17 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0877 | 0.99991 | 99.9008 | 0.13095 | -0.005836 | 0.13108 | 130.858 | 1 | 0.1309 | -0.22972 | 8.6e-05 | -4.9e-05 | -0.1752 | 0.0086 | -0.0377
voltage_sweep | 3 | 2 | 18 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0877 | 0.99991 | 99.9008 | 0.13095 | -0.005836 | 0.13108 | 130.858 | 1 | 0.13087 | -0.22972 | 8.6e-05 | -7.1e-05 | -0.1752 | 0.0086 | -0.0545
voltage_sweep | 3 | 3 | 1 | 1 | 131.1 | 1 | 0.999 | 60 | 131.0995 | 0.99997 | 99.9001 | 0.13096 | -0.0058459 | 0.1311 | 130.897 | 2.97 | 0.38875 | -0.20247 | 1.97 | 0.25778 | -0.1544 | -0.99723 | -1.0556
voltage_sweep | 3 | 3 | 2 | 1 | 131.1 | 1 | 0.999 | 60 | 131.0995 | 0.99997 | 99.9001 | 0.13096 | -0.0058459 | 0.1311 | 130.897 | 1.001 | 0.13114 | -0.20247 | 0.001028 | 0.00017 | -0.1544 | 0.1028 | 0.1299
voltage_sweep | 3 | 3 | 3 | 1 | 131.1 | 1 | 0.999 | 60 | 131.0995 | 0.99997 | 99.9001 | 0.13096 | -0.0058459 | 0.1311 | 130.897 | 1.001 | 0.13108 | -0.20247 | 0.001028 | 0.000111 | -0.1544 | 0.1028 | 0.0849
voltage_sweep | 3 | 3 | 4 | 1 | 131.1 | 1 | 0.999 | 60 | 131.0995 | 0.99997 | 99.9001 | 0.13096 | -0.0058459 | 0.1311 | 130.897 | 1 | 0.13099 | -0.20247 | 2.8e-05 | 2.6e-05 | -0.1544 | 0.0028 | 0.02
voltage_sweep | 3 | 3 | 5 | 1 | 131.1 | 1 | 0.999 | 60 | 131.0995 | 0.99997 | 99.9001 | 0.13096 | -0.0058459 | 0.1311 | 130.897 | 1.001 | 0.13111 | -0.20247 | 0.001028 | 0.000141 | -0.1544 | 0.1028 | 0.1078
voltage_sweep | 3 | 3 | 6 | 1 | 131.1 | 1 | 0.999 | 60 | 131.0995 | 0.99997 | 99.9001 | 0.13096 | -0.0058459 | 0.1311 | 130.897 | 0.998 | 0.13069 | -0.20247 | -0.001972 | -0.000277 | -0.1544 | -0.1972 | -0.2114
voltage_sweep | 3 | 3 | 7 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0792 | 1.0002 | 99.9038 | 0.13097 | -0.0057474 | 0.1311 | 130.855 | 2.979 | 0.38987 | -0.22418 | 1.9788 | 0.2589 | -0.171 | -0.71604 | -0.77645
voltage_sweep | 3 | 3 | 8 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0792 | 1.0002 | 99.9038 | 0.13097 | -0.0057474 | 0.1311 | 130.855 | 1 | 0.1309 | -0.22418 | -0.000162 | -7.2e-05 | -0.171 | -0.0162 | -0.0552
voltage_sweep | 3 | 3 | 9 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0792 | 1.0002 | 99.9038 | 0.13097 | -0.0057474 | 0.1311 | 130.855 | 1.001 | 0.13102 | -0.22418 | 0.000838 | 4.7e-05 | -0.171 | 0.0838 | 0.0357
voltage_sweep | 3 | 3 | 10 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0792 | 1.0002 | 99.9038 | 0.13097 | -0.0057474 | 0.1311 | 130.855 | 1 | 0.13095 | -0.22418 | -0.000162 | -2.1e-05 | -0.171 | -0.0162 | -0.0162
voltage_sweep | 3 | 3 | 11 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0792 | 1.0002 | 99.9038 | 0.13097 | -0.0057474 | 0.1311 | 130.855 | 1 | 0.13087 | -0.22418 | -0.000162 | -0.000104 | -0.171 | -0.0162 | -0.0796
voltage_sweep | 3 | 3 | 12 | 2 | 131.1 | 1 | 0.999 | 60 | 131.0792 | 1.0002 | 99.9038 | 0.13097 | -0.0057474 | 0.1311 | 130.855 | 1 | 0.13093 | -0.22418 | -0.000162 | -4e-05 | -0.171 | -0.0162 | -0.0308
voltage_sweep | 3 | 3 | 13 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0906 | 0.99989 | 99.9006 | 0.13095 | -0.0058409 | 0.13108 | 130.852 | 0.999 | 0.13084 | -0.23856 | -0.000894 | -0.000101 | -0.182 | -0.0894 | -0.0774
voltage_sweep | 3 | 3 | 14 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0906 | 0.99989 | 99.9006 | 0.13095 | -0.0058409 | 0.13108 | 130.852 | 0.999 | 0.13081 | -0.23856 | -0.000894 | -0.00014 | -0.182 | -0.0894 | -0.1072
voltage_sweep | 3 | 3 | 15 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0906 | 0.99989 | 99.9006 | 0.13095 | -0.0058409 | 0.13108 | 130.852 | 1 | 0.13098 | -0.23856 | 0.000106 | 3.7e-05 | -0.182 | 0.0106 | 0.028
voltage_sweep | 3 | 3 | 16 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0906 | 0.99989 | 99.9006 | 0.13095 | -0.0058409 | 0.13108 | 130.852 | 0.999 | 0.1308 | -0.23856 | -0.000894 | -0.000148 | -0.182 | -0.0894 | -0.1133
voltage_sweep | 3 | 3 | 17 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0906 | 0.99989 | 99.9006 | 0.13095 | -0.0058409 | 0.13108 | 130.852 | 1 | 0.13086 | -0.23856 | 0.000106 | -8.4e-05 | -0.182 | 0.0106 | -0.0644
voltage_sweep | 3 | 3 | 18 | 3 | 131.1 | 1 | 0.999 | 60 | 131.0906 | 0.99989 | 99.9006 | 0.13095 | -0.0058409 | 0.13108 | 130.852 | 1 | 0.13087 | -0.23856 | 0.000106 | -7.8e-05 | -0.182 | 0.0106 | -0.0598
voltage_sweep | 4 | 1 | 1 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7227 | 0.99998 | 99.8999 | 0.14657 | -0.006543 | 0.14672 | 146.449 | 2.97 | 0.43515 | -0.2737 | 1.97 | 0.28858 | -0.1865 | -0.99791 | -1.0389
voltage_sweep | 4 | 1 | 2 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7227 | 0.99998 | 99.8999 | 0.14657 | -0.006543 | 0.14672 | 146.449 | 1.001 | 0.14665 | -0.2737 | 0.001021 | 7.5e-05 | -0.1865 | 0.1021 | 0.0514
voltage_sweep | 4 | 1 | 3 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7227 | 0.99998 | 99.8999 | 0.14657 | -0.006543 | 0.14672 | 146.449 | 1.001 | 0.14673 | -0.2737 | 0.001021 | 0.000155 | -0.1865 | 0.1021 | 0.106
voltage_sweep | 4 | 1 | 4 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7227 | 0.99998 | 99.8999 | 0.14657 | -0.006543 | 0.14672 | 146.449 | 1 | 0.14662 | -0.2737 | 2.1e-05 | 4.4e-05 | -0.1865 | 0.0021 | 0.0302
voltage_sweep | 4 | 1 | 5 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7227 | 0.99998 | 99.8999 | 0.14657 | -0.006543 | 0.14672 | 146.449 | 1.001 | 0.14662 | -0.2737 | 0.001021 | 4.3e-05 | -0.1865 | 0.1021 | 0.0295
voltage_sweep | 4 | 1 | 6 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7227 | 0.99998 | 99.8999 | 0.14657 | -0.006543 | 0.14672 | 146.449 | 0.998 | 0.14629 | -0.2737 | -0.001979 | -0.000283 | -0.1865 | -0.1979 | -0.1929
voltage_sweep | 4 | 1 | 7 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6843 | 1.0002 | 99.9036 | 0.14657 | -0.0064373 | 0.14671 | 146.427 | 2.979 | 0.43627 | -0.25734 | 1.9788 | 0.2897 | -0.1754 | -0.7169 | -0.78163
voltage_sweep | 4 | 1 | 8 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6843 | 1.0002 | 99.9036 | 0.14657 | -0.0064373 | 0.14671 | 146.427 | 1 | 0.14645 | -0.25734 | -0.00017 | -0.000119 | -0.1754 | -0.017 | -0.0812
voltage_sweep | 4 | 1 | 9 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6843 | 1.0002 | 99.9036 | 0.14657 | -0.0064373 | 0.14671 | 146.427 | 1 | 0.14657 | -0.25734 | -0.00017 | -1e-06 | -0.1754 | -0.017 | -0.0007
voltage_sweep | 4 | 1 | 10 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6843 | 1.0002 | 99.9036 | 0.14657 | -0.0064373 | 0.14671 | 146.427 | 1 | 0.14653 | -0.25734 | -0.00017 | -3.6e-05 | -0.1754 | -0.017 | -0.0245
voltage_sweep | 4 | 1 | 11 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6843 | 1.0002 | 99.9036 | 0.14657 | -0.0064373 | 0.14671 | 146.427 | 1 | 0.14644 | -0.25734 | -0.00017 | -0.000132 | -0.1754 | -0.017 | -0.09
voltage_sweep | 4 | 1 | 12 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6843 | 1.0002 | 99.9036 | 0.14657 | -0.0064373 | 0.14671 | 146.427 | 1 | 0.14648 | -0.25734 | -0.00017 | -9.2e-05 | -0.1754 | -0.017 | -0.0627
voltage_sweep | 4 | 1 | 13 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6838 | 0.99992 | 99.9008 | 0.14653 | -0.006531 | 0.14667 | 146.452 | 1 | 0.14643 | -0.23178 | 8.3e-05 | -9.2e-05 | -0.158 | 0.0083 | -0.0629
voltage_sweep | 4 | 1 | 14 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6838 | 0.99992 | 99.9008 | 0.14653 | -0.006531 | 0.14667 | 146.452 | 0.999 | 0.14639 | -0.23178 | -0.000917 | -0.000139 | -0.158 | -0.0917 | -0.095
voltage_sweep | 4 | 1 | 15 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6838 | 0.99992 | 99.9008 | 0.14653 | -0.006531 | 0.14667 | 146.452 | 1.001 | 0.1466 | -0.23178 | 0.001083 | 7.5e-05 | -0.158 | 0.1083 | 0.0511
voltage_sweep | 4 | 1 | 16 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6838 | 0.99992 | 99.9008 | 0.14653 | -0.006531 | 0.14667 | 146.452 | 0.999 | 0.14638 | -0.23178 | -0.000917 | -0.000147 | -0.158 | -0.0917 | -0.1004
voltage_sweep | 4 | 1 | 17 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6838 | 0.99992 | 99.9008 | 0.14653 | -0.006531 | 0.14667 | 146.452 | 1 | 0.14647 | -0.23178 | 8.3e-05 | -5.6e-05 | -0.158 | 0.0083 | -0.0383
voltage_sweep | 4 | 1 | 18 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6838 | 0.99992 | 99.9008 | 0.14653 | -0.006531 | 0.14667 | 146.452 | 1 | 0.14648 | -0.23178 | 8.3e-05 | -4.7e-05 | -0.158 | 0.0083 | -0.0322
voltage_sweep | 4 | 2 | 1 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7092 | 0.99998 | 99.9002 | 0.14656 | -0.0065361 | 0.14671 | 146.378 | 2.97 | 0.43487 | -0.33123 | 1.97 | 0.28831 | -0.2258 | -0.99817 | -1.0947
voltage_sweep | 4 | 2 | 2 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7092 | 0.99998 | 99.9002 | 0.14656 | -0.0065361 | 0.14671 | 146.378 | 1.001 | 0.14666 | -0.33123 | 0.001018 | 9.5e-05 | -0.2258 | 0.1018 | 0.0647
voltage_sweep | 4 | 2 | 3 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7092 | 0.99998 | 99.9002 | 0.14656 | -0.0065361 | 0.14671 | 146.378 | 1.002 | 0.14668 | -0.33123 | 0.002018 | 0.000116 | -0.2258 | 0.2019 | 0.0791
voltage_sweep | 4 | 2 | 4 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7092 | 0.99998 | 99.9002 | 0.14656 | -0.0065361 | 0.14671 | 146.378 | 1 | 0.14651 | -0.33123 | 1.8e-05 | -4.5e-05 | -0.2258 | 0.0018 | -0.0308
voltage_sweep | 4 | 2 | 5 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7092 | 0.99998 | 99.9002 | 0.14656 | -0.0065361 | 0.14671 | 146.378 | 1.001 | 0.1466 | -0.33123 | 0.001018 | 4.4e-05 | -0.2258 | 0.1018 | 0.0299
voltage_sweep | 4 | 2 | 6 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7092 | 0.99998 | 99.9002 | 0.14656 | -0.0065361 | 0.14671 | 146.378 | 0.998 | 0.14618 | -0.33123 | -0.001982 | -0.000376 | -0.2258 | -0.1982 | -0.2566
voltage_sweep | 4 | 2 | 7 | 2 | 146.7 | 1 | 0.999 | 60 | 146.679 | 1.0002 | 99.9037 | 0.14656 | -0.0064352 | 0.1467 | 146.437 | 2.979 | 0.43626 | -0.24196 | 1.9788 | 0.28969 | -0.165 | -0.71735 | -0.78055
voltage_sweep | 4 | 2 | 8 | 2 | 146.7 | 1 | 0.999 | 60 | 146.679 | 1.0002 | 99.9037 | 0.14656 | -0.0064352 | 0.1467 | 146.437 | 1 | 0.14642 | -0.24196 | -0.000175 | -0.000144 | -0.165 | -0.0175 | -0.0985
voltage_sweep | 4 | 2 | 9 | 2 | 146.7 | 1 | 0.999 | 60 | 146.679 | 1.0002 | 99.9037 | 0.14656 | -0.0064352 | 0.1467 | 146.437 | 1 | 0.14657 | -0.24196 | -0.000175 | 4e-06 | -0.165 | -0.0175 | 0.0025
voltage_sweep | 4 | 2 | 10 | 2 | 146.7 | 1 | 0.999 | 60 | 146.679 | 1.0002 | 99.9037 | 0.14656 | -0.0064352 | 0.1467 | 146.437 | 1 | 0.14652 | -0.24196 | -0.000175 | -4.3e-05 | -0.165 | -0.0175 | -0.0296
voltage_sweep | 4 | 2 | 11 | 2 | 146.7 | 1 | 0.999 | 60 | 146.679 | 1.0002 | 99.9037 | 0.14656 | -0.0064352 | 0.1467 | 146.437 | 1 | 0.14642 | -0.24196 | -0.000175 | -0.000138 | -0.165 | -0.0175 | -0.0944
voltage_sweep | 4 | 2 | 12 | 2 | 146.7 | 1 | 0.999 | 60 | 146.679 | 1.0002 | 99.9037 | 0.14656 | -0.0064352 | 0.1467 | 146.437 | 1 | 0.14651 | -0.24196 | -0.000175 | -5.6e-05 | -0.165 | -0.0175 | -0.0384
voltage_sweep | 4 | 2 | 13 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6892 | 0.99991 | 99.9008 | 0.14653 | -0.0065302 | 0.14668 | 146.451 | 1 | 0.14645 | -0.23821 | 8.6e-05 | -8.1e-05 | -0.1624 | 0.0086 | -0.0554
voltage_sweep | 4 | 2 | 14 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6892 | 0.99991 | 99.9008 | 0.14653 | -0.0065302 | 0.14668 | 146.451 | 0.999 | 0.1464 | -0.23821 | -0.000914 | -0.000128 | -0.1624 | -0.0914 | -0.0874
voltage_sweep | 4 | 2 | 15 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6892 | 0.99991 | 99.9008 | 0.14653 | -0.0065302 | 0.14668 | 146.451 | 1.001 | 0.1466 | -0.23821 | 0.001086 | 6.7e-05 | -0.1624 | 0.1086 | 0.0457
voltage_sweep | 4 | 2 | 16 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6892 | 0.99991 | 99.9008 | 0.14653 | -0.0065302 | 0.14668 | 146.451 | 0.999 | 0.14638 | -0.23821 | -0.000914 | -0.000149 | -0.1624 | -0.0914 | -0.1018
voltage_sweep | 4 | 2 | 17 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6892 | 0.99991 | 99.9008 | 0.14653 | -0.0065302 | 0.14668 | 146.451 | 1 | 0.14648 | -0.23821 | 8.6e-05 | -5.1e-05 | -0.1624 | 0.0086 | -0.0349
voltage_sweep | 4 | 2 | 18 | 3 | 146.7 | 1 | 0.999 | 60 | 146.6892 | 0.99991 | 99.9008 | 0.14653 | -0.0065302 | 0.14668 | 146.451 | 1 | 0.14647 | -0.23821 | 8.6e-05 | -5.9e-05 | -0.1624 | 0.0086 | -0.0403
voltage_sweep | 4 | 3 | 1 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7156 | 0.99999 | 99.9004 | 0.14657 | -0.0065389 | 0.14671 | 146.411 | 2.97 | 0.43489 | -0.30462 | 1.97 | 0.28832 | -0.2076 | -0.9991 | -1.0951
voltage_sweep | 4 | 3 | 2 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7156 | 0.99999 | 99.9004 | 0.14657 | -0.0065389 | 0.14671 | 146.411 | 1.001 | 0.14668 | -0.30462 | 0.001009 | 0.000114 | -0.2076 | 0.1009 | 0.0777
voltage_sweep | 4 | 3 | 3 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7156 | 0.99999 | 99.9004 | 0.14657 | -0.0065389 | 0.14671 | 146.411 | 1.001 | 0.14669 | -0.30462 | 0.001009 | 0.000126 | -0.2076 | 0.1009 | 0.0859
voltage_sweep | 4 | 3 | 4 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7156 | 0.99999 | 99.9004 | 0.14657 | -0.0065389 | 0.14671 | 146.411 | 1 | 0.14652 | -0.30462 | 9e-06 | -5e-05 | -0.2076 | 0.0009 | -0.0342
voltage_sweep | 4 | 3 | 5 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7156 | 0.99999 | 99.9004 | 0.14657 | -0.0065389 | 0.14671 | 146.411 | 1.001 | 0.14668 | -0.30462 | 0.001009 | 0.000108 | -0.2076 | 0.1009 | 0.0736
voltage_sweep | 4 | 3 | 6 | 1 | 146.7 | 1 | 0.999 | 60 | 146.7156 | 0.99999 | 99.9004 | 0.14657 | -0.0065389 | 0.14671 | 146.411 | 0.998 | 0.14624 | -0.30462 | -0.001991 | -0.000328 | -0.2076 | -0.1991 | -0.2239
voltage_sweep | 4 | 3 | 7 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6791 | 1.0001 | 99.9036 | 0.14656 | -0.0064379 | 0.1467 | 146.441 | 2.979 | 0.43634 | -0.23814 | 1.9789 | 0.28978 | -0.1624 | -0.71431 | -0.75847
voltage_sweep | 4 | 3 | 8 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6791 | 1.0001 | 99.9036 | 0.14656 | -0.0064379 | 0.1467 | 146.441 | 1 | 0.14646 | -0.23814 | -0.000144 | -0.000101 | -0.1624 | -0.0144 | -0.0689
voltage_sweep | 4 | 3 | 9 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6791 | 1.0001 | 99.9036 | 0.14656 | -0.0064379 | 0.1467 | 146.441 | 1.001 | 0.14659 | -0.23814 | 0.000856 | 2.7e-05 | -0.1624 | 0.0856 | 0.0185
voltage_sweep | 4 | 3 | 10 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6791 | 1.0001 | 99.9036 | 0.14656 | -0.0064379 | 0.1467 | 146.441 | 1 | 0.14653 | -0.23814 | -0.000144 | -3e-05 | -0.1624 | -0.0144 | -0.0204
voltage_sweep | 4 | 3 | 11 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6791 | 1.0001 | 99.9036 | 0.14656 | -0.0064379 | 0.1467 | 146.441 | 1 | 0.14643 | -0.23814 | -0.000144 | -0.000126 | -0.1624 | -0.0144 | -0.0859
voltage_sweep | 4 | 3 | 12 | 2 | 146.7 | 1 | 0.999 | 60 | 146.6791 | 1.0001 | 99.9036 | 0.14656 | -0.0064379 | 0.1467 | 146.441 | 1 | 0.1465 | -0.23814 | -0.000144 | -5.4e-05 | -0.1624 | -0.0144 | -0.0368
voltage_sweep | 4 | 3 | 13 | 3 | 146.7 | 1 | 0.999 | 60 | 146.681 | 0.99989 | 99.9008 | 0.14652 | -0.0065314 | 0.14667 | 146.428 | 1 | 0.14643 | -0.25305 | 0.000106 | -9.2e-05 | -0.1725 | 0.0106 | -0.0628
voltage_sweep | 4 | 3 | 14 | 3 | 146.7 | 1 | 0.999 | 60 | 146.681 | 0.99989 | 99.9008 | 0.14652 | -0.0065314 | 0.14667 | 146.428 | 0.999 | 0.14642 | -0.25305 | -0.000894 | -0.000104 | -0.1725 | -0.0894 | -0.071
voltage_sweep | 4 | 3 | 15 | 3 | 146.7 | 1 | 0.999 | 60 | 146.681 | 0.99989 | 99.9008 | 0.14652 | -0.0065314 | 0.14667 | 146.428 | 1.001 | 0.14659 | -0.25305 | 0.001106 | 7.1e-05 | -0.1725 | 0.1106 | 0.0485
voltage_sweep | 4 | 3 | 16 | 3 | 146.7 | 1 | 0.999 | 60 | 146.681 | 0.99989 | 99.9008 | 0.14652 | -0.0065314 | 0.14667 | 146.428 | 0.999 | 0.14638 | -0.25305 | -0.000894 | -0.000141 | -0.1725 | -0.0894 | -0.0962
voltage_sweep | 4 | 3 | 17 | 3 | 146.7 | 1 | 0.999 | 60 | 146.681 | 0.99989 | 99.9008 | 0.14652 | -0.0065314 | 0.14667 | 146.428 | 1 | 0.14648 | -0.25305 | 0.000106 | -4.4e-05 | -0.1725 | 0.0106 | -0.03
voltage_sweep | 4 | 3 | 18 | 3 | 146.7 | 1 | 0.999 | 60 | 146.681 | 0.99989 | 99.9008 | 0.14652 | -0.0065314 | 0.14667 | 146.428 | 1 | 0.14645 | -0.25305 | 0.000106 | -6.9e-05 | -0.1725 | 0.0106 | -0.0471
voltage_sweep | 5 | 1 | 1 | 1 | 162.2 | 1 | 0.999 | 60 | 162.242 | 0.99997 | 99.9005 | 0.16208 | -0.0072298 | 0.16224 | 161.924 | 2.97 | 0.48101 | -0.31804 | 1.97 | 0.31894 | -0.196 | -0.99727 | -1.0724
voltage_sweep | 5 | 1 | 2 | 1 | 162.2 | 1 | 0.999 | 60 | 162.242 | 0.99997 | 99.9005 | 0.16208 | -0.0072298 | 0.16224 | 161.924 | 1.001 | 0.16221 | -0.31804 | 0.001028 | 0.000133 | -0.196 | 0.1028 | 0.082
voltage_sweep | 5 | 1 | 3 | 1 | 162.2 | 1 | 0.999 | 60 | 162.242 | 0.99997 | 99.9005 | 0.16208 | -0.0072298 | 0.16224 | 161.924 | 1.001 | 0.16221 | -0.31804 | 0.001028 | 0.000134 | -0.196 | 0.1028 | 0.0826
voltage_sweep | 5 | 1 | 4 | 1 | 162.2 | 1 | 0.999 | 60 | 162.242 | 0.99997 | 99.9005 | 0.16208 | -0.0072298 | 0.16224 | 161.924 | 1 | 0.16209 | -0.31804 | 2.8e-05 | 1.2e-05 | -0.196 | 0.0028 | 0.0073
voltage_sweep | 5 | 1 | 5 | 1 | 162.2 | 1 | 0.999 | 60 | 162.242 | 0.99997 | 99.9005 | 0.16208 | -0.0072298 | 0.16224 | 161.924 | 1.001 | 0.16219 | -0.31804 | 0.001028 | 0.000111 | -0.196 | 0.1028 | 0.0684
voltage_sweep | 5 | 1 | 6 | 1 | 162.2 | 1 | 0.999 | 60 | 162.242 | 0.99997 | 99.9005 | 0.16208 | -0.0072298 | 0.16224 | 161.924 | 0.999 | 0.16176 | -0.31804 | -0.000972 | -0.000314 | -0.196 | -0.0972 | -0.1938
voltage_sweep | 5 | 1 | 7 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1793 | 1.0002 | 99.9037 | 0.16205 | -0.0071148 | 0.1622 | 161.895 | 2.979 | 0.48236 | -0.28432 | 1.9788 | 0.32031 | -0.1753 | -0.7154 | -0.77837
voltage_sweep | 5 | 1 | 8 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1793 | 1.0002 | 99.9037 | 0.16205 | -0.0071148 | 0.1622 | 161.895 | 0.999 | 0.16189 | -0.28432 | -0.001155 | -0.000154 | -0.1753 | -0.1155 | -0.0952
voltage_sweep | 5 | 1 | 9 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1793 | 1.0002 | 99.9037 | 0.16205 | -0.0071148 | 0.1622 | 161.895 | 1.001 | 0.1621 | -0.28432 | 0.000845 | 5.2e-05 | -0.1753 | 0.0845 | 0.0319
voltage_sweep | 5 | 1 | 10 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1793 | 1.0002 | 99.9037 | 0.16205 | -0.0071148 | 0.1622 | 161.895 | 1 | 0.16199 | -0.28432 | -0.000155 | -5.8e-05 | -0.1753 | -0.0155 | -0.036
voltage_sweep | 5 | 1 | 11 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1793 | 1.0002 | 99.9037 | 0.16205 | -0.0071148 | 0.1622 | 161.895 | 1 | 0.1619 | -0.28432 | -0.000155 | -0.000143 | -0.1753 | -0.0155 | -0.0885
voltage_sweep | 5 | 1 | 12 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1793 | 1.0002 | 99.9037 | 0.16205 | -0.0071148 | 0.1622 | 161.895 | 1 | 0.162 | -0.28432 | -0.000155 | -4.8e-05 | -0.1753 | -0.0155 | -0.0298
voltage_sweep | 5 | 1 | 13 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1861 | 0.99992 | 99.9008 | 0.16201 | -0.0072227 | 0.16217 | 161.937 | 1 | 0.16194 | -0.24911 | 7.8e-05 | -7.6e-05 | -0.1536 | 0.0078 | -0.0472
voltage_sweep | 5 | 1 | 14 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1861 | 0.99992 | 99.9008 | 0.16201 | -0.0072227 | 0.16217 | 161.937 | 0.999 | 0.1619 | -0.24911 | -0.000922 | -0.000111 | -0.1536 | -0.0922 | -0.0688
voltage_sweep | 5 | 1 | 15 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1861 | 0.99992 | 99.9008 | 0.16201 | -0.0072227 | 0.16217 | 161.937 | 1 | 0.16207 | -0.24911 | 7.8e-05 | 5.4e-05 | -0.1536 | 0.0078 | 0.033
voltage_sweep | 5 | 1 | 16 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1861 | 0.99992 | 99.9008 | 0.16201 | -0.0072227 | 0.16217 | 161.937 | 0.999 | 0.16187 | -0.24911 | -0.000922 | -0.000137 | -0.1536 | -0.0922 | -0.0849
voltage_sweep | 5 | 1 | 17 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1861 | 0.99992 | 99.9008 | 0.16201 | -0.0072227 | 0.16217 | 161.937 | 1 | 0.16197 | -0.24911 | 7.8e-05 | -4.4e-05 | -0.1536 | 0.0078 | -0.0275
voltage_sweep | 5 | 1 | 18 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1861 | 0.99992 | 99.9008 | 0.16201 | -0.0072227 | 0.16217 | 161.937 | 1 | 0.16192 | -0.24911 | 7.8e-05 | -8.8e-05 | -0.1536 | 0.0078 | -0.0546
voltage_sweep | 5 | 2 | 1 | 1 | 162.2 | 1 | 0.999 | 60 | 162.2028 | 0.99997 | 99.9006 | 0.16204 | -0.0072282 | 0.1622 | 161.898 | 2.97 | 0.48093 | -0.30484 | 1.97 | 0.31889 | -0.1879 | -0.99665 | -1.0653
voltage_sweep | 5 | 2 | 2 | 1 | 162.2 | 1 | 0.999 | 60 | 162.2028 | 0.99997 | 99.9006 | 0.16204 | -0.0072282 | 0.1622 | 161.898 | 1.001 | 0.16219 | -0.30484 | 0.001034 | 0.000158 | -0.1879 | 0.1034 | 0.0974
voltage_sweep | 5 | 2 | 3 | 1 | 162.2 | 1 | 0.999 | 60 | 162.2028 | 0.99997 | 99.9006 | 0.16204 | -0.0072282 | 0.1622 | 161.898 | 1.002 | 0.1622 | -0.30484 | 0.002034 | 0.000167 | -0.1879 | 0.2034 | 0.103
voltage_sweep | 5 | 2 | 4 | 1 | 162.2 | 1 | 0.999 | 60 | 162.2028 | 0.99997 | 99.9006 | 0.16204 | -0.0072282 | 0.1622 | 161.898 | 1 | 0.16207 | -0.30484 | 3.4e-05 | 3.1e-05 | -0.1879 | 0.0034 | 0.0191
voltage_sweep | 5 | 2 | 5 | 1 | 162.2 | 1 | 0.999 | 60 | 162.2028 | 0.99997 | 99.9006 | 0.16204 | -0.0072282 | 0.1622 | 161.898 | 1.001 | 0.16215 | -0.30484 | 0.001034 | 0.000114 | -0.1879 | 0.1034 | 0.0703
voltage_sweep | 5 | 2 | 6 | 1 | 162.2 | 1 | 0.999 | 60 | 162.2028 | 0.99997 | 99.9006 | 0.16204 | -0.0072282 | 0.1622 | 161.898 | 0.998 | 0.16169 | -0.30484 | -0.001966 | -0.000347 | -0.1879 | -0.1966 | -0.2142
voltage_sweep | 5 | 2 | 7 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1803 | 1.0002 | 99.9036 | 0.16205 | -0.0071207 | 0.16221 | 161.895 | 2.978 | 0.48228 | -0.28528 | 1.9778 | 0.32023 | -0.1759 | -0.75122 | -0.79776
voltage_sweep | 5 | 2 | 8 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1803 | 1.0002 | 99.9036 | 0.16205 | -0.0071207 | 0.16221 | 161.895 | 0.999 | 0.16189 | -0.28528 | -0.00118 | -0.000161 | -0.1759 | -0.118 | -0.0994
voltage_sweep | 5 | 2 | 9 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1803 | 1.0002 | 99.9036 | 0.16205 | -0.0071207 | 0.16221 | 161.895 | 1.001 | 0.16206 | -0.28528 | 0.00082 | 9e-06 | -0.1759 | 0.082 | 0.0055
voltage_sweep | 5 | 2 | 10 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1803 | 1.0002 | 99.9036 | 0.16205 | -0.0071207 | 0.16221 | 161.895 | 1 | 0.16198 | -0.28528 | -0.00018 | -7.2e-05 | -0.1759 | -0.018 | -0.0445
voltage_sweep | 5 | 2 | 11 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1803 | 1.0002 | 99.9036 | 0.16205 | -0.0071207 | 0.16221 | 161.895 | 1 | 0.16193 | -0.28528 | -0.00018 | -0.000123 | -0.1759 | -0.018 | -0.076
voltage_sweep | 5 | 2 | 12 | 2 | 162.2 | 1 | 0.999 | 60 | 162.1803 | 1.0002 | 99.9036 | 0.16205 | -0.0071207 | 0.16221 | 161.895 | 1 | 0.16199 | -0.28528 | -0.00018 | -6.2e-05 | -0.1759 | -0.018 | -0.0383
voltage_sweep | 5 | 2 | 13 | 3 | 162.2 | 1 | 0.999 | 60 | 162.184 | 0.9999 | 99.9007 | 0.16201 | -0.0072234 | 0.16217 | 161.933 | 1 | 0.16196 | -0.25102 | 9.9e-05 | -4.9e-05 | -0.1548 | 0.0099 | -0.0302
voltage_sweep | 5 | 2 | 14 | 3 | 162.2 | 1 | 0.999 | 60 | 162.184 | 0.9999 | 99.9007 | 0.16201 | -0.0072234 | 0.16217 | 161.933 | 0.999 | 0.1619 | -0.25102 | -0.000901 | -0.000105 | -0.1548 | -0.0901 | -0.0648
voltage_sweep | 5 | 2 | 15 | 3 | 162.2 | 1 | 0.999 | 60 | 162.184 | 0.9999 | 99.9007 | 0.16201 | -0.0072234 | 0.16217 | 161.933 | 1 | 0.16206 | -0.25102 | 9.9e-05 | 5.6e-05 | -0.1548 | 0.0099 | 0.0346
voltage_sweep | 5 | 2 | 16 | 3 | 162.2 | 1 | 0.999 | 60 | 162.184 | 0.9999 | 99.9007 | 0.16201 | -0.0072234 | 0.16217 | 161.933 | 0.999 | 0.16188 | -0.25102 | -0.000901 | -0.000126 | -0.1548 | -0.0901 | -0.0777
voltage_sweep | 5 | 2 | 17 | 3 | 162.2 | 1 | 0.999 | 60 | 162.184 | 0.9999 | 99.9007 | 0.16201 | -0.0072234 | 0.16217 | 161.933 | 1 | 0.16196 | -0.25102 | 9.9e-05 | -4.5e-05 | -0.1548 | 0.0099 | -0.0277
voltage_sweep | 5 | 2 | 18 | 3 | 162.2 | 1 | 0.999 | 60 | 162.184 | 0.9999 | 99.9007 | 0.16201 | -0.0072234 | 0.16217 | 161.933 | 1 | 0.16193 | -0.25102 | 9.9e-05 | -8.1e-05 | -0.1548 | 0.0099 | -0.05
voltage_sweep | 5 | 3 | 1 | 1 | 162.2 | 1 | 0.999 | 60 | 162.1986 | 0.99998 | 99.9004 | 0.16203 | -0.0072324 | 0.16219 | 161.864 | 2.97 | 0.48095 | -0.33456 | 1.97 | 0.31892 | -0.2063 | -0.99753 | -1.0584
voltage_sweep | 5 | 3 | 2 | 1 | 162.2 | 1 | 0.999 | 60 | 162.1986 | 0.99998 | 99.9004 | 0.16203 | -0.0072324 | 0.16219 | 161.864 | 1.001 | 0.16212 | -0.33456 | 0.001025 | 8.4e-05 | -0.2063 | 0.1025 | 0.0519
voltage_sweep | 5 | 3 | 3 | 1 | 162.2 | 1 | 0.999 | 60 | 162.1986 | 0.99998 | 99.9004 | 0.16203 | -0.0072324 | 0.16219 | 161.864 | 1.001 | 0.16225 | -0.33456 | 0.001025 | 0.000215 | -0.2063 | 0.1025 | 0.1327
voltage_sweep | 5 | 3 | 4 | 1 | 162.2 | 1 | 0.999 | 60 | 162.1986 | 0.99998 | 99.9004 | 0.16203 | -0.0072324 | 0.16219 | 161.864 | 1 | 0.16206 | -0.33456 | 2.5e-05 | 2.8e-05 | -0.2063 | 0.0025 | 0.0173
voltage_sweep | 5 | 3 | 5 | 1 | 162.2 | 1 | 0.999 | 60 | 162.1986 | 0.99998 | 99.9004 | 0.16203 | -0.0072324 | 0.16219 | 161.864 | 1.001 | 0.16215 | -0.33456 | 0.001025 | 0.000114 | -0.2063 | 0.1025 | 0.0704
voltage_sweep | 5 | 3 | 6 | 1 | 162.2 | 1 | 0.999 | 60 | 162.1986 | 0.99998 | 99.9004 | 0.16203 | -0.0072324 | 0.16219 | 161.864 | 0.999 | 0.16179 | -0.33456 | -0.000975 | -0.000242 | -0.2063 | -0.0975 | -0.1493
voltage_sweep | 5 | 3 | 7 | 2 | 162.2 | 1 | 0.999 | 60 | 162.179 | 1.0002 | 99.9038 | 0.16205 | -0.0071135 | 0.1622 | 161.913 | 2.979 | 0.48231 | -0.26595 | 1.9788 | 0.32026 | -0.164 | -0.71521 | -0.78808
voltage_sweep | 5 | 3 | 8 | 2 | 162.2 | 1 | 0.999 | 60 | 162.179 | 1.0002 | 99.9038 | 0.16205 | -0.0071135 | 0.1622 | 161.913 | 1 | 0.16194 | -0.26595 | -0.000153 | -0.000104 | -0.164 | -0.0153 | -0.064
voltage_sweep | 5 | 3 | 9 | 2 | 162.2 | 1 | 0.999 | 60 | 162.179 | 1.0002 | 99.9038 | 0.16205 | -0.0071135 | 0.1622 | 161.913 | 1 | 0.16204 | -0.26595 | -0.000153 | -5e-06 | -0.164 | -0.0153 | -0.0029
voltage_sweep | 5 | 3 | 10 | 2 | 162.2 | 1 | 0.999 | 60 | 162.179 | 1.0002 | 99.9038 | 0.16205 | -0.0071135 | 0.1622 | 161.913 | 1 | 0.16199 | -0.26595 | -0.000153 | -5.6e-05 | -0.164 | -0.0153 | -0.0344
voltage_sweep | 5 | 3 | 11 | 2 | 162.2 | 1 | 0.999 | 60 | 162.179 | 1.0002 | 99.9038 | 0.16205 | -0.0071135 | 0.1622 | 161.913 | 1 | 0.16196 | -0.26595 | -0.000153 | -8.8e-05 | -0.164 | -0.0153 | -0.0541
voltage_sweep | 5 | 3 | 12 | 2 | 162.2 | 1 | 0.999 | 60 | 162.179 | 1.0002 | 99.9038 | 0.16205 | -0.0071135 | 0.1622 | 161.913 | 1 | 0.16195 | -0.26595 | -0.000153 | -9.9e-05 | -0.164 | -0.0153 | -0.0609
voltage_sweep | 5 | 3 | 13 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1843 | 0.99992 | 99.9007 | 0.16201 | -0.0072236 | 0.16217 | 161.922 | 1 | 0.16197 | -0.26234 | 8.5e-05 | -3.9e-05 | -0.1618 | 0.0085 | -0.0238
voltage_sweep | 5 | 3 | 14 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1843 | 0.99992 | 99.9007 | 0.16201 | -0.0072236 | 0.16217 | 161.922 | 0.999 | 0.16187 | -0.26234 | -0.000915 | -0.000137 | -0.1618 | -0.0915 | -0.0843
voltage_sweep | 5 | 3 | 15 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1843 | 0.99992 | 99.9007 | 0.16201 | -0.0072236 | 0.16217 | 161.922 | 1.001 | 0.1621 | -0.26234 | 0.001085 | 8.6e-05 | -0.1618 | 0.1085 | 0.0533
voltage_sweep | 5 | 3 | 16 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1843 | 0.99992 | 99.9007 | 0.16201 | -0.0072236 | 0.16217 | 161.922 | 0.999 | 0.16192 | -0.26234 | -0.000915 | -9.2e-05 | -0.1618 | -0.0915 | -0.0565
voltage_sweep | 5 | 3 | 17 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1843 | 0.99992 | 99.9007 | 0.16201 | -0.0072236 | 0.16217 | 161.922 | 1 | 0.16193 | -0.26234 | 8.5e-05 | -8e-05 | -0.1618 | 0.0085 | -0.0491
voltage_sweep | 5 | 3 | 18 | 3 | 162.2 | 1 | 0.999 | 60 | 162.1843 | 0.99992 | 99.9007 | 0.16201 | -0.0072236 | 0.16217 | 161.922 | 1 | 0.16195 | -0.26234 | 8.5e-05 | -6e-05 | -0.1618 | 0.0085 | -0.0368
voltage_sweep | 6 | 1 | 1 | 1 | 177.8 | 1 | 0.999 | 60 | 177.825 | 1 | 99.9005 | 0.17765 | -0.0079281 | 0.17782 | 177.524 | 2.969 | 0.52726 | -0.30101 | 1.969 | 0.34961 | -0.1693 | -1.0331 | -1.066
voltage_sweep | 6 | 1 | 2 | 1 | 177.8 | 1 | 0.999 | 60 | 177.825 | 1 | 99.9005 | 0.17765 | -0.0079281 | 0.17782 | 177.524 | 1.001 | 0.17785 | -0.30101 | 0.001002 | 0.000201 | -0.1693 | 0.1002 | 0.1133
voltage_sweep | 6 | 1 | 3 | 1 | 177.8 | 1 | 0.999 | 60 | 177.825 | 1 | 99.9005 | 0.17765 | -0.0079281 | 0.17782 | 177.524 | 1.002 | 0.17791 | -0.30101 | 0.002002 | 0.000257 | -0.1693 | 0.2002 | 0.1448
voltage_sweep | 6 | 1 | 4 | 1 | 177.8 | 1 | 0.999 | 60 | 177.825 | 1 | 99.9005 | 0.17765 | -0.0079281 | 0.17782 | 177.524 | 1 | 0.17768 | -0.30101 | 2e-06 | 2.8e-05 | -0.1693 | 0.0002 | 0.0159
voltage_sweep | 6 | 1 | 5 | 1 | 177.8 | 1 | 0.999 | 60 | 177.825 | 1 | 99.9005 | 0.17765 | -0.0079281 | 0.17782 | 177.524 | 1.001 | 0.17779 | -0.30101 | 0.001002 | 0.000139 | -0.1693 | 0.1002 | 0.0784
voltage_sweep | 6 | 1 | 6 | 1 | 177.8 | 1 | 0.999 | 60 | 177.825 | 1 | 99.9005 | 0.17765 | -0.0079281 | 0.17782 | 177.524 | 0.998 | 0.17733 | -0.30101 | -0.001998 | -0.000319 | -0.1693 | -0.1998 | -0.1794
voltage_sweep | 6 | 1 | 7 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7774 | 1.0002 | 99.9037 | 0.17764 | -0.0078013 | 0.17781 | 177.469 | 2.979 | 0.52883 | -0.3084 | 1.9788 | 0.3512 | -0.1735 | -0.71651 | -0.76451
voltage_sweep | 6 | 1 | 8 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7774 | 1.0002 | 99.9037 | 0.17764 | -0.0078013 | 0.17781 | 177.469 | 1 | 0.17749 | -0.3084 | -0.000166 | -0.000145 | -0.1735 | -0.0166 | -0.0815
voltage_sweep | 6 | 1 | 9 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7774 | 1.0002 | 99.9037 | 0.17764 | -0.0078013 | 0.17781 | 177.469 | 1 | 0.17763 | -0.3084 | -0.000166 | -7e-06 | -0.1735 | -0.0166 | -0.0038
voltage_sweep | 6 | 1 | 10 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7774 | 1.0002 | 99.9037 | 0.17764 | -0.0078013 | 0.17781 | 177.469 | 1 | 0.17763 | -0.3084 | -0.000166 | -8e-06 | -0.1735 | -0.0166 | -0.0043
voltage_sweep | 6 | 1 | 11 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7774 | 1.0002 | 99.9037 | 0.17764 | -0.0078013 | 0.17781 | 177.469 | 1 | 0.17753 | -0.3084 | -0.000166 | -0.000105 | -0.1735 | -0.0166 | -0.0589
voltage_sweep | 6 | 1 | 12 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7774 | 1.0002 | 99.9037 | 0.17764 | -0.0078013 | 0.17781 | 177.469 | 1 | 0.17754 | -0.3084 | -0.000166 | -9.7e-05 | -0.1735 | -0.0166 | -0.0544
voltage_sweep | 6 | 1 | 13 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7858 | 0.99993 | 99.9009 | 0.1776 | -0.0079132 | 0.17777 | 177.504 | 1 | 0.17752 | -0.28183 | 7.4e-05 | -7.9e-05 | -0.1585 | 0.0074 | -0.0447
voltage_sweep | 6 | 1 | 14 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7858 | 0.99993 | 99.9009 | 0.1776 | -0.0079132 | 0.17777 | 177.504 | 0.999 | 0.17743 | -0.28183 | -0.000926 | -0.000166 | -0.1585 | -0.0926 | -0.0937
voltage_sweep | 6 | 1 | 15 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7858 | 0.99993 | 99.9009 | 0.1776 | -0.0079132 | 0.17777 | 177.504 | 1 | 0.17767 | -0.28183 | 7.4e-05 | 7.9e-05 | -0.1585 | 0.0074 | 0.0443
voltage_sweep | 6 | 1 | 16 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7858 | 0.99993 | 99.9009 | 0.1776 | -0.0079132 | 0.17777 | 177.504 | 0.999 | 0.17746 | -0.28183 | -0.000926 | -0.000138 | -0.1585 | -0.0926 | -0.0779
voltage_sweep | 6 | 1 | 17 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7858 | 0.99993 | 99.9009 | 0.1776 | -0.0079132 | 0.17777 | 177.504 | 1 | 0.17751 | -0.28183 | 7.4e-05 | -9e-05 | -0.1585 | 0.0074 | -0.0509
voltage_sweep | 6 | 1 | 18 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7858 | 0.99993 | 99.9009 | 0.1776 | -0.0079132 | 0.17777 | 177.504 | 1 | 0.17752 | -0.28183 | 7.4e-05 | -7.4e-05 | -0.1585 | 0.0074 | -0.0419
voltage_sweep | 6 | 2 | 1 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8389 | 0.99997 | 99.9007 | 0.17766 | -0.0079218 | 0.17783 | 177.544 | 2.97 | 0.52731 | -0.29488 | 1.97 | 0.34966 | -0.1658 | -0.99714 | -1.0615
voltage_sweep | 6 | 2 | 2 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8389 | 0.99997 | 99.9007 | 0.17766 | -0.0079218 | 0.17783 | 177.544 | 1.001 | 0.1778 | -0.29488 | 0.001029 | 0.000143 | -0.1658 | 0.1029 | 0.0804
voltage_sweep | 6 | 2 | 3 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8389 | 0.99997 | 99.9007 | 0.17766 | -0.0079218 | 0.17783 | 177.544 | 1.002 | 0.17785 | -0.29488 | 0.002029 | 0.000197 | -0.1658 | 0.2029 | 0.1108
voltage_sweep | 6 | 2 | 4 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8389 | 0.99997 | 99.9007 | 0.17766 | -0.0079218 | 0.17783 | 177.544 | 1 | 0.1777 | -0.29488 | 2.9e-05 | 4.2e-05 | -0.1658 | 0.0029 | 0.0236
voltage_sweep | 6 | 2 | 5 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8389 | 0.99997 | 99.9007 | 0.17766 | -0.0079218 | 0.17783 | 177.544 | 1.001 | 0.17778 | -0.29488 | 0.001029 | 0.000119 | -0.1658 | 0.1029 | 0.0669
voltage_sweep | 6 | 2 | 6 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8389 | 0.99997 | 99.9007 | 0.17766 | -0.0079218 | 0.17783 | 177.544 | 0.998 | 0.17729 | -0.29488 | -0.001971 | -0.000369 | -0.1658 | -0.1971 | -0.2078
voltage_sweep | 6 | 2 | 7 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7764 | 1.0002 | 99.9037 | 0.17763 | -0.007801 | 0.17781 | 177.47 | 2.979 | 0.52877 | -0.30637 | 1.9788 | 0.35114 | -0.1723 | -0.71613 | -0.77482
voltage_sweep | 6 | 2 | 8 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7764 | 1.0002 | 99.9037 | 0.17763 | -0.007801 | 0.17781 | 177.47 | 0.999 | 0.17747 | -0.30637 | -0.001162 | -0.000167 | -0.1723 | -0.1162 | -0.094
voltage_sweep | 6 | 2 | 9 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7764 | 1.0002 | 99.9037 | 0.17763 | -0.007801 | 0.17781 | 177.47 | 1 | 0.17764 | -0.30637 | -0.000162 | 4e-06 | -0.1723 | -0.0162 | 0.0022
voltage_sweep | 6 | 2 | 10 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7764 | 1.0002 | 99.9037 | 0.17763 | -0.007801 | 0.17781 | 177.47 | 1 | 0.17757 | -0.30637 | -0.000162 | -6e-05 | -0.1723 | -0.0162 | -0.0338
voltage_sweep | 6 | 2 | 11 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7764 | 1.0002 | 99.9037 | 0.17763 | -0.007801 | 0.17781 | 177.47 | 1 | 0.1775 | -0.30637 | -0.000162 | -0.00013 | -0.1723 | -0.0162 | -0.0732
voltage_sweep | 6 | 2 | 12 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7764 | 1.0002 | 99.9037 | 0.17763 | -0.007801 | 0.17781 | 177.47 | 1 | 0.17752 | -0.30637 | -0.000162 | -0.000109 | -0.1723 | -0.0162 | -0.0614
voltage_sweep | 6 | 2 | 13 | 3 | 177.8 | 1 | 0.999 | 60 | 177.786 | 0.99992 | 99.9007 | 0.1776 | -0.0079199 | 0.17777 | 177.523 | 1 | 0.17756 | -0.26299 | 8.1e-05 | -3.2e-05 | -0.1479 | 0.0081 | -0.018
voltage_sweep | 6 | 2 | 14 | 3 | 177.8 | 1 | 0.999 | 60 | 177.786 | 0.99992 | 99.9007 | 0.1776 | -0.0079199 | 0.17777 | 177.523 | 0.999 | 0.17744 | -0.26299 | -0.000919 | -0.000158 | -0.1479 | -0.0919 | -0.089
voltage_sweep | 6 | 2 | 15 | 3 | 177.8 | 1 | 0.999 | 60 | 177.786 | 0.99992 | 99.9007 | 0.1776 | -0.0079199 | 0.17777 | 177.523 | 1.001 | 0.17769 | -0.26299 | 0.001081 | 0.0001 | -0.1479 | 0.1081 | 0.0563
voltage_sweep | 6 | 2 | 16 | 3 | 177.8 | 1 | 0.999 | 60 | 177.786 | 0.99992 | 99.9007 | 0.1776 | -0.0079199 | 0.17777 | 177.523 | 0.999 | 0.17748 | -0.26299 | -0.000919 | -0.000116 | -0.1479 | -0.0919 | -0.0653
voltage_sweep | 6 | 2 | 17 | 3 | 177.8 | 1 | 0.999 | 60 | 177.786 | 0.99992 | 99.9007 | 0.1776 | -0.0079199 | 0.17777 | 177.523 | 1 | 0.1775 | -0.26299 | 8.1e-05 | -9.4e-05 | -0.1479 | 0.0081 | -0.0529
voltage_sweep | 6 | 2 | 18 | 3 | 177.8 | 1 | 0.999 | 60 | 177.786 | 0.99992 | 99.9007 | 0.1776 | -0.0079199 | 0.17777 | 177.523 | 1 | 0.17751 | -0.26299 | 8.1e-05 | -8.6e-05 | -0.1479 | 0.0081 | -0.0484
voltage_sweep | 6 | 3 | 1 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8111 | 0.99996 | 99.9004 | 0.17763 | -0.0079315 | 0.1778 | 177.449 | 2.97 | 0.52718 | -0.36211 | 1.97 | 0.34956 | -0.2036 | -0.99621 | -1.069
voltage_sweep | 6 | 3 | 2 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8111 | 0.99996 | 99.9004 | 0.17763 | -0.0079315 | 0.1778 | 177.449 | 1.001 | 0.17776 | -0.36211 | 0.001038 | 0.000132 | -0.2036 | 0.1038 | 0.0742
voltage_sweep | 6 | 3 | 3 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8111 | 0.99996 | 99.9004 | 0.17763 | -0.0079315 | 0.1778 | 177.449 | 1.001 | 0.17788 | -0.36211 | 0.001038 | 0.000248 | -0.2036 | 0.1038 | 0.1395
voltage_sweep | 6 | 3 | 4 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8111 | 0.99996 | 99.9004 | 0.17763 | -0.0079315 | 0.1778 | 177.449 | 1 | 0.17766 | -0.36211 | 3.8e-05 | 3.6e-05 | -0.2036 | 0.0038 | 0.0202
voltage_sweep | 6 | 3 | 5 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8111 | 0.99996 | 99.9004 | 0.17763 | -0.0079315 | 0.1778 | 177.449 | 1.001 | 0.17771 | -0.36211 | 0.001038 | 8.3e-05 | -0.2036 | 0.1038 | 0.0466
voltage_sweep | 6 | 3 | 6 | 1 | 177.8 | 1 | 0.999 | 60 | 177.8111 | 0.99996 | 99.9004 | 0.17763 | -0.0079315 | 0.1778 | 177.449 | 0.999 | 0.17736 | -0.36211 | -0.000962 | -0.000263 | -0.2036 | -0.0962 | -0.1481
voltage_sweep | 6 | 3 | 7 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7806 | 1.0002 | 99.9038 | 0.17764 | -0.0077957 | 0.17781 | 177.486 | 2.979 | 0.52875 | -0.29461 | 1.9788 | 0.35112 | -0.1657 | -0.71509 | -0.78004
voltage_sweep | 6 | 3 | 8 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7806 | 1.0002 | 99.9038 | 0.17764 | -0.0077957 | 0.17781 | 177.486 | 0.999 | 0.17746 | -0.29461 | -0.001152 | -0.000175 | -0.1657 | -0.1152 | -0.0983
voltage_sweep | 6 | 3 | 9 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7806 | 1.0002 | 99.9038 | 0.17764 | -0.0077957 | 0.17781 | 177.486 | 1 | 0.17764 | -0.29461 | -0.000152 | 0 | -0.1657 | -0.0152 | 0.0002
voltage_sweep | 6 | 3 | 10 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7806 | 1.0002 | 99.9038 | 0.17764 | -0.0077957 | 0.17781 | 177.486 | 1 | 0.17757 | -0.29461 | -0.000152 | -6.4e-05 | -0.1657 | -0.0152 | -0.0358
voltage_sweep | 6 | 3 | 11 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7806 | 1.0002 | 99.9038 | 0.17764 | -0.0077957 | 0.17781 | 177.486 | 1 | 0.17751 | -0.29461 | -0.000152 | -0.000125 | -0.1657 | -0.0152 | -0.0702
voltage_sweep | 6 | 3 | 12 | 2 | 177.8 | 1 | 0.999 | 60 | 177.7806 | 1.0002 | 99.9038 | 0.17764 | -0.0077957 | 0.17781 | 177.486 | 1 | 0.17755 | -0.29461 | -0.000152 | -8.5e-05 | -0.1657 | -0.0152 | -0.0476
voltage_sweep | 6 | 3 | 13 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7793 | 0.99992 | 99.9008 | 0.17759 | -0.0079145 | 0.17776 | 177.515 | 1 | 0.17754 | -0.26434 | 8.1e-05 | -4.6e-05 | -0.1487 | 0.0081 | -0.0257
voltage_sweep | 6 | 3 | 14 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7793 | 0.99992 | 99.9008 | 0.17759 | -0.0079145 | 0.17776 | 177.515 | 0.999 | 0.17747 | -0.26434 | -0.000919 | -0.000121 | -0.1487 | -0.0919 | -0.0679
voltage_sweep | 6 | 3 | 15 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7793 | 0.99992 | 99.9008 | 0.17759 | -0.0079145 | 0.17776 | 177.515 | 1 | 0.17768 | -0.26434 | 8.1e-05 | 9.1e-05 | -0.1487 | 0.0081 | 0.0515
voltage_sweep | 6 | 3 | 16 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7793 | 0.99992 | 99.9008 | 0.17759 | -0.0079145 | 0.17776 | 177.515 | 0.999 | 0.17749 | -0.26434 | -0.000919 | -0.0001 | -0.1487 | -0.0919 | -0.0561
voltage_sweep | 6 | 3 | 17 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7793 | 0.99992 | 99.9008 | 0.17759 | -0.0079145 | 0.17776 | 177.515 | 1 | 0.17757 | -0.26434 | 8.1e-05 | -2e-05 | -0.1487 | 0.0081 | -0.011
voltage_sweep | 6 | 3 | 18 | 3 | 177.8 | 1 | 0.999 | 60 | 177.7793 | 0.99992 | 99.9008 | 0.17759 | -0.0079145 | 0.17776 | 177.515 | 1 | 0.17752 | -0.26434 | 8.1e-05 | -6.7e-05 | -0.1487 | 0.0081 | -0.0375
voltage_sweep | 7 | 1 | 1 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3218 | 0.99997 | 99.9001 | 0.19312 | -0.0086225 | 0.19332 | 193.023 | 2.97 | 0.57314 | -0.29878 | 1.97 | 0.38002 | -0.1545 | -0.99708 | -1.0749
voltage_sweep | 7 | 1 | 2 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3218 | 0.99997 | 99.9001 | 0.19312 | -0.0086225 | 0.19332 | 193.023 | 1.001 | 0.19342 | -0.29878 | 0.001029 | 0.000294 | -0.1545 | 0.1029 | 0.1523
voltage_sweep | 7 | 1 | 3 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3218 | 0.99997 | 99.9001 | 0.19312 | -0.0086225 | 0.19332 | 193.023 | 1.002 | 0.19341 | -0.29878 | 0.002029 | 0.00029 | -0.1545 | 0.203 | 0.1502
voltage_sweep | 7 | 1 | 4 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3218 | 0.99997 | 99.9001 | 0.19312 | -0.0086225 | 0.19332 | 193.023 | 1.001 | 0.19317 | -0.29878 | 0.001029 | 4.5e-05 | -0.1545 | 0.1029 | 0.0233
voltage_sweep | 7 | 1 | 5 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3218 | 0.99997 | 99.9001 | 0.19312 | -0.0086225 | 0.19332 | 193.023 | 1.001 | 0.1934 | -0.29878 | 0.001029 | 0.000279 | -0.1545 | 0.1029 | 0.1445
voltage_sweep | 7 | 1 | 6 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3218 | 0.99997 | 99.9001 | 0.19312 | -0.0086225 | 0.19332 | 193.023 | 0.998 | 0.19279 | -0.29878 | -0.001971 | -0.000337 | -0.1545 | -0.1971 | -0.1745
voltage_sweep | 7 | 1 | 7 | 2 | 193.3 | 1 | 0.999 | 60 | 193.276 | 1.0002 | 99.9038 | 0.19312 | -0.0084783 | 0.19331 | 192.938 | 2.979 | 0.57487 | -0.33803 | 1.9788 | 0.38174 | -0.1749 | -0.71649 | -0.77677
voltage_sweep | 7 | 1 | 8 | 2 | 193.3 | 1 | 0.999 | 60 | 193.276 | 1.0002 | 99.9038 | 0.19312 | -0.0084783 | 0.19331 | 192.938 | 0.999 | 0.19288 | -0.33803 | -0.001166 | -0.000239 | -0.1749 | -0.1166 | -0.1238
voltage_sweep | 7 | 1 | 9 | 2 | 193.3 | 1 | 0.999 | 60 | 193.276 | 1.0002 | 99.9038 | 0.19312 | -0.0084783 | 0.19331 | 192.938 | 1 | 0.19311 | -0.33803 | -0.000166 | -1.5e-05 | -0.1749 | -0.0166 | -0.0078
voltage_sweep | 7 | 1 | 10 | 2 | 193.3 | 1 | 0.999 | 60 | 193.276 | 1.0002 | 99.9038 | 0.19312 | -0.0084783 | 0.19331 | 192.938 | 1 | 0.19308 | -0.33803 | -0.000166 | -4.3e-05 | -0.1749 | -0.0166 | -0.0223
voltage_sweep | 7 | 1 | 11 | 2 | 193.3 | 1 | 0.999 | 60 | 193.276 | 1.0002 | 99.9038 | 0.19312 | -0.0084783 | 0.19331 | 192.938 | 1 | 0.19294 | -0.33803 | -0.000166 | -0.000185 | -0.1749 | -0.0166 | -0.0959
voltage_sweep | 7 | 1 | 12 | 2 | 193.3 | 1 | 0.999 | 60 | 193.276 | 1.0002 | 99.9038 | 0.19312 | -0.0084783 | 0.19331 | 192.938 | 1 | 0.19303 | -0.33803 | -0.000166 | -9.4e-05 | -0.1749 | -0.0166 | -0.0487
voltage_sweep | 7 | 1 | 13 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2843 | 0.99992 | 99.9008 | 0.19308 | -0.0086067 | 0.19327 | 193.007 | 0.999 | 0.19297 | -0.27729 | -0.000915 | -0.000106 | -0.1435 | -0.0915 | -0.0549
voltage_sweep | 7 | 1 | 14 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2843 | 0.99992 | 99.9008 | 0.19308 | -0.0086067 | 0.19327 | 193.007 | 0.999 | 0.19295 | -0.27729 | -0.000915 | -0.00013 | -0.1435 | -0.0915 | -0.0674
voltage_sweep | 7 | 1 | 15 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2843 | 0.99992 | 99.9008 | 0.19308 | -0.0086067 | 0.19327 | 193.007 | 1.001 | 0.19323 | -0.27729 | 0.001085 | 0.000152 | -0.1435 | 0.1085 | 0.0787
voltage_sweep | 7 | 1 | 16 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2843 | 0.99992 | 99.9008 | 0.19308 | -0.0086067 | 0.19327 | 193.007 | 0.999 | 0.19292 | -0.27729 | -0.000915 | -0.000159 | -0.1435 | -0.0915 | -0.0824
voltage_sweep | 7 | 1 | 17 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2843 | 0.99992 | 99.9008 | 0.19308 | -0.0086067 | 0.19327 | 193.007 | 1 | 0.19301 | -0.27729 | 8.5e-05 | -6.1e-05 | -0.1435 | 0.0085 | -0.0316
voltage_sweep | 7 | 1 | 18 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2843 | 0.99992 | 99.9008 | 0.19308 | -0.0086067 | 0.19327 | 193.007 | 1 | 0.19306 | -0.27729 | 8.5e-05 | -1.1e-05 | -0.1435 | 0.0085 | -0.0057
voltage_sweep | 7 | 2 | 1 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3162 | 0.99998 | 99.9004 | 0.19312 | -0.0086169 | 0.19331 | 192.994 | 2.97 | 0.57326 | -0.32216 | 1.97 | 0.38014 | -0.1667 | -0.99786 | -1.0525
voltage_sweep | 7 | 2 | 2 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3162 | 0.99998 | 99.9004 | 0.19312 | -0.0086169 | 0.19331 | 192.994 | 1.001 | 0.1933 | -0.32216 | 0.001022 | 0.000178 | -0.1667 | 0.1022 | 0.0924
voltage_sweep | 7 | 2 | 3 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3162 | 0.99998 | 99.9004 | 0.19312 | -0.0086169 | 0.19331 | 192.994 | 1.001 | 0.19336 | -0.32216 | 0.001022 | 0.000242 | -0.1667 | 0.1022 | 0.1255
voltage_sweep | 7 | 2 | 4 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3162 | 0.99998 | 99.9004 | 0.19312 | -0.0086169 | 0.19331 | 192.994 | 1 | 0.19318 | -0.32216 | 2.2e-05 | 5.6e-05 | -0.1667 | 0.0022 | 0.0292
voltage_sweep | 7 | 2 | 5 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3162 | 0.99998 | 99.9004 | 0.19312 | -0.0086169 | 0.19331 | 192.994 | 1.001 | 0.1933 | -0.32216 | 0.001022 | 0.000178 | -0.1667 | 0.1022 | 0.0924
voltage_sweep | 7 | 2 | 6 | 1 | 193.3 | 1 | 0.999 | 60 | 193.3162 | 0.99998 | 99.9004 | 0.19312 | -0.0086169 | 0.19331 | 192.994 | 0.999 | 0.19281 | -0.32216 | -0.000978 | -0.000314 | -0.1667 | -0.0978 | -0.1624
voltage_sweep | 7 | 2 | 7 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2775 | 1.0002 | 99.9036 | 0.19312 | -0.0084853 | 0.19331 | 192.937 | 2.979 | 0.57479 | -0.34053 | 1.9788 | 0.38167 | -0.1762 | -0.71584 | -0.78965
voltage_sweep | 7 | 2 | 8 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2775 | 1.0002 | 99.9036 | 0.19312 | -0.0084853 | 0.19331 | 192.937 | 0.999 | 0.1929 | -0.34053 | -0.00116 | -0.000225 | -0.1762 | -0.1159 | -0.1165
voltage_sweep | 7 | 2 | 9 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2775 | 1.0002 | 99.9036 | 0.19312 | -0.0084853 | 0.19331 | 192.937 | 1 | 0.19311 | -0.34053 | -0.00016 | -8e-06 | -0.1762 | -0.0159 | -0.0041
voltage_sweep | 7 | 2 | 10 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2775 | 1.0002 | 99.9036 | 0.19312 | -0.0084853 | 0.19331 | 192.937 | 1 | 0.19307 | -0.34053 | -0.00016 | -5.2e-05 | -0.1762 | -0.0159 | -0.0269
voltage_sweep | 7 | 2 | 11 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2775 | 1.0002 | 99.9036 | 0.19312 | -0.0084853 | 0.19331 | 192.937 | 1 | 0.19297 | -0.34053 | -0.00016 | -0.000152 | -0.1762 | -0.0159 | -0.0787
voltage_sweep | 7 | 2 | 12 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2775 | 1.0002 | 99.9036 | 0.19312 | -0.0084853 | 0.19331 | 192.937 | 1 | 0.19302 | -0.34053 | -0.00016 | -0.000106 | -0.1762 | -0.0159 | -0.0549
voltage_sweep | 7 | 2 | 13 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2906 | 0.99993 | 99.9009 | 0.19309 | -0.0086037 | 0.19328 | 193.003 | 1 | 0.19301 | -0.28756 | 6.9e-05 | -7.3e-05 | -0.1488 | 0.0069 | -0.0376
voltage_sweep | 7 | 2 | 14 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2906 | 0.99993 | 99.9009 | 0.19309 | -0.0086037 | 0.19328 | 193.003 | 0.999 | 0.19293 | -0.28756 | -0.000931 | -0.000154 | -0.1488 | -0.0931 | -0.0796
voltage_sweep | 7 | 2 | 15 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2906 | 0.99993 | 99.9009 | 0.19309 | -0.0086037 | 0.19328 | 193.003 | 1.001 | 0.19322 | -0.28756 | 0.001069 | 0.00013 | -0.1488 | 0.1069 | 0.0675
voltage_sweep | 7 | 2 | 16 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2906 | 0.99993 | 99.9009 | 0.19309 | -0.0086037 | 0.19328 | 193.003 | 0.999 | 0.19294 | -0.28756 | -0.000931 | -0.000149 | -0.1488 | -0.0931 | -0.077
voltage_sweep | 7 | 2 | 17 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2906 | 0.99993 | 99.9009 | 0.19309 | -0.0086037 | 0.19328 | 193.003 | 1 | 0.19303 | -0.28756 | 6.9e-05 | -5.3e-05 | -0.1488 | 0.0069 | -0.0272
voltage_sweep | 7 | 2 | 18 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2906 | 0.99993 | 99.9009 | 0.19309 | -0.0086037 | 0.19328 | 193.003 | 1 | 0.19302 | -0.28756 | 6.9e-05 | -6.3e-05 | -0.1488 | 0.0069 | -0.0324
voltage_sweep | 7 | 3 | 1 | 1 | 193.3 | 1 | 0.999 | 60 | 193.299 | 0.99998 | 99.8998 | 0.1931 | -0.0086191 | 0.1933 | 192.941 | 2.97 | 0.5733 | -0.35803 | 1.97 | 0.3802 | -0.1852 | -0.99822 | -1.0371
voltage_sweep | 7 | 3 | 2 | 1 | 193.3 | 1 | 0.999 | 60 | 193.299 | 0.99998 | 99.8998 | 0.1931 | -0.0086191 | 0.1933 | 192.941 | 1.001 | 0.19333 | -0.35803 | 0.001018 | 0.000224 | -0.1852 | 0.1018 | 0.116
voltage_sweep | 7 | 3 | 3 | 1 | 193.3 | 1 | 0.999 | 60 | 193.299 | 0.99998 | 99.8998 | 0.1931 | -0.0086191 | 0.1933 | 192.941 | 1.001 | 0.19332 | -0.35803 | 0.001018 | 0.00022 | -0.1852 | 0.1018 | 0.114
voltage_sweep | 7 | 3 | 4 | 1 | 193.3 | 1 | 0.999 | 60 | 193.299 | 0.99998 | 99.8998 | 0.1931 | -0.0086191 | 0.1933 | 192.941 | 1 | 0.19319 | -0.35803 | 1.8e-05 | 8.9e-05 | -0.1852 | 0.0018 | 0.0461
voltage_sweep | 7 | 3 | 5 | 1 | 193.3 | 1 | 0.999 | 60 | 193.299 | 0.99998 | 99.8998 | 0.1931 | -0.0086191 | 0.1933 | 192.941 | 1.001 | 0.19331 | -0.35803 | 0.001018 | 0.000206 | -0.1852 | 0.1018 | 0.1067
voltage_sweep | 7 | 3 | 6 | 1 | 193.3 | 1 | 0.999 | 60 | 193.299 | 0.99998 | 99.8998 | 0.1931 | -0.0086191 | 0.1933 | 192.941 | 0.998 | 0.19269 | -0.35803 | -0.001982 | -0.000411 | -0.1852 | -0.1982 | -0.2128
voltage_sweep | 7 | 3 | 7 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2802 | 1.0002 | 99.9038 | 0.19313 | -0.0084754 | 0.19331 | 192.943 | 2.979 | 0.57489 | -0.33717 | 1.9788 | 0.38177 | -0.1744 | -0.71591 | -0.77355
voltage_sweep | 7 | 3 | 8 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2802 | 1.0002 | 99.9038 | 0.19313 | -0.0084754 | 0.19331 | 192.943 | 0.999 | 0.19288 | -0.33717 | -0.00116 | -0.000243 | -0.1744 | -0.116 | -0.126
voltage_sweep | 7 | 3 | 9 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2802 | 1.0002 | 99.9038 | 0.19313 | -0.0084754 | 0.19331 | 192.943 | 1.001 | 0.19314 | -0.33717 | 0.00084 | 1.7e-05 | -0.1744 | 0.084 | 0.0087
voltage_sweep | 7 | 3 | 10 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2802 | 1.0002 | 99.9038 | 0.19313 | -0.0084754 | 0.19331 | 192.943 | 1 | 0.19305 | -0.33717 | -0.00016 | -7.2e-05 | -0.1744 | -0.016 | -0.0374
voltage_sweep | 7 | 3 | 11 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2802 | 1.0002 | 99.9038 | 0.19313 | -0.0084754 | 0.19331 | 192.943 | 1 | 0.19295 | -0.33717 | -0.00016 | -0.00017 | -0.1744 | -0.016 | -0.0882
voltage_sweep | 7 | 3 | 12 | 2 | 193.3 | 1 | 0.999 | 60 | 193.2802 | 1.0002 | 99.9038 | 0.19313 | -0.0084754 | 0.19331 | 192.943 | 1 | 0.19303 | -0.33717 | -0.00016 | -9.8e-05 | -0.1744 | -0.016 | -0.0509
voltage_sweep | 7 | 3 | 13 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2839 | 0.99991 | 99.9007 | 0.19308 | -0.0086079 | 0.19327 | 193.002 | 1 | 0.193 | -0.28192 | 8.7e-05 | -7.9e-05 | -0.1459 | 0.0087 | -0.0411
voltage_sweep | 7 | 3 | 14 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2839 | 0.99991 | 99.9007 | 0.19308 | -0.0086079 | 0.19327 | 193.002 | 0.999 | 0.19295 | -0.28192 | -0.000913 | -0.000126 | -0.1459 | -0.0913 | -0.0654
voltage_sweep | 7 | 3 | 15 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2839 | 0.99991 | 99.9007 | 0.19308 | -0.0086079 | 0.19327 | 193.002 | 1.001 | 0.19321 | -0.28192 | 0.001087 | 0.00014 | -0.1459 | 0.1087 | 0.0723
voltage_sweep | 7 | 3 | 16 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2839 | 0.99991 | 99.9007 | 0.19308 | -0.0086079 | 0.19327 | 193.002 | 0.999 | 0.19293 | -0.28192 | -0.000913 | -0.000148 | -0.1459 | -0.0913 | -0.0768
voltage_sweep | 7 | 3 | 17 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2839 | 0.99991 | 99.9007 | 0.19308 | -0.0086079 | 0.19327 | 193.002 | 1 | 0.19303 | -0.28192 | 8.7e-05 | -4.1e-05 | -0.1459 | 0.0087 | -0.0214
voltage_sweep | 7 | 3 | 18 | 3 | 193.3 | 1 | 0.999 | 60 | 193.2839 | 0.99991 | 99.9007 | 0.19308 | -0.0086079 | 0.19327 | 193.002 | 1 | 0.19304 | -0.28192 | 8.7e-05 | -3.7e-05 | -0.1459 | 0.0087 | -0.0193
voltage_sweep | 8 | 1 | 1 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9434 | 0.99999 | 99.9006 | 0.20873 | -0.0093132 | 0.20894 | 208.535 | 2.97 | 0.61944 | -0.40844 | 1.97 | 0.41071 | -0.1955 | -0.99875 | -1.0788
voltage_sweep | 8 | 1 | 2 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9434 | 0.99999 | 99.9006 | 0.20873 | -0.0093132 | 0.20894 | 208.535 | 1.001 | 0.20893 | -0.40844 | 0.001013 | 0.000193 | -0.1955 | 0.1013 | 0.0924
voltage_sweep | 8 | 1 | 3 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9434 | 0.99999 | 99.9006 | 0.20873 | -0.0093132 | 0.20894 | 208.535 | 1.002 | 0.20904 | -0.40844 | 0.002013 | 0.000307 | -0.1955 | 0.2013 | 0.147
voltage_sweep | 8 | 1 | 4 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9434 | 0.99999 | 99.9006 | 0.20873 | -0.0093132 | 0.20894 | 208.535 | 1 | 0.2087 | -0.40844 | 1.3e-05 | -3e-05 | -0.1955 | 0.0013 | -0.0144
voltage_sweep | 8 | 1 | 5 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9434 | 0.99999 | 99.9006 | 0.20873 | -0.0093132 | 0.20894 | 208.535 | 1.001 | 0.20886 | -0.40844 | 0.001013 | 0.000126 | -0.1955 | 0.1013 | 0.0603
voltage_sweep | 8 | 1 | 6 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9434 | 0.99999 | 99.9006 | 0.20873 | -0.0093132 | 0.20894 | 208.535 | 0.998 | 0.20838 | -0.40844 | -0.001987 | -0.000355 | -0.1955 | -0.1987 | -0.1701
voltage_sweep | 8 | 1 | 7 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8646 | 1.0002 | 99.9038 | 0.2087 | -0.0091571 | 0.2089 | 208.5 | 2.979 | 0.62115 | -0.36459 | 1.9788 | 0.41246 | -0.1746 | -0.71585 | -0.78873
voltage_sweep | 8 | 1 | 8 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8646 | 1.0002 | 99.9038 | 0.2087 | -0.0091571 | 0.2089 | 208.5 | 0.999 | 0.20843 | -0.36459 | -0.00116 | -0.000263 | -0.1746 | -0.1159 | -0.1261
voltage_sweep | 8 | 1 | 9 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8646 | 1.0002 | 99.9038 | 0.2087 | -0.0091571 | 0.2089 | 208.5 | 1 | 0.20871 | -0.36459 | -0.00016 | 1.7e-05 | -0.1746 | -0.016 | 0.0081
voltage_sweep | 8 | 1 | 10 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8646 | 1.0002 | 99.9038 | 0.2087 | -0.0091571 | 0.2089 | 208.5 | 1 | 0.20864 | -0.36459 | -0.00016 | -5.6e-05 | -0.1746 | -0.016 | -0.0269
voltage_sweep | 8 | 1 | 11 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8646 | 1.0002 | 99.9038 | 0.2087 | -0.0091571 | 0.2089 | 208.5 | 1 | 0.20852 | -0.36459 | -0.00016 | -0.000178 | -0.1746 | -0.016 | -0.0853
voltage_sweep | 8 | 1 | 12 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8646 | 1.0002 | 99.9038 | 0.2087 | -0.0091571 | 0.2089 | 208.5 | 1 | 0.20862 | -0.36459 | -0.00016 | -8e-05 | -0.1746 | -0.016 | -0.0384
voltage_sweep | 8 | 1 | 13 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8833 | 0.99992 | 99.9008 | 0.20866 | -0.0093014 | 0.20887 | 208.586 | 1 | 0.20861 | -0.2973 | 7.7e-05 | -4.8e-05 | -0.1423 | 0.0077 | -0.023
voltage_sweep | 8 | 1 | 14 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8833 | 0.99992 | 99.9008 | 0.20866 | -0.0093014 | 0.20887 | 208.586 | 0.999 | 0.20854 | -0.2973 | -0.000923 | -0.000123 | -0.1423 | -0.0923 | -0.0589
voltage_sweep | 8 | 1 | 15 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8833 | 0.99992 | 99.9008 | 0.20866 | -0.0093014 | 0.20887 | 208.586 | 1 | 0.20875 | -0.2973 | 7.7e-05 | 9.2e-05 | -0.1423 | 0.0077 | 0.0441
voltage_sweep | 8 | 1 | 16 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8833 | 0.99992 | 99.9008 | 0.20866 | -0.0093014 | 0.20887 | 208.586 | 0.999 | 0.20853 | -0.2973 | -0.000923 | -0.000132 | -0.1423 | -0.0923 | -0.0632
voltage_sweep | 8 | 1 | 17 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8833 | 0.99992 | 99.9008 | 0.20866 | -0.0093014 | 0.20887 | 208.586 | 1 | 0.20863 | -0.2973 | 7.7e-05 | -3.1e-05 | -0.1423 | 0.0077 | -0.0148
voltage_sweep | 8 | 1 | 18 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8833 | 0.99992 | 99.9008 | 0.20866 | -0.0093014 | 0.20887 | 208.586 | 1 | 0.20859 | -0.2973 | 7.7e-05 | -7.3e-05 | -0.1423 | 0.0077 | -0.035
voltage_sweep | 8 | 2 | 1 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9166 | 0.99995 | 99.9004 | 0.2087 | -0.0093161 | 0.20891 | 208.565 | 2.97 | 0.61959 | -0.35158 | 1.97 | 0.41089 | -0.1683 | -0.99552 | -1.0401
voltage_sweep | 8 | 2 | 2 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9166 | 0.99995 | 99.9004 | 0.2087 | -0.0093161 | 0.20891 | 208.565 | 1.001 | 0.20894 | -0.35158 | 0.001045 | 0.000245 | -0.1683 | 0.1045 | 0.1174
voltage_sweep | 8 | 2 | 3 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9166 | 0.99995 | 99.9004 | 0.2087 | -0.0093161 | 0.20891 | 208.565 | 1.001 | 0.20901 | -0.35158 | 0.001045 | 0.000314 | -0.1683 | 0.1045 | 0.1505
voltage_sweep | 8 | 2 | 4 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9166 | 0.99995 | 99.9004 | 0.2087 | -0.0093161 | 0.20891 | 208.565 | 1 | 0.20878 | -0.35158 | 4.5e-05 | 8.3e-05 | -0.1683 | 0.0045 | 0.0398
voltage_sweep | 8 | 2 | 5 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9166 | 0.99995 | 99.9004 | 0.2087 | -0.0093161 | 0.20891 | 208.565 | 1.001 | 0.20889 | -0.35158 | 0.001045 | 0.000196 | -0.1683 | 0.1045 | 0.0939
voltage_sweep | 8 | 2 | 6 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9166 | 0.99995 | 99.9004 | 0.2087 | -0.0093161 | 0.20891 | 208.565 | 0.998 | 0.20836 | -0.35158 | -0.001955 | -0.00034 | -0.1683 | -0.1955 | -0.1629
voltage_sweep | 8 | 2 | 7 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8679 | 1.0001 | 99.9038 | 0.2087 | -0.0091575 | 0.2089 | 208.499 | 2.979 | 0.62117 | -0.36892 | 1.9789 | 0.41247 | -0.1766 | -0.71382 | -0.78587
voltage_sweep | 8 | 2 | 8 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8679 | 1.0001 | 99.9038 | 0.2087 | -0.0091575 | 0.2089 | 208.499 | 0.999 | 0.20843 | -0.36892 | -0.001139 | -0.000269 | -0.1766 | -0.1139 | -0.1289
voltage_sweep | 8 | 2 | 9 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8679 | 1.0001 | 99.9038 | 0.2087 | -0.0091575 | 0.2089 | 208.499 | 1.001 | 0.20876 | -0.36892 | 0.000861 | 6.8e-05 | -0.1766 | 0.0861 | 0.0325
voltage_sweep | 8 | 2 | 10 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8679 | 1.0001 | 99.9038 | 0.2087 | -0.0091575 | 0.2089 | 208.499 | 1 | 0.20864 | -0.36892 | -0.000139 | -5.8e-05 | -0.1766 | -0.0139 | -0.0278
voltage_sweep | 8 | 2 | 11 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8679 | 1.0001 | 99.9038 | 0.2087 | -0.0091575 | 0.2089 | 208.499 | 1 | 0.20851 | -0.36892 | -0.000139 | -0.000181 | -0.1766 | -0.0139 | -0.0868
voltage_sweep | 8 | 2 | 12 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8679 | 1.0001 | 99.9038 | 0.2087 | -0.0091575 | 0.2089 | 208.499 | 1 | 0.20862 | -0.36892 | -0.000139 | -7.6e-05 | -0.1766 | -0.0139 | -0.0365
voltage_sweep | 8 | 2 | 13 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8806 | 0.99993 | 99.9008 | 0.20866 | -0.0092989 | 0.20887 | 208.576 | 1 | 0.20856 | -0.30455 | 6.9e-05 | -9.6e-05 | -0.1458 | 0.0069 | -0.046
voltage_sweep | 8 | 2 | 14 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8806 | 0.99993 | 99.9008 | 0.20866 | -0.0092989 | 0.20887 | 208.576 | 0.999 | 0.20854 | -0.30455 | -0.000931 | -0.000117 | -0.1458 | -0.0931 | -0.0561
voltage_sweep | 8 | 2 | 15 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8806 | 0.99993 | 99.9008 | 0.20866 | -0.0092989 | 0.20887 | 208.576 | 1 | 0.20876 | -0.30455 | 6.9e-05 | 0.000103 | -0.1458 | 0.0069 | 0.0494
voltage_sweep | 8 | 2 | 16 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8806 | 0.99993 | 99.9008 | 0.20866 | -0.0092989 | 0.20887 | 208.576 | 0.999 | 0.20849 | -0.30455 | -0.000931 | -0.000172 | -0.1458 | -0.0931 | -0.0824
voltage_sweep | 8 | 2 | 17 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8806 | 0.99993 | 99.9008 | 0.20866 | -0.0092989 | 0.20887 | 208.576 | 1 | 0.20862 | -0.30455 | 6.9e-05 | -3.7e-05 | -0.1458 | 0.0069 | -0.0177
voltage_sweep | 8 | 2 | 18 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8806 | 0.99993 | 99.9008 | 0.20866 | -0.0092989 | 0.20887 | 208.576 | 1 | 0.20856 | -0.30455 | 6.9e-05 | -0.000103 | -0.1458 | 0.0069 | -0.0494
voltage_sweep | 8 | 3 | 1 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9256 | 0.99996 | 99.9004 | 0.20871 | -0.0093179 | 0.20892 | 208.467 | 2.97 | 0.61937 | -0.4586 | 1.97 | 0.41066 | -0.2195 | -0.99622 | -1.0799
voltage_sweep | 8 | 3 | 2 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9256 | 0.99996 | 99.9004 | 0.20871 | -0.0093179 | 0.20892 | 208.467 | 1.001 | 0.20882 | -0.4586 | 0.001038 | 0.000112 | -0.2195 | 0.1038 | 0.0539
voltage_sweep | 8 | 3 | 3 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9256 | 0.99996 | 99.9004 | 0.20871 | -0.0093179 | 0.20892 | 208.467 | 1.002 | 0.209 | -0.4586 | 0.002038 | 0.000292 | -0.2195 | 0.2038 | 0.1401
voltage_sweep | 8 | 3 | 4 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9256 | 0.99996 | 99.9004 | 0.20871 | -0.0093179 | 0.20892 | 208.467 | 1 | 0.20872 | -0.4586 | 3.8e-05 | 1.4e-05 | -0.2195 | 0.0038 | 0.0069
voltage_sweep | 8 | 3 | 5 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9256 | 0.99996 | 99.9004 | 0.20871 | -0.0093179 | 0.20892 | 208.467 | 1.001 | 0.20876 | -0.4586 | 0.001038 | 4.9e-05 | -0.2195 | 0.1038 | 0.0237
voltage_sweep | 8 | 3 | 6 | 1 | 208.9 | 1 | 0.999 | 60 | 208.9256 | 0.99996 | 99.9004 | 0.20871 | -0.0093179 | 0.20892 | 208.467 | 0.998 | 0.20836 | -0.4586 | -0.001962 | -0.000349 | -0.2195 | -0.1962 | -0.167
voltage_sweep | 8 | 3 | 7 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8711 | 1.0002 | 99.904 | 0.20871 | -0.0091538 | 0.20891 | 208.502 | 2.979 | 0.62121 | -0.36914 | 1.9788 | 0.4125 | -0.1767 | -0.71715 | -0.78398
voltage_sweep | 8 | 3 | 8 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8711 | 1.0002 | 99.904 | 0.20871 | -0.0091538 | 0.20891 | 208.502 | 0.999 | 0.20844 | -0.36914 | -0.001173 | -0.000271 | -0.1767 | -0.1173 | -0.1296
voltage_sweep | 8 | 3 | 9 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8711 | 1.0002 | 99.904 | 0.20871 | -0.0091538 | 0.20891 | 208.502 | 1.001 | 0.20872 | -0.36914 | 0.000827 | 8e-06 | -0.1767 | 0.0827 | 0.004
voltage_sweep | 8 | 3 | 10 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8711 | 1.0002 | 99.904 | 0.20871 | -0.0091538 | 0.20891 | 208.502 | 1 | 0.20863 | -0.36914 | -0.000173 | -7.5e-05 | -0.1767 | -0.0173 | -0.0357
voltage_sweep | 8 | 3 | 11 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8711 | 1.0002 | 99.904 | 0.20871 | -0.0091538 | 0.20891 | 208.502 | 1 | 0.20854 | -0.36914 | -0.000173 | -0.000165 | -0.1767 | -0.0173 | -0.0788
voltage_sweep | 8 | 3 | 12 | 2 | 208.9 | 1 | 0.999 | 60 | 208.8711 | 1.0002 | 99.904 | 0.20871 | -0.0091538 | 0.20891 | 208.502 | 1 | 0.20859 | -0.36914 | -0.000173 | -0.000113 | -0.1767 | -0.0173 | -0.0539
voltage_sweep | 8 | 3 | 13 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8825 | 0.99991 | 99.9008 | 0.20866 | -0.0092989 | 0.20886 | 208.579 | 1 | 0.20863 | -0.30346 | 8.8e-05 | -3.1e-05 | -0.1453 | 0.0088 | -0.0148
voltage_sweep | 8 | 3 | 14 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8825 | 0.99991 | 99.9008 | 0.20866 | -0.0092989 | 0.20886 | 208.579 | 0.999 | 0.20855 | -0.30346 | -0.000912 | -0.000109 | -0.1453 | -0.0912 | -0.0522
voltage_sweep | 8 | 3 | 15 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8825 | 0.99991 | 99.9008 | 0.20866 | -0.0092989 | 0.20886 | 208.579 | 1 | 0.20875 | -0.30346 | 8.8e-05 | 8.8e-05 | -0.1453 | 0.0088 | 0.0422
voltage_sweep | 8 | 3 | 16 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8825 | 0.99991 | 99.9008 | 0.20866 | -0.0092989 | 0.20886 | 208.579 | 0.999 | 0.20855 | -0.30346 | -0.000912 | -0.000106 | -0.1453 | -0.0912 | -0.0507
voltage_sweep | 8 | 3 | 17 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8825 | 0.99991 | 99.9008 | 0.20866 | -0.0092989 | 0.20886 | 208.579 | 1 | 0.20863 | -0.30346 | 8.8e-05 | -2.3e-05 | -0.1453 | 0.0088 | -0.0109
voltage_sweep | 8 | 3 | 18 | 3 | 208.9 | 1 | 0.999 | 60 | 208.8825 | 0.99991 | 99.9008 | 0.20866 | -0.0092989 | 0.20886 | 208.579 | 1 | 0.20857 | -0.30346 | 8.8e-05 | -9.1e-05 | -0.1453 | 0.0088 | -0.0435
voltage_sweep | 9 | 1 | 1 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4098 | 1 | 99.9005 | 0.22419 | -0.010006 | 0.22441 | 224.049 | 2.97 | 0.66544 | -0.3608 | 1.97 | 0.44126 | -0.1608 | -1.0006 | -1.0586
voltage_sweep | 9 | 1 | 2 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4098 | 1 | 99.9005 | 0.22419 | -0.010006 | 0.22441 | 224.049 | 1.001 | 0.22445 | -0.3608 | 0.000994 | 0.000257 | -0.1608 | 0.0994 | 0.1147
voltage_sweep | 9 | 1 | 3 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4098 | 1 | 99.9005 | 0.22419 | -0.010006 | 0.22441 | 224.049 | 1.002 | 0.22445 | -0.3608 | 0.001994 | 0.000257 | -0.1608 | 0.1994 | 0.1147
voltage_sweep | 9 | 1 | 4 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4098 | 1 | 99.9005 | 0.22419 | -0.010006 | 0.22441 | 224.049 | 1.001 | 0.22428 | -0.3608 | 0.000994 | 8.8e-05 | -0.1608 | 0.0994 | 0.0393
voltage_sweep | 9 | 1 | 5 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4098 | 1 | 99.9005 | 0.22419 | -0.010006 | 0.22441 | 224.049 | 1.001 | 0.2244 | -0.3608 | 0.000994 | 0.000216 | -0.1608 | 0.0994 | 0.0964
voltage_sweep | 9 | 1 | 6 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4098 | 1 | 99.9005 | 0.22419 | -0.010006 | 0.22441 | 224.049 | 0.998 | 0.22371 | -0.3608 | -0.002006 | -0.000482 | -0.1608 | -0.2006 | -0.2149
voltage_sweep | 9 | 1 | 7 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3638 | 1.0001 | 99.9038 | 0.22418 | -0.0098394 | 0.2244 | 223.971 | 2.979 | 0.66727 | -0.39285 | 1.9789 | 0.44309 | -0.1751 | -0.71462 | -0.78448
voltage_sweep | 9 | 1 | 8 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3638 | 1.0001 | 99.9038 | 0.22418 | -0.0098394 | 0.2244 | 223.971 | 0.999 | 0.22386 | -0.39285 | -0.001147 | -0.000321 | -0.1751 | -0.1147 | -0.1432
voltage_sweep | 9 | 1 | 9 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3638 | 1.0001 | 99.9038 | 0.22418 | -0.0098394 | 0.2244 | 223.971 | 1 | 0.2242 | -0.39285 | -0.000147 | 2.3e-05 | -0.1751 | -0.0147 | 0.0103
voltage_sweep | 9 | 1 | 10 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3638 | 1.0001 | 99.9038 | 0.22418 | -0.0098394 | 0.2244 | 223.971 | 1 | 0.22412 | -0.39285 | -0.000147 | -6.3e-05 | -0.1751 | -0.0147 | -0.0281
voltage_sweep | 9 | 1 | 11 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3638 | 1.0001 | 99.9038 | 0.22418 | -0.0098394 | 0.2244 | 223.971 | 1 | 0.22398 | -0.39285 | -0.000147 | -0.000198 | -0.1751 | -0.0147 | -0.0883
voltage_sweep | 9 | 1 | 12 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3638 | 1.0001 | 99.9038 | 0.22418 | -0.0098394 | 0.2244 | 223.971 | 1 | 0.22408 | -0.39285 | -0.000147 | -0.0001 | -0.1751 | -0.0147 | -0.0446
voltage_sweep | 9 | 1 | 13 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3886 | 0.99993 | 99.9008 | 0.22415 | -0.0099878 | 0.22437 | 224.064 | 1 | 0.22408 | -0.32456 | 6.8e-05 | -6.7e-05 | -0.1446 | 0.0068 | -0.0298
voltage_sweep | 9 | 1 | 14 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3886 | 0.99993 | 99.9008 | 0.22415 | -0.0099878 | 0.22437 | 224.064 | 0.999 | 0.22399 | -0.32456 | -0.000932 | -0.000157 | -0.1446 | -0.0932 | -0.0699
voltage_sweep | 9 | 1 | 15 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3886 | 0.99993 | 99.9008 | 0.22415 | -0.0099878 | 0.22437 | 224.064 | 1.001 | 0.2243 | -0.32456 | 0.001068 | 0.000148 | -0.1446 | 0.1068 | 0.0661
voltage_sweep | 9 | 1 | 16 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3886 | 0.99993 | 99.9008 | 0.22415 | -0.0099878 | 0.22437 | 224.064 | 0.999 | 0.224 | -0.32456 | -0.000932 | -0.000148 | -0.1446 | -0.0932 | -0.0659
voltage_sweep | 9 | 1 | 17 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3886 | 0.99993 | 99.9008 | 0.22415 | -0.0099878 | 0.22437 | 224.064 | 1 | 0.22409 | -0.32456 | 6.8e-05 | -6.1e-05 | -0.1446 | 0.0068 | -0.0271
voltage_sweep | 9 | 1 | 18 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3886 | 0.99993 | 99.9008 | 0.22415 | -0.0099878 | 0.22437 | 224.064 | 1 | 0.22409 | -0.32456 | 6.8e-05 | -5.8e-05 | -0.1446 | 0.0068 | -0.0258
voltage_sweep | 9 | 2 | 1 | 1 | 224.4 | 1 | 0.999 | 60 | 224.396 | 1 | 99.9004 | 0.22417 | -0.010007 | 0.2244 | 223.814 | 2.969 | 0.66465 | -0.582 | 1.969 | 0.44048 | -0.2594 | -1.0344 | -1.1709
voltage_sweep | 9 | 2 | 2 | 1 | 224.4 | 1 | 0.999 | 60 | 224.396 | 1 | 99.9004 | 0.22417 | -0.010007 | 0.2244 | 223.814 | 1.001 | 0.22445 | -0.582 | 0.00099 | 0.000273 | -0.2594 | 0.099 | 0.1218
voltage_sweep | 9 | 2 | 3 | 1 | 224.4 | 1 | 0.999 | 60 | 224.396 | 1 | 99.9004 | 0.22417 | -0.010007 | 0.2244 | 223.814 | 1.001 | 0.22444 | -0.582 | 0.00099 | 0.000267 | -0.2594 | 0.099 | 0.1192
voltage_sweep | 9 | 2 | 4 | 1 | 224.4 | 1 | 0.999 | 60 | 224.396 | 1 | 99.9004 | 0.22417 | -0.010007 | 0.2244 | 223.814 | 1 | 0.22399 | -0.582 | -1e-05 | -0.000181 | -0.2594 | -0.001 | -0.0807
voltage_sweep | 9 | 2 | 5 | 1 | 224.4 | 1 | 0.999 | 60 | 224.396 | 1 | 99.9004 | 0.22417 | -0.010007 | 0.2244 | 223.814 | 1.001 | 0.22438 | -0.582 | 0.00099 | 0.000201 | -0.2594 | 0.099 | 0.0897
voltage_sweep | 9 | 2 | 6 | 1 | 224.4 | 1 | 0.999 | 60 | 224.396 | 1 | 99.9004 | 0.22417 | -0.010007 | 0.2244 | 223.814 | 0.998 | 0.22372 | -0.582 | -0.00201 | -0.000457 | -0.2594 | -0.201 | -0.2038
voltage_sweep | 9 | 2 | 7 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3693 | 1.0002 | 99.9039 | 0.22419 | -0.0098347 | 0.2244 | 223.973 | 2.979 | 0.66738 | -0.39628 | 1.9788 | 0.44319 | -0.1766 | -0.71513 | -0.77116
voltage_sweep | 9 | 2 | 8 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3693 | 1.0002 | 99.9039 | 0.22419 | -0.0098347 | 0.2244 | 223.973 | 0.999 | 0.22388 | -0.39628 | -0.001152 | -0.000312 | -0.1766 | -0.1152 | -0.1391
voltage_sweep | 9 | 2 | 9 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3693 | 1.0002 | 99.9039 | 0.22419 | -0.0098347 | 0.2244 | 223.973 | 1 | 0.22418 | -0.39628 | -0.000152 | -7e-06 | -0.1766 | -0.0152 | -0.0031
voltage_sweep | 9 | 2 | 10 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3693 | 1.0002 | 99.9039 | 0.22419 | -0.0098347 | 0.2244 | 223.973 | 1 | 0.22413 | -0.39628 | -0.000152 | -5.8e-05 | -0.1766 | -0.0152 | -0.0258
voltage_sweep | 9 | 2 | 11 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3693 | 1.0002 | 99.9039 | 0.22419 | -0.0098347 | 0.2244 | 223.973 | 1 | 0.224 | -0.39628 | -0.000152 | -0.000186 | -0.1766 | -0.0152 | -0.0829
voltage_sweep | 9 | 2 | 12 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3693 | 1.0002 | 99.9039 | 0.22419 | -0.0098347 | 0.2244 | 223.973 | 1 | 0.22404 | -0.39628 | -0.000152 | -0.000143 | -0.1766 | -0.0152 | -0.0637
voltage_sweep | 9 | 2 | 13 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3832 | 0.99992 | 99.9008 | 0.22414 | -0.0099932 | 0.22436 | 224.057 | 1 | 0.22408 | -0.32622 | 8.2e-05 | -5.9e-05 | -0.1454 | 0.0082 | -0.0263
voltage_sweep | 9 | 2 | 14 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3832 | 0.99992 | 99.9008 | 0.22414 | -0.0099932 | 0.22436 | 224.057 | 0.999 | 0.22399 | -0.32622 | -0.000918 | -0.000155 | -0.1454 | -0.0918 | -0.0692
voltage_sweep | 9 | 2 | 15 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3832 | 0.99992 | 99.9008 | 0.22414 | -0.0099932 | 0.22436 | 224.057 | 1.001 | 0.22429 | -0.32622 | 0.001082 | 0.000148 | -0.1454 | 0.1083 | 0.066
voltage_sweep | 9 | 2 | 16 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3832 | 0.99992 | 99.9008 | 0.22414 | -0.0099932 | 0.22436 | 224.057 | 0.999 | 0.224 | -0.32622 | -0.000918 | -0.00014 | -0.1454 | -0.0918 | -0.0625
voltage_sweep | 9 | 2 | 17 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3832 | 0.99992 | 99.9008 | 0.22414 | -0.0099932 | 0.22436 | 224.057 | 1 | 0.22409 | -0.32622 | 8.2e-05 | -5.7e-05 | -0.1454 | 0.0082 | -0.0254
voltage_sweep | 9 | 2 | 18 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3832 | 0.99992 | 99.9008 | 0.22414 | -0.0099932 | 0.22436 | 224.057 | 1 | 0.22408 | -0.32622 | 8.2e-05 | -6.4e-05 | -0.1454 | 0.0082 | -0.0286
voltage_sweep | 9 | 3 | 1 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4295 | 1 | 99.9006 | 0.22421 | -0.0099991 | 0.22443 | 224.047 | 2.969 | 0.66529 | -0.38247 | 1.969 | 0.44109 | -0.1704 | -1.0338 | -1.0897
voltage_sweep | 9 | 3 | 2 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4295 | 1 | 99.9006 | 0.22421 | -0.0099991 | 0.22443 | 224.047 | 1.001 | 0.22445 | -0.38247 | 0.000995 | 0.000238 | -0.1704 | 0.0995 | 0.1059
voltage_sweep | 9 | 3 | 3 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4295 | 1 | 99.9006 | 0.22421 | -0.0099991 | 0.22443 | 224.047 | 1.001 | 0.22438 | -0.38247 | 0.000995 | 0.000172 | -0.1704 | 0.0995 | 0.0765
voltage_sweep | 9 | 3 | 4 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4295 | 1 | 99.9006 | 0.22421 | -0.0099991 | 0.22443 | 224.047 | 1 | 0.22421 | -0.38247 | -5e-06 | 6e-06 | -0.1704 | -0.0005 | 0.0025
voltage_sweep | 9 | 3 | 5 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4295 | 1 | 99.9006 | 0.22421 | -0.0099991 | 0.22443 | 224.047 | 1.001 | 0.22442 | -0.38247 | 0.000995 | 0.000217 | -0.1704 | 0.0995 | 0.0966
voltage_sweep | 9 | 3 | 6 | 1 | 224.4 | 1 | 0.999 | 60 | 224.4295 | 1 | 99.9006 | 0.22421 | -0.0099991 | 0.22443 | 224.047 | 0.998 | 0.22369 | -0.38247 | -0.002005 | -0.000517 | -0.1704 | -0.2005 | -0.2308
voltage_sweep | 9 | 3 | 7 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3831 | 1.0002 | 99.9036 | 0.2242 | -0.0098498 | 0.22442 | 223.999 | 2.979 | 0.66744 | -0.38413 | 1.9788 | 0.44324 | -0.1712 | -0.71645 | -0.76883
voltage_sweep | 9 | 3 | 8 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3831 | 1.0002 | 99.9036 | 0.2242 | -0.0098498 | 0.22442 | 223.999 | 0.999 | 0.22389 | -0.38413 | -0.001166 | -0.000311 | -0.1712 | -0.1166 | -0.1387
voltage_sweep | 9 | 3 | 9 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3831 | 1.0002 | 99.9036 | 0.2242 | -0.0098498 | 0.22442 | 223.999 | 1.001 | 0.22422 | -0.38413 | 0.000834 | 1.3e-05 | -0.1712 | 0.0834 | 0.0058
voltage_sweep | 9 | 3 | 10 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3831 | 1.0002 | 99.9036 | 0.2242 | -0.0098498 | 0.22442 | 223.999 | 1 | 0.22416 | -0.38413 | -0.000166 | -4.1e-05 | -0.1712 | -0.0166 | -0.0183
voltage_sweep | 9 | 3 | 11 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3831 | 1.0002 | 99.9036 | 0.2242 | -0.0098498 | 0.22442 | 223.999 | 1 | 0.22405 | -0.38413 | -0.000166 | -0.000153 | -0.1712 | -0.0166 | -0.0683
voltage_sweep | 9 | 3 | 12 | 2 | 224.4 | 1 | 0.999 | 60 | 224.3831 | 1.0002 | 99.9036 | 0.2242 | -0.0098498 | 0.22442 | 223.999 | 1 | 0.22409 | -0.38413 | -0.000166 | -0.000117 | -0.1712 | -0.0166 | -0.0522
voltage_sweep | 9 | 3 | 13 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3751 | 0.9999 | 99.9009 | 0.22413 | -0.0099867 | 0.22435 | 224.018 | 1 | 0.22408 | -0.35706 | 9.9e-05 | -5.4e-05 | -0.1591 | 0.0099 | -0.0243
voltage_sweep | 9 | 3 | 14 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3751 | 0.9999 | 99.9009 | 0.22413 | -0.0099867 | 0.22435 | 224.018 | 0.999 | 0.22394 | -0.35706 | -0.000901 | -0.00019 | -0.1591 | -0.0901 | -0.0849
voltage_sweep | 9 | 3 | 15 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3751 | 0.9999 | 99.9009 | 0.22413 | -0.0099867 | 0.22435 | 224.018 | 1.001 | 0.2243 | -0.35706 | 0.001099 | 0.00017 | -0.1591 | 0.1099 | 0.0757
voltage_sweep | 9 | 3 | 16 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3751 | 0.9999 | 99.9009 | 0.22413 | -0.0099867 | 0.22435 | 224.018 | 0.999 | 0.22397 | -0.35706 | -0.000901 | -0.000156 | -0.1591 | -0.0901 | -0.0698
voltage_sweep | 9 | 3 | 17 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3751 | 0.9999 | 99.9009 | 0.22413 | -0.0099867 | 0.22435 | 224.018 | 1 | 0.22404 | -0.35706 | 9.9e-05 | -9.4e-05 | -0.1591 | 0.0099 | -0.0421
voltage_sweep | 9 | 3 | 18 | 3 | 224.4 | 1 | 0.999 | 60 | 224.3751 | 0.9999 | 99.9009 | 0.22413 | -0.0099867 | 0.22435 | 224.018 | 1 | 0.2241 | -0.35706 | 9.9e-05 | -3.2e-05 | -0.1591 | 0.0099 | -0.0144
voltage_sweep | 10 | 1 | 1 | 1 | 240 | 1 | 0.999 | 60 | 240.0302 | 0.99997 | 99.9005 | 0.23979 | -0.010704 | 0.24002 | 239.554 | 2.969 | 0.71164 | -0.47623 | 1.969 | 0.47186 | -0.1984 | -1.0308 | -1.0719
voltage_sweep | 10 | 1 | 2 | 1 | 240 | 1 | 0.999 | 60 | 240.0302 | 0.99997 | 99.9005 | 0.23979 | -0.010704 | 0.24002 | 239.554 | 1.001 | 0.24004 | -0.47623 | 0.001025 | 0.000253 | -0.1984 | 0.1025 | 0.1054
voltage_sweep | 10 | 1 | 3 | 1 | 240 | 1 | 0.999 | 60 | 240.0302 | 0.99997 | 99.9005 | 0.23979 | -0.010704 | 0.24002 | 239.554 | 1.001 | 0.24002 | -0.47623 | 0.001025 | 0.000233 | -0.1984 | 0.1025 | 0.097
voltage_sweep | 10 | 1 | 4 | 1 | 240 | 1 | 0.999 | 60 | 240.0302 | 0.99997 | 99.9005 | 0.23979 | -0.010704 | 0.24002 | 239.554 | 1.001 | 0.23986 | -0.47623 | 0.001025 | 7.4e-05 | -0.1984 | 0.1025 | 0.0307
voltage_sweep | 10 | 1 | 5 | 1 | 240 | 1 | 0.999 | 60 | 240.0302 | 0.99997 | 99.9005 | 0.23979 | -0.010704 | 0.24002 | 239.554 | 1.001 | 0.23995 | -0.47623 | 0.001025 | 0.000169 | -0.1984 | 0.1025 | 0.0703
voltage_sweep | 10 | 1 | 6 | 1 | 240 | 1 | 0.999 | 60 | 240.0302 | 0.99997 | 99.9005 | 0.23979 | -0.010704 | 0.24002 | 239.554 | 0.998 | 0.23919 | -0.47623 | -0.001975 | -0.000591 | -0.1984 | -0.1975 | -0.2466
voltage_sweep | 10 | 1 | 7 | 2 | 240 | 1 | 0.999 | 60 | 239.9731 | 1.0002 | 99.9037 | 0.23978 | -0.010532 | 0.24001 | 239.568 | 2.979 | 0.71362 | -0.40507 | 1.9788 | 0.47384 | -0.1688 | -0.71611 | -0.79579
voltage_sweep | 10 | 1 | 8 | 2 | 240 | 1 | 0.999 | 60 | 239.9731 | 1.0002 | 99.9037 | 0.23978 | -0.010532 | 0.24001 | 239.568 | 0.999 | 0.23938 | -0.40507 | -0.001162 | -0.000402 | -0.1688 | -0.1162 | -0.1676
voltage_sweep | 10 | 1 | 9 | 2 | 240 | 1 | 0.999 | 60 | 239.9731 | 1.0002 | 99.9037 | 0.23978 | -0.010532 | 0.24001 | 239.568 | 1.001 | 0.23982 | -0.40507 | 0.000838 | 3.4e-05 | -0.1688 | 0.0838 | 0.0143
voltage_sweep | 10 | 1 | 10 | 2 | 240 | 1 | 0.999 | 60 | 239.9731 | 1.0002 | 99.9037 | 0.23978 | -0.010532 | 0.24001 | 239.568 | 1 | 0.23966 | -0.40507 | -0.000162 | -0.000124 | -0.1688 | -0.0162 | -0.0516
voltage_sweep | 10 | 1 | 11 | 2 | 240 | 1 | 0.999 | 60 | 239.9731 | 1.0002 | 99.9037 | 0.23978 | -0.010532 | 0.24001 | 239.568 | 1 | 0.23956 | -0.40507 | -0.000162 | -0.000223 | -0.1688 | -0.0162 | -0.0929
voltage_sweep | 10 | 1 | 12 | 2 | 240 | 1 | 0.999 | 60 | 239.9731 | 1.0002 | 99.9037 | 0.23978 | -0.010532 | 0.24001 | 239.568 | 1 | 0.23971 | -0.40507 | -0.000162 | -7.5e-05 | -0.1688 | -0.0162 | -0.0312
voltage_sweep | 10 | 1 | 13 | 3 | 240 | 1 | 0.999 | 60 | 239.9827 | 0.99992 | 99.9008 | 0.23972 | -0.010684 | 0.23996 | 239.623 | 1 | 0.23967 | -0.35965 | 8.4e-05 | -5.1e-05 | -0.1499 | 0.0084 | -0.0215
voltage_sweep | 10 | 1 | 14 | 3 | 240 | 1 | 0.999 | 60 | 239.9827 | 0.99992 | 99.9008 | 0.23972 | -0.010684 | 0.23996 | 239.623 | 0.999 | 0.2396 | -0.35965 | -0.000916 | -0.000129 | -0.1499 | -0.0916 | -0.054
voltage_sweep | 10 | 1 | 15 | 3 | 240 | 1 | 0.999 | 60 | 239.9827 | 0.99992 | 99.9008 | 0.23972 | -0.010684 | 0.23996 | 239.623 | 1 | 0.23985 | -0.35965 | 8.4e-05 | 0.000123 | -0.1499 | 0.0084 | 0.0511
voltage_sweep | 10 | 1 | 16 | 3 | 240 | 1 | 0.999 | 60 | 239.9827 | 0.99992 | 99.9008 | 0.23972 | -0.010684 | 0.23996 | 239.623 | 0.999 | 0.23956 | -0.35965 | -0.000916 | -0.000165 | -0.1499 | -0.0916 | -0.069
voltage_sweep | 10 | 1 | 17 | 3 | 240 | 1 | 0.999 | 60 | 239.9827 | 0.99992 | 99.9008 | 0.23972 | -0.010684 | 0.23996 | 239.623 | 1 | 0.23971 | -0.35965 | 8.4e-05 | -1.3e-05 | -0.1499 | 0.0084 | -0.0056
voltage_sweep | 10 | 1 | 18 | 3 | 240 | 1 | 0.999 | 60 | 239.9827 | 0.99992 | 99.9008 | 0.23972 | -0.010684 | 0.23996 | 239.623 | 1 | 0.23967 | -0.35965 | 8.4e-05 | -5.3e-05 | -0.1499 | 0.0084 | -0.0223
voltage_sweep | 10 | 2 | 1 | 1 | 240 | 1 | 0.999 | 60 | 240.0279 | 0.99999 | 99.9005 | 0.23979 | -0.010706 | 0.24003 | 239.608 | 2.969 | 0.71131 | -0.41994 | 1.969 | 0.47152 | -0.175 | -1.0321 | -1.1188
voltage_sweep | 10 | 2 | 2 | 1 | 240 | 1 | 0.999 | 60 | 240.0279 | 0.99999 | 99.9005 | 0.23979 | -0.010706 | 0.24003 | 239.608 | 1.001 | 0.24002 | -0.41994 | 0.001012 | 0.000231 | -0.175 | 0.1012 | 0.0963
voltage_sweep | 10 | 2 | 3 | 1 | 240 | 1 | 0.999 | 60 | 240.0279 | 0.99999 | 99.9005 | 0.23979 | -0.010706 | 0.24003 | 239.608 | 1.002 | 0.2401 | -0.41994 | 0.002012 | 0.000316 | -0.175 | 0.2012 | 0.1317
voltage_sweep | 10 | 2 | 4 | 1 | 240 | 1 | 0.999 | 60 | 240.0279 | 0.99999 | 99.9005 | 0.23979 | -0.010706 | 0.24003 | 239.608 | 1 | 0.2397 | -0.41994 | 1.2e-05 | -8.8e-05 | -0.175 | 0.0012 | -0.0367
voltage_sweep | 10 | 2 | 5 | 1 | 240 | 1 | 0.999 | 60 | 240.0279 | 0.99999 | 99.9005 | 0.23979 | -0.010706 | 0.24003 | 239.608 | 1.001 | 0.23991 | -0.41994 | 0.001012 | 0.000123 | -0.175 | 0.1012 | 0.0512
voltage_sweep | 10 | 2 | 6 | 1 | 240 | 1 | 0.999 | 60 | 240.0279 | 0.99999 | 99.9005 | 0.23979 | -0.010706 | 0.24003 | 239.608 | 0.998 | 0.23935 | -0.41994 | -0.001988 | -0.000441 | -0.175 | -0.1988 | -0.184
voltage_sweep | 10 | 2 | 7 | 2 | 240 | 1 | 0.999 | 60 | 239.9734 | 1.0002 | 99.9039 | 0.23978 | -0.010518 | 0.24001 | 239.537 | 2.979 | 0.71373 | -0.43639 | 1.9788 | 0.47395 | -0.1818 | -0.7162 | -0.78084
voltage_sweep | 10 | 2 | 8 | 2 | 240 | 1 | 0.999 | 60 | 239.9734 | 1.0002 | 99.9039 | 0.23978 | -0.010518 | 0.24001 | 239.537 | 0.999 | 0.23944 | -0.43639 | -0.001163 | -0.000345 | -0.1818 | -0.1163 | -0.1439
voltage_sweep | 10 | 2 | 9 | 2 | 240 | 1 | 0.999 | 60 | 239.9734 | 1.0002 | 99.9039 | 0.23978 | -0.010518 | 0.24001 | 239.537 | 1 | 0.23977 | -0.43639 | -0.000163 | -1.2e-05 | -0.1818 | -0.0163 | -0.005
voltage_sweep | 10 | 2 | 10 | 2 | 240 | 1 | 0.999 | 60 | 239.9734 | 1.0002 | 99.9039 | 0.23978 | -0.010518 | 0.24001 | 239.537 | 1 | 0.23972 | -0.43639 | -0.000163 | -6.5e-05 | -0.1818 | -0.0163 | -0.0271
voltage_sweep | 10 | 2 | 11 | 2 | 240 | 1 | 0.999 | 60 | 239.9734 | 1.0002 | 99.9039 | 0.23978 | -0.010518 | 0.24001 | 239.537 | 1 | 0.2396 | -0.43639 | -0.000163 | -0.000186 | -0.1818 | -0.0163 | -0.0776
voltage_sweep | 10 | 2 | 12 | 2 | 240 | 1 | 0.999 | 60 | 239.9734 | 1.0002 | 99.9039 | 0.23978 | -0.010518 | 0.24001 | 239.537 | 1 | 0.23965 | -0.43639 | -0.000163 | -0.000134 | -0.1818 | -0.0163 | -0.0559
voltage_sweep | 10 | 2 | 13 | 3 | 240 | 1 | 0.999 | 60 | 239.9824 | 0.99991 | 99.9009 | 0.23972 | -0.010681 | 0.23996 | 239.598 | 1 | 0.23966 | -0.38444 | 8.6e-05 | -6.3e-05 | -0.1602 | 0.0086 | -0.0262
voltage_sweep | 10 | 2 | 14 | 3 | 240 | 1 | 0.999 | 60 | 239.9824 | 0.99991 | 99.9009 | 0.23972 | -0.010681 | 0.23996 | 239.598 | 0.999 | 0.23953 | -0.38444 | -0.000914 | -0.000191 | -0.1602 | -0.0914 | -0.0796
voltage_sweep | 10 | 2 | 15 | 3 | 240 | 1 | 0.999 | 60 | 239.9824 | 0.99991 | 99.9009 | 0.23972 | -0.010681 | 0.23996 | 239.598 | 1.001 | 0.23985 | -0.38444 | 0.001086 | 0.000129 | -0.1602 | 0.1086 | 0.0539
voltage_sweep | 10 | 2 | 16 | 3 | 240 | 1 | 0.999 | 60 | 239.9824 | 0.99991 | 99.9009 | 0.23972 | -0.010681 | 0.23996 | 239.598 | 0.999 | 0.23957 | -0.38444 | -0.000914 | -0.000149 | -0.1602 | -0.0914 | -0.0621
voltage_sweep | 10 | 2 | 17 | 3 | 240 | 1 | 0.999 | 60 | 239.9824 | 0.99991 | 99.9009 | 0.23972 | -0.010681 | 0.23996 | 239.598 | 1 | 0.23964 | -0.38444 | 8.6e-05 | -8.8e-05 | -0.1602 | 0.0086 | -0.0367
voltage_sweep | 10 | 2 | 18 | 3 | 240 | 1 | 0.999 | 60 | 239.9824 | 0.99991 | 99.9009 | 0.23972 | -0.010681 | 0.23996 | 239.598 | 1 | 0.23961 | -0.38444 | 8.6e-05 | -0.000112 | -0.1602 | 0.0086 | -0.0467
voltage_sweep | 10 | 3 | 1 | 1 | 240 | 1 | 0.999 | 60 | 240.023 | 0.99998 | 99.9004 | 0.23978 | -0.010703 | 0.24002 | 239.597 | 2.969 | 0.71155 | -0.426 | 1.969 | 0.47177 | -0.1775 | -1.0314 | -1.0821
voltage_sweep | 10 | 3 | 2 | 1 | 240 | 1 | 0.999 | 60 | 240.023 | 0.99998 | 99.9004 | 0.23978 | -0.010703 | 0.24002 | 239.597 | 1.001 | 0.24003 | -0.426 | 0.00102 | 0.00025 | -0.1775 | 0.102 | 0.1041
voltage_sweep | 10 | 3 | 3 | 1 | 240 | 1 | 0.999 | 60 | 240.023 | 0.99998 | 99.9004 | 0.23978 | -0.010703 | 0.24002 | 239.597 | 1.001 | 0.24009 | -0.426 | 0.00102 | 0.000311 | -0.1775 | 0.102 | 0.1296
voltage_sweep | 10 | 3 | 4 | 1 | 240 | 1 | 0.999 | 60 | 240.023 | 0.99998 | 99.9004 | 0.23978 | -0.010703 | 0.24002 | 239.597 | 1 | 0.2398 | -0.426 | 2e-05 | 2.1e-05 | -0.1775 | 0.002 | 0.0086
voltage_sweep | 10 | 3 | 5 | 1 | 240 | 1 | 0.999 | 60 | 240.023 | 0.99998 | 99.9004 | 0.23978 | -0.010703 | 0.24002 | 239.597 | 1.001 | 0.23998 | -0.426 | 0.00102 | 0.0002 | -0.1775 | 0.102 | 0.0833
voltage_sweep | 10 | 3 | 6 | 1 | 240 | 1 | 0.999 | 60 | 240.023 | 0.99998 | 99.9004 | 0.23978 | -0.010703 | 0.24002 | 239.597 | 0.998 | 0.23932 | -0.426 | -0.00198 | -0.00046 | -0.1775 | -0.198 | -0.192
voltage_sweep | 10 | 3 | 7 | 2 | 240 | 1 | 0.999 | 60 | 239.9575 | 1.0002 | 99.9035 | 0.23976 | -0.010541 | 0.24 | 239.535 | 2.979 | 0.71372 | -0.42246 | 1.9788 | 0.47396 | -0.1761 | -0.71578 | -0.77437
voltage_sweep | 10 | 3 | 8 | 2 | 240 | 1 | 0.999 | 60 | 239.9575 | 1.0002 | 99.9035 | 0.23976 | -0.010541 | 0.24 | 239.535 | 0.999 | 0.23935 | -0.42246 | -0.001159 | -0.000409 | -0.1761 | -0.1159 | -0.1706
voltage_sweep | 10 | 3 | 9 | 2 | 240 | 1 | 0.999 | 60 | 239.9575 | 1.0002 | 99.9035 | 0.23976 | -0.010541 | 0.24 | 239.535 | 1 | 0.23973 | -0.42246 | -0.000159 | -3.1e-05 | -0.1761 | -0.0159 | -0.0129
voltage_sweep | 10 | 3 | 10 | 2 | 240 | 1 | 0.999 | 60 | 239.9575 | 1.0002 | 99.9035 | 0.23976 | -0.010541 | 0.24 | 239.535 | 1 | 0.23971 | -0.42246 | -0.000159 | -5.1e-05 | -0.1761 | -0.0159 | -0.0213
voltage_sweep | 10 | 3 | 11 | 2 | 240 | 1 | 0.999 | 60 | 239.9575 | 1.0002 | 99.9035 | 0.23976 | -0.010541 | 0.24 | 239.535 | 1 | 0.23956 | -0.42246 | -0.000159 | -0.000203 | -0.1761 | -0.0159 | -0.0847
voltage_sweep | 10 | 3 | 12 | 2 | 240 | 1 | 0.999 | 60 | 239.9575 | 1.0002 | 99.9035 | 0.23976 | -0.010541 | 0.24 | 239.535 | 1 | 0.2396 | -0.42246 | -0.000159 | -0.000163 | -0.1761 | -0.0159 | -0.068
voltage_sweep | 10 | 3 | 13 | 3 | 240 | 1 | 0.999 | 60 | 239.9852 | 0.99992 | 99.9008 | 0.23973 | -0.010685 | 0.23997 | 239.624 | 1 | 0.23964 | -0.36118 | 8e-05 | -8.3e-05 | -0.1505 | 0.008 | -0.0346
voltage_sweep | 10 | 3 | 14 | 3 | 240 | 1 | 0.999 | 60 | 239.9852 | 0.99992 | 99.9008 | 0.23973 | -0.010685 | 0.23997 | 239.624 | 0.999 | 0.23956 | -0.36118 | -0.00092 | -0.000168 | -0.1505 | -0.092 | -0.07
voltage_sweep | 10 | 3 | 15 | 3 | 240 | 1 | 0.999 | 60 | 239.9852 | 0.99992 | 99.9008 | 0.23973 | -0.010685 | 0.23997 | 239.624 | 1 | 0.23986 | -0.36118 | 8e-05 | 0.000132 | -0.1505 | 0.008 | 0.0551
voltage_sweep | 10 | 3 | 16 | 3 | 240 | 1 | 0.999 | 60 | 239.9852 | 0.99992 | 99.9008 | 0.23973 | -0.010685 | 0.23997 | 239.624 | 0.999 | 0.23958 | -0.36118 | -0.00092 | -0.000149 | -0.1505 | -0.092 | -0.0621
voltage_sweep | 10 | 3 | 17 | 3 | 240 | 1 | 0.999 | 60 | 239.9852 | 0.99992 | 99.9008 | 0.23973 | -0.010685 | 0.23997 | 239.624 | 1 | 0.23966 | -0.36118 | 8e-05 | -6.3e-05 | -0.1505 | 0.008 | -0.0262
voltage_sweep | 10 | 3 | 18 | 3 | 240 | 1 | 0.999 | 60 | 239.9852 | 0.99992 | 99.9008 | 0.23973 | -0.010685 | 0.23997 | 239.624 | 1 | 0.23969 | -0.36118 | 8e-05 | -3.7e-05 | -0.1505 | 0.008 | -0.0154

## PF_Sweep
sweep | point | sample | channel | phase | source_vln | source_i | source_pf | source_freq | ref_vln | ref_i | ref_pf | ref_kW | ref_kVAR | ref_kVA | dut_vln | dut_i | dut_kW | err_vln | err_i | err_kW | pct_err_vln | pct_err_i | pct_err_kW
pf_sweep | 1 | 1 | 1 | 1 | 120 | 1 | -0.999 | 60 | 120.0009 | 1 | -99.899 | 0.11988 | 0.0053869 | 0.12 | 119.737 | 2.966 | 0.35535 | -0.26389 | 1.966 | 0.23547 | -0.2199 | -1.1338 | -1.1927
pf_sweep | 1 | 1 | 2 | 1 | 120 | 1 | -0.999 | 60 | 120.0009 | 1 | -99.899 | 0.11988 | 0.0053869 | 0.12 | 119.737 | 1.001 | 0.1199 | -0.26389 | 0.000996 | 2.5e-05 | -0.2199 | 0.0996 | 0.0207
pf_sweep | 1 | 1 | 3 | 1 | 120 | 1 | -0.999 | 60 | 120.0009 | 1 | -99.899 | 0.11988 | 0.0053869 | 0.12 | 119.737 | 1.001 | 0.11997 | -0.26389 | 0.000996 | 8.7e-05 | -0.2199 | 0.0996 | 0.0724
pf_sweep | 1 | 1 | 4 | 1 | 120 | 1 | -0.999 | 60 | 120.0009 | 1 | -99.899 | 0.11988 | 0.0053869 | 0.12 | 119.737 | 1 | 0.11985 | -0.26389 | -4e-06 | -3e-05 | -0.2199 | -0.0004 | -0.0252
pf_sweep | 1 | 1 | 5 | 1 | 120 | 1 | -0.999 | 60 | 120.0009 | 1 | -99.899 | 0.11988 | 0.0053869 | 0.12 | 119.737 | 1.001 | 0.1199 | -0.26389 | 0.000996 | 2.3e-05 | -0.2199 | 0.0996 | 0.019
pf_sweep | 1 | 1 | 6 | 1 | 120 | 1 | -0.999 | 60 | 120.0009 | 1 | -99.899 | 0.11988 | 0.0053869 | 0.12 | 119.737 | 0.998 | 0.11961 | -0.26389 | -0.002004 | -0.000271 | -0.2199 | -0.2004 | -0.2262
pf_sweep | 1 | 1 | 7 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054657 | 0.12 | 119.752 | 2.977 | 0.35657 | -0.23247 | 1.9768 | 0.23669 | -0.1937 | -0.78344 | -0.8524
pf_sweep | 1 | 1 | 8 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054657 | 0.12 | 119.752 | 1 | 0.11979 | -0.23247 | -0.000169 | -9.3e-05 | -0.1937 | -0.0169 | -0.0777
pf_sweep | 1 | 1 | 9 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054657 | 0.12 | 119.752 | 1.001 | 0.11988 | -0.23247 | 0.000831 | 3e-06 | -0.1937 | 0.0831 | 0.0023
pf_sweep | 1 | 1 | 10 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054657 | 0.12 | 119.752 | 1 | 0.11982 | -0.23247 | -0.000169 | -6e-05 | -0.1937 | -0.0169 | -0.0502
pf_sweep | 1 | 1 | 11 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054657 | 0.12 | 119.752 | 1 | 0.11976 | -0.23247 | -0.000169 | -0.000121 | -0.1937 | -0.0169 | -0.1011
pf_sweep | 1 | 1 | 12 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054657 | 0.12 | 119.752 | 1 | 0.11981 | -0.23247 | -0.000169 | -7e-05 | -0.1937 | -0.0169 | -0.0586
pf_sweep | 1 | 1 | 13 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99992 | -99.8993 | 0.11986 | 0.0053815 | 0.11998 | 119.776 | 0.999 | 0.11975 | -0.21563 | -0.000925 | -0.000115 | -0.1797 | -0.0925 | -0.0958
pf_sweep | 1 | 1 | 14 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99992 | -99.8993 | 0.11986 | 0.0053815 | 0.11998 | 119.776 | 0.999 | 0.11972 | -0.21563 | -0.000925 | -0.000146 | -0.1797 | -0.0925 | -0.1217
pf_sweep | 1 | 1 | 15 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99992 | -99.8993 | 0.11986 | 0.0053815 | 0.11998 | 119.776 | 1 | 0.11984 | -0.21563 | 7.5e-05 | -2e-05 | -0.1797 | 0.0075 | -0.0165
pf_sweep | 1 | 1 | 16 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99992 | -99.8993 | 0.11986 | 0.0053815 | 0.11998 | 119.776 | 0.999 | 0.11973 | -0.21563 | -0.000925 | -0.000129 | -0.1797 | -0.0925 | -0.1075
pf_sweep | 1 | 1 | 17 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99992 | -99.8993 | 0.11986 | 0.0053815 | 0.11998 | 119.776 | 1 | 0.1198 | -0.21563 | 7.5e-05 | -6.3e-05 | -0.1797 | 0.0075 | -0.0524
pf_sweep | 1 | 1 | 18 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99992 | -99.8993 | 0.11986 | 0.0053815 | 0.11998 | 119.776 | 0.999 | 0.11975 | -0.21563 | -0.000925 | -0.000113 | -0.1797 | -0.0925 | -0.0941
pf_sweep | 1 | 2 | 1 | 1 | 120 | 1 | -0.999 | 60 | 120.0329 | 0.99998 | -99.8989 | 0.11991 | 0.0053938 | 0.12003 | 119.73 | 2.967 | 0.35528 | -0.30288 | 1.967 | 0.23538 | -0.2523 | -1.0983 | -1.2355
pf_sweep | 1 | 2 | 2 | 1 | 120 | 1 | -0.999 | 60 | 120.0329 | 0.99998 | -99.8989 | 0.11991 | 0.0053938 | 0.12003 | 119.73 | 1.001 | 0.11989 | -0.30288 | 0.001018 | -1.7e-05 | -0.2523 | 0.1018 | -0.0145
pf_sweep | 1 | 2 | 3 | 1 | 120 | 1 | -0.999 | 60 | 120.0329 | 0.99998 | -99.8989 | 0.11991 | 0.0053938 | 0.12003 | 119.73 | 1.001 | 0.11992 | -0.30288 | 0.001018 | 1.2e-05 | -0.2523 | 0.1018 | 0.0096
pf_sweep | 1 | 2 | 4 | 1 | 120 | 1 | -0.999 | 60 | 120.0329 | 0.99998 | -99.8989 | 0.11991 | 0.0053938 | 0.12003 | 119.73 | 1 | 0.11982 | -0.30288 | 1.8e-05 | -9.1e-05 | -0.2523 | 0.0018 | -0.0763
pf_sweep | 1 | 2 | 5 | 1 | 120 | 1 | -0.999 | 60 | 120.0329 | 0.99998 | -99.8989 | 0.11991 | 0.0053938 | 0.12003 | 119.73 | 1.001 | 0.1199 | -0.30288 | 0.001018 | -1.3e-05 | -0.2523 | 0.1018 | -0.0112
pf_sweep | 1 | 2 | 6 | 1 | 120 | 1 | -0.999 | 60 | 120.0329 | 0.99998 | -99.8989 | 0.11991 | 0.0053938 | 0.12003 | 119.73 | 0.998 | 0.11958 | -0.30288 | -0.001982 | -0.000331 | -0.2523 | -0.1983 | -0.2764
pf_sweep | 1 | 2 | 7 | 2 | 120 | 1 | -0.999 | 60 | 119.9759 | 1.0002 | -99.896 | 0.11987 | 0.0054696 | 0.12 | 119.751 | 2.978 | 0.35668 | -0.22494 | 1.9778 | 0.23681 | -0.1875 | -0.75016 | -0.81658
pf_sweep | 1 | 2 | 8 | 2 | 120 | 1 | -0.999 | 60 | 119.9759 | 1.0002 | -99.896 | 0.11987 | 0.0054696 | 0.12 | 119.751 | 1 | 0.11978 | -0.22494 | -0.00017 | -9e-05 | -0.1875 | -0.0169 | -0.0747
pf_sweep | 1 | 2 | 9 | 2 | 120 | 1 | -0.999 | 60 | 119.9759 | 1.0002 | -99.896 | 0.11987 | 0.0054696 | 0.12 | 119.751 | 1.001 | 0.11989 | -0.22494 | 0.00083 | 2.1e-05 | -0.1875 | 0.083 | 0.0179
pf_sweep | 1 | 2 | 10 | 2 | 120 | 1 | -0.999 | 60 | 119.9759 | 1.0002 | -99.896 | 0.11987 | 0.0054696 | 0.12 | 119.751 | 1 | 0.11984 | -0.22494 | -0.00017 | -3.3e-05 | -0.1875 | -0.0169 | -0.0271
pf_sweep | 1 | 2 | 11 | 2 | 120 | 1 | -0.999 | 60 | 119.9759 | 1.0002 | -99.896 | 0.11987 | 0.0054696 | 0.12 | 119.751 | 1 | 0.11976 | -0.22494 | -0.00017 | -0.000113 | -0.1875 | -0.0169 | -0.0939
pf_sweep | 1 | 2 | 12 | 2 | 120 | 1 | -0.999 | 60 | 119.9759 | 1.0002 | -99.896 | 0.11987 | 0.0054696 | 0.12 | 119.751 | 1 | 0.11983 | -0.22494 | -0.00017 | -4.5e-05 | -0.1875 | -0.0169 | -0.0371
pf_sweep | 1 | 2 | 13 | 3 | 120 | 1 | -0.999 | 60 | 119.991 | 0.99991 | -99.8994 | 0.11986 | 0.0053797 | 0.11998 | 119.778 | 0.999 | 0.11974 | -0.21298 | -0.000913 | -0.000119 | -0.1775 | -0.0913 | -0.0992
pf_sweep | 1 | 2 | 14 | 3 | 120 | 1 | -0.999 | 60 | 119.991 | 0.99991 | -99.8994 | 0.11986 | 0.0053797 | 0.11998 | 119.778 | 0.999 | 0.11971 | -0.21298 | -0.000913 | -0.000151 | -0.1775 | -0.0913 | -0.1259
pf_sweep | 1 | 2 | 15 | 3 | 120 | 1 | -0.999 | 60 | 119.991 | 0.99991 | -99.8994 | 0.11986 | 0.0053797 | 0.11998 | 119.778 | 1 | 0.11985 | -0.21298 | 8.7e-05 | -6e-06 | -0.1775 | 0.0087 | -0.0049
pf_sweep | 1 | 2 | 16 | 3 | 120 | 1 | -0.999 | 60 | 119.991 | 0.99991 | -99.8994 | 0.11986 | 0.0053797 | 0.11998 | 119.778 | 0.999 | 0.11971 | -0.21298 | -0.000913 | -0.000146 | -0.1775 | -0.0913 | -0.1217
pf_sweep | 1 | 2 | 17 | 3 | 120 | 1 | -0.999 | 60 | 119.991 | 0.99991 | -99.8994 | 0.11986 | 0.0053797 | 0.11998 | 119.778 | 1 | 0.11979 | -0.21298 | 8.7e-05 | -7.2e-05 | -0.1775 | 0.0087 | -0.06
pf_sweep | 1 | 2 | 18 | 3 | 120 | 1 | -0.999 | 60 | 119.991 | 0.99991 | -99.8994 | 0.11986 | 0.0053797 | 0.11998 | 119.778 | 0.999 | 0.11975 | -0.21298 | -0.000913 | -0.000112 | -0.1775 | -0.0913 | -0.0934
pf_sweep | 1 | 3 | 1 | 1 | 120 | 1 | -0.999 | 60 | 120.0032 | 0.99996 | -99.8988 | 0.11988 | 0.0053863 | 0.12 | 119.716 | 2.967 | 0.35535 | -0.28718 | 1.967 | 0.23548 | -0.2393 | -1.0962 | -1.1891
pf_sweep | 1 | 3 | 2 | 1 | 120 | 1 | -0.999 | 60 | 120.0032 | 0.99996 | -99.8988 | 0.11988 | 0.0053863 | 0.12 | 119.716 | 1.001 | 0.11981 | -0.28718 | 0.001038 | -6.2e-05 | -0.2393 | 0.1038 | -0.0519
pf_sweep | 1 | 3 | 3 | 1 | 120 | 1 | -0.999 | 60 | 120.0032 | 0.99996 | -99.8988 | 0.11988 | 0.0053863 | 0.12 | 119.716 | 1.001 | 0.11991 | -0.28718 | 0.001038 | 3e-05 | -0.2393 | 0.1038 | 0.0249
pf_sweep | 1 | 3 | 4 | 1 | 120 | 1 | -0.999 | 60 | 120.0032 | 0.99996 | -99.8988 | 0.11988 | 0.0053863 | 0.12 | 119.716 | 1 | 0.11983 | -0.28718 | 3.8e-05 | -4.6e-05 | -0.2393 | 0.0038 | -0.0385
pf_sweep | 1 | 3 | 5 | 1 | 120 | 1 | -0.999 | 60 | 120.0032 | 0.99996 | -99.8988 | 0.11988 | 0.0053863 | 0.12 | 119.716 | 1.001 | 0.11983 | -0.28718 | 0.001038 | -5.2e-05 | -0.2393 | 0.1038 | -0.0435
pf_sweep | 1 | 3 | 6 | 1 | 120 | 1 | -0.999 | 60 | 120.0032 | 0.99996 | -99.8988 | 0.11988 | 0.0053863 | 0.12 | 119.716 | 0.998 | 0.11953 | -0.28718 | -0.001962 | -0.000352 | -0.2393 | -0.1962 | -0.2938
pf_sweep | 1 | 3 | 7 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054659 | 0.12 | 119.758 | 2.977 | 0.35661 | -0.22649 | 1.9768 | 0.23673 | -0.1888 | -0.78163 | -0.84089
pf_sweep | 1 | 3 | 8 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054659 | 0.12 | 119.758 | 1 | 0.11978 | -0.22649 | -0.000151 | -9.8e-05 | -0.1888 | -0.0151 | -0.0818
pf_sweep | 1 | 3 | 9 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054659 | 0.12 | 119.758 | 1 | 0.11986 | -0.22649 | -0.000151 | -1.4e-05 | -0.1888 | -0.0151 | -0.0117
pf_sweep | 1 | 3 | 10 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054659 | 0.12 | 119.758 | 1 | 0.11983 | -0.22649 | -0.000151 | -4.9e-05 | -0.1888 | -0.0151 | -0.0409
pf_sweep | 1 | 3 | 11 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054659 | 0.12 | 119.758 | 1 | 0.11975 | -0.22649 | -0.000151 | -0.000129 | -0.1888 | -0.0151 | -0.1076
pf_sweep | 1 | 3 | 12 | 2 | 120 | 1 | -0.999 | 60 | 119.9845 | 1.0002 | -99.8962 | 0.11988 | 0.0054659 | 0.12 | 119.758 | 1 | 0.1198 | -0.22649 | -0.000151 | -8e-05 | -0.1888 | -0.0151 | -0.0668
pf_sweep | 1 | 3 | 13 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99993 | -99.8993 | 0.11986 | 0.005383 | 0.11998 | 119.777 | 0.999 | 0.11975 | -0.21461 | -0.000928 | -0.00011 | -0.1789 | -0.0928 | -0.0919
pf_sweep | 1 | 3 | 14 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99993 | -99.8993 | 0.11986 | 0.005383 | 0.11998 | 119.777 | 0.999 | 0.11972 | -0.21461 | -0.000928 | -0.000143 | -0.1789 | -0.0928 | -0.1194
pf_sweep | 1 | 3 | 15 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99993 | -99.8993 | 0.11986 | 0.005383 | 0.11998 | 119.777 | 1 | 0.11989 | -0.21461 | 7.2e-05 | 2.3e-05 | -0.1789 | 0.0072 | 0.0191
pf_sweep | 1 | 3 | 16 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99993 | -99.8993 | 0.11986 | 0.005383 | 0.11998 | 119.777 | 0.999 | 0.11972 | -0.21461 | -0.000928 | -0.000137 | -0.1789 | -0.0928 | -0.1144
pf_sweep | 1 | 3 | 17 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99993 | -99.8993 | 0.11986 | 0.005383 | 0.11998 | 119.777 | 1 | 0.11979 | -0.21461 | 7.2e-05 | -6.8e-05 | -0.1789 | 0.0072 | -0.0569
pf_sweep | 1 | 3 | 18 | 3 | 120 | 1 | -0.999 | 60 | 119.9916 | 0.99993 | -99.8993 | 0.11986 | 0.005383 | 0.11998 | 119.777 | 1 | 0.11979 | -0.21461 | 7.2e-05 | -7.5e-05 | -0.1789 | 0.0072 | -0.0627
pf_sweep | 2 | 1 | 1 | 1 | 120 | 1 | -0.899 | 60 | 120.0246 | 0.99997 | -89.8903 | 0.10789 | 0.052587 | 0.12002 | 119.76 | 2.967 | 0.3149 | -0.2646 | 1.967 | 0.20701 | -0.2205 | -1.0975 | -2.7072
pf_sweep | 2 | 1 | 2 | 1 | 120 | 1 | -0.899 | 60 | 120.0246 | 0.99997 | -89.8903 | 0.10789 | 0.052587 | 0.12002 | 119.76 | 1.001 | 0.10657 | -0.2646 | 0.001026 | -0.001321 | -0.2205 | 0.1026 | -1.2241
pf_sweep | 2 | 1 | 3 | 1 | 120 | 1 | -0.899 | 60 | 120.0246 | 0.99997 | -89.8903 | 0.10789 | 0.052587 | 0.12002 | 119.76 | 1.001 | 0.10658 | -0.2646 | 0.001026 | -0.001307 | -0.2205 | 0.1026 | -1.2112
pf_sweep | 2 | 1 | 4 | 1 | 120 | 1 | -0.899 | 60 | 120.0246 | 0.99997 | -89.8903 | 0.10789 | 0.052587 | 0.12002 | 119.76 | 1 | 0.10648 | -0.2646 | 2.6e-05 | -0.001403 | -0.2205 | 0.0026 | -1.3001
pf_sweep | 2 | 1 | 5 | 1 | 120 | 1 | -0.899 | 60 | 120.0246 | 0.99997 | -89.8903 | 0.10789 | 0.052587 | 0.12002 | 119.76 | 1.001 | 0.10661 | -0.2646 | 0.001026 | -0.001281 | -0.2205 | 0.1026 | -1.1871
pf_sweep | 2 | 1 | 6 | 1 | 120 | 1 | -0.899 | 60 | 120.0246 | 0.99997 | -89.8903 | 0.10789 | 0.052587 | 0.12002 | 119.76 | 0.998 | 0.10654 | -0.2646 | -0.001974 | -0.001344 | -0.2205 | -0.1974 | -1.2455
pf_sweep | 2 | 1 | 7 | 2 | 120 | 1 | -0.899 | 60 | 119.9804 | 1.0002 | -89.8628 | 0.10784 | 0.052646 | 0.12 | 119.771 | 2.978 | 0.31624 | -0.20936 | 1.9778 | 0.2084 | -0.1745 | -0.75157 | -2.2474
pf_sweep | 2 | 1 | 8 | 2 | 120 | 1 | -0.899 | 60 | 119.9804 | 1.0002 | -89.8628 | 0.10784 | 0.052646 | 0.12 | 119.771 | 1 | 0.10652 | -0.20936 | -0.000184 | -0.001314 | -0.1745 | -0.0184 | -1.218
pf_sweep | 2 | 1 | 9 | 2 | 120 | 1 | -0.899 | 60 | 119.9804 | 1.0002 | -89.8628 | 0.10784 | 0.052646 | 0.12 | 119.771 | 1.001 | 0.10669 | -0.20936 | 0.000816 | -0.001149 | -0.1745 | 0.0816 | -1.065
pf_sweep | 2 | 1 | 10 | 2 | 120 | 1 | -0.899 | 60 | 119.9804 | 1.0002 | -89.8628 | 0.10784 | 0.052646 | 0.12 | 119.771 | 1 | 0.10662 | -0.20936 | -0.000184 | -0.00122 | -0.1745 | -0.0184 | -1.1309
pf_sweep | 2 | 1 | 11 | 2 | 120 | 1 | -0.899 | 60 | 119.9804 | 1.0002 | -89.8628 | 0.10784 | 0.052646 | 0.12 | 119.771 | 1 | 0.10664 | -0.20936 | -0.000184 | -0.001198 | -0.1745 | -0.0184 | -1.1105
pf_sweep | 2 | 1 | 12 | 2 | 120 | 1 | -0.899 | 60 | 119.9804 | 1.0002 | -89.8628 | 0.10784 | 0.052646 | 0.12 | 119.771 | 1 | 0.10655 | -0.20936 | -0.000184 | -0.001284 | -0.1745 | -0.0184 | -1.1902
pf_sweep | 2 | 1 | 13 | 3 | 120 | 1 | -0.899 | 60 | 119.993 | 0.9999 | -89.8934 | 0.10785 | 0.052562 | 0.11998 | 119.761 | 0.999 | 0.10654 | -0.23197 | -0.000899 | -0.001316 | -0.1933 | -0.0899 | -1.2201
pf_sweep | 2 | 1 | 14 | 3 | 120 | 1 | -0.899 | 60 | 119.993 | 0.9999 | -89.8934 | 0.10785 | 0.052562 | 0.11998 | 119.761 | 0.999 | 0.10663 | -0.23197 | -0.000899 | -0.001226 | -0.1933 | -0.0899 | -1.1366
pf_sweep | 2 | 1 | 15 | 3 | 120 | 1 | -0.899 | 60 | 119.993 | 0.9999 | -89.8934 | 0.10785 | 0.052562 | 0.11998 | 119.761 | 1 | 0.10671 | -0.23197 | 0.000101 | -0.001149 | -0.1933 | 0.0101 | -1.0652
pf_sweep | 2 | 1 | 16 | 3 | 120 | 1 | -0.899 | 60 | 119.993 | 0.9999 | -89.8934 | 0.10785 | 0.052562 | 0.11998 | 119.761 | 0.999 | 0.10664 | -0.23197 | -0.000899 | -0.001213 | -0.1933 | -0.0899 | -1.1246
pf_sweep | 2 | 1 | 17 | 3 | 120 | 1 | -0.899 | 60 | 119.993 | 0.9999 | -89.8934 | 0.10785 | 0.052562 | 0.11998 | 119.761 | 1 | 0.10667 | -0.23197 | 0.000101 | -0.001188 | -0.1933 | 0.0101 | -1.1014
pf_sweep | 2 | 1 | 18 | 3 | 120 | 1 | -0.899 | 60 | 119.993 | 0.9999 | -89.8934 | 0.10785 | 0.052562 | 0.11998 | 119.761 | 1 | 0.10667 | -0.23197 | 0.000101 | -0.001188 | -0.1933 | 0.0101 | -1.1014
pf_sweep | 2 | 2 | 1 | 1 | 120 | 1 | -0.899 | 60 | 120.0252 | 0.99997 | -89.8911 | 0.10789 | 0.052585 | 0.12002 | 119.736 | 2.968 | 0.31498 | -0.28921 | 1.968 | 0.20709 | -0.241 | -1.0632 | -2.6845
pf_sweep | 2 | 2 | 2 | 1 | 120 | 1 | -0.899 | 60 | 120.0252 | 0.99997 | -89.8911 | 0.10789 | 0.052585 | 0.12002 | 119.736 | 1.001 | 0.10656 | -0.28921 | 0.001035 | -0.001328 | -0.241 | 0.1035 | -1.2311
pf_sweep | 2 | 2 | 3 | 1 | 120 | 1 | -0.899 | 60 | 120.0252 | 0.99997 | -89.8911 | 0.10789 | 0.052585 | 0.12002 | 119.736 | 1.001 | 0.10656 | -0.28921 | 0.001035 | -0.00133 | -0.241 | 0.1035 | -1.233
pf_sweep | 2 | 2 | 4 | 1 | 120 | 1 | -0.899 | 60 | 120.0252 | 0.99997 | -89.8911 | 0.10789 | 0.052585 | 0.12002 | 119.736 | 1 | 0.1065 | -0.28921 | 3.5e-05 | -0.001383 | -0.241 | 0.0035 | -1.2821
pf_sweep | 2 | 2 | 5 | 1 | 120 | 1 | -0.899 | 60 | 120.0252 | 0.99997 | -89.8911 | 0.10789 | 0.052585 | 0.12002 | 119.736 | 1.001 | 0.10658 | -0.28921 | 0.001035 | -0.001306 | -0.241 | 0.1035 | -1.2107
pf_sweep | 2 | 2 | 6 | 1 | 120 | 1 | -0.899 | 60 | 120.0252 | 0.99997 | -89.8911 | 0.10789 | 0.052585 | 0.12002 | 119.736 | 0.998 | 0.10651 | -0.28921 | -0.001965 | -0.001377 | -0.241 | -0.1965 | -1.2765
pf_sweep | 2 | 2 | 7 | 2 | 120 | 1 | -0.899 | 60 | 119.9796 | 1.0002 | -89.8611 | 0.10784 | 0.052651 | 0.12 | 119.757 | 2.978 | 0.31613 | -0.22256 | 1.9778 | 0.2083 | -0.1855 | -0.75274 | -2.28
pf_sweep | 2 | 2 | 8 | 2 | 120 | 1 | -0.899 | 60 | 119.9796 | 1.0002 | -89.8611 | 0.10784 | 0.052651 | 0.12 | 119.757 | 1 | 0.10658 | -0.22256 | -0.000196 | -0.001254 | -0.1855 | -0.0195 | -1.1629
pf_sweep | 2 | 2 | 9 | 2 | 120 | 1 | -0.899 | 60 | 119.9796 | 1.0002 | -89.8611 | 0.10784 | 0.052651 | 0.12 | 119.757 | 1.001 | 0.1067 | -0.22256 | 0.000804 | -0.00114 | -0.1855 | 0.0804 | -1.0572
pf_sweep | 2 | 2 | 10 | 2 | 120 | 1 | -0.899 | 60 | 119.9796 | 1.0002 | -89.8611 | 0.10784 | 0.052651 | 0.12 | 119.757 | 1 | 0.1066 | -0.22256 | -0.000196 | -0.00124 | -0.1855 | -0.0195 | -1.1499
pf_sweep | 2 | 2 | 11 | 2 | 120 | 1 | -0.899 | 60 | 119.9796 | 1.0002 | -89.8611 | 0.10784 | 0.052651 | 0.12 | 119.757 | 1 | 0.10669 | -0.22256 | -0.000196 | -0.001148 | -0.1855 | -0.0195 | -1.0646
pf_sweep | 2 | 2 | 12 | 2 | 120 | 1 | -0.899 | 60 | 119.9796 | 1.0002 | -89.8611 | 0.10784 | 0.052651 | 0.12 | 119.757 | 1 | 0.10655 | -0.22256 | -0.000196 | -0.001286 | -0.1855 | -0.0195 | -1.1925
pf_sweep | 2 | 2 | 13 | 3 | 120 | 1 | -0.899 | 60 | 119.9844 | 0.9999 | -89.8934 | 0.10785 | 0.052558 | 0.11997 | 119.768 | 0.999 | 0.10654 | -0.21638 | -0.0009 | -0.001308 | -0.1803 | -0.09 | -1.213
pf_sweep | 2 | 2 | 14 | 3 | 120 | 1 | -0.899 | 60 | 119.9844 | 0.9999 | -89.8934 | 0.10785 | 0.052558 | 0.11997 | 119.768 | 0.999 | 0.10663 | -0.21638 | -0.0009 | -0.001215 | -0.1803 | -0.09 | -1.1268
pf_sweep | 2 | 2 | 15 | 3 | 120 | 1 | -0.899 | 60 | 119.9844 | 0.9999 | -89.8934 | 0.10785 | 0.052558 | 0.11997 | 119.768 | 1 | 0.10667 | -0.21638 | 0.0001 | -0.00118 | -0.1803 | 0.01 | -1.0944
pf_sweep | 2 | 2 | 16 | 3 | 120 | 1 | -0.899 | 60 | 119.9844 | 0.9999 | -89.8934 | 0.10785 | 0.052558 | 0.11997 | 119.768 | 0.999 | 0.10664 | -0.21638 | -0.0009 | -0.001205 | -0.1803 | -0.09 | -1.1175
pf_sweep | 2 | 2 | 17 | 3 | 120 | 1 | -0.899 | 60 | 119.9844 | 0.9999 | -89.8934 | 0.10785 | 0.052558 | 0.11997 | 119.768 | 1 | 0.10665 | -0.21638 | 0.0001 | -0.001196 | -0.1803 | 0.01 | -1.1092
pf_sweep | 2 | 2 | 18 | 3 | 120 | 1 | -0.899 | 60 | 119.9844 | 0.9999 | -89.8934 | 0.10785 | 0.052558 | 0.11997 | 119.768 | 1 | 0.10664 | -0.21638 | 0.0001 | -0.00121 | -0.1803 | 0.01 | -1.1222
pf_sweep | 2 | 3 | 1 | 1 | 120 | 1 | -0.899 | 60 | 119.9938 | 0.99998 | -89.8922 | 0.10786 | 0.052568 | 0.11999 | 119.749 | 2.967 | 0.31478 | -0.24477 | 1.967 | 0.20692 | -0.204 | -1.098 | -2.7212
pf_sweep | 2 | 3 | 2 | 1 | 120 | 1 | -0.899 | 60 | 119.9938 | 0.99998 | -89.8922 | 0.10786 | 0.052568 | 0.11999 | 119.749 | 1.001 | 0.10658 | -0.24477 | 0.00102 | -0.001286 | -0.204 | 0.102 | -1.1921
pf_sweep | 2 | 3 | 3 | 1 | 120 | 1 | -0.899 | 60 | 119.9938 | 0.99998 | -89.8922 | 0.10786 | 0.052568 | 0.11999 | 119.749 | 1.001 | 0.10656 | -0.24477 | 0.00102 | -0.001303 | -0.204 | 0.102 | -1.2079
pf_sweep | 2 | 3 | 4 | 1 | 120 | 1 | -0.899 | 60 | 119.9938 | 0.99998 | -89.8922 | 0.10786 | 0.052568 | 0.11999 | 119.749 | 1 | 0.10646 | -0.24477 | 2e-05 | -0.001401 | -0.204 | 0.002 | -1.2987
pf_sweep | 2 | 3 | 5 | 1 | 120 | 1 | -0.899 | 60 | 119.9938 | 0.99998 | -89.8922 | 0.10786 | 0.052568 | 0.11999 | 119.749 | 1.001 | 0.10663 | -0.24477 | 0.00102 | -0.001236 | -0.204 | 0.102 | -1.1458
pf_sweep | 2 | 3 | 6 | 1 | 120 | 1 | -0.899 | 60 | 119.9938 | 0.99998 | -89.8922 | 0.10786 | 0.052568 | 0.11999 | 119.749 | 0.998 | 0.10651 | -0.24477 | -0.00198 | -0.00135 | -0.204 | -0.198 | -1.2514
pf_sweep | 2 | 3 | 7 | 2 | 120 | 1 | -0.899 | 60 | 119.9866 | 1.0002 | -89.862 | 0.10784 | 0.052651 | 0.12001 | 119.753 | 2.978 | 0.31617 | -0.23362 | 1.9778 | 0.20833 | -0.1947 | -0.75149 | -2.2747
pf_sweep | 2 | 3 | 8 | 2 | 120 | 1 | -0.899 | 60 | 119.9866 | 1.0002 | -89.862 | 0.10784 | 0.052651 | 0.12001 | 119.753 | 1 | 0.10653 | -0.23362 | -0.000183 | -0.001313 | -0.1947 | -0.0183 | -1.2176
pf_sweep | 2 | 3 | 9 | 2 | 120 | 1 | -0.899 | 60 | 119.9866 | 1.0002 | -89.862 | 0.10784 | 0.052651 | 0.12001 | 119.753 | 1.001 | 0.10668 | -0.23362 | 0.000817 | -0.001165 | -0.1947 | 0.0817 | -1.0804
pf_sweep | 2 | 3 | 10 | 2 | 120 | 1 | -0.899 | 60 | 119.9866 | 1.0002 | -89.862 | 0.10784 | 0.052651 | 0.12001 | 119.753 | 1 | 0.10661 | -0.23362 | -0.000183 | -0.001233 | -0.1947 | -0.0183 | -1.1434
pf_sweep | 2 | 3 | 11 | 2 | 120 | 1 | -0.899 | 60 | 119.9866 | 1.0002 | -89.862 | 0.10784 | 0.052651 | 0.12001 | 119.753 | 1 | 0.10664 | -0.23362 | -0.000183 | -0.001197 | -0.1947 | -0.0183 | -1.1101
pf_sweep | 2 | 3 | 12 | 2 | 120 | 1 | -0.899 | 60 | 119.9866 | 1.0002 | -89.862 | 0.10784 | 0.052651 | 0.12001 | 119.753 | 1 | 0.10653 | -0.23362 | -0.000183 | -0.00131 | -0.1947 | -0.0183 | -1.2148
pf_sweep | 2 | 3 | 13 | 3 | 120 | 1 | -0.899 | 60 | 119.9906 | 0.9999 | -89.8925 | 0.10785 | 0.052563 | 0.11998 | 119.764 | 0.999 | 0.10654 | -0.22663 | -0.000899 | -0.001316 | -0.1889 | -0.0899 | -1.2199
pf_sweep | 2 | 3 | 14 | 3 | 120 | 1 | -0.899 | 60 | 119.9906 | 0.9999 | -89.8925 | 0.10785 | 0.052563 | 0.11998 | 119.764 | 0.999 | 0.10659 | -0.22663 | -0.000899 | -0.001266 | -0.1889 | -0.0899 | -1.1736
pf_sweep | 2 | 3 | 15 | 3 | 120 | 1 | -0.899 | 60 | 119.9906 | 0.9999 | -89.8925 | 0.10785 | 0.052563 | 0.11998 | 119.764 | 1 | 0.1067 | -0.22663 | 0.000101 | -0.00115 | -0.1889 | 0.0101 | -1.066
pf_sweep | 2 | 3 | 16 | 3 | 120 | 1 | -0.899 | 60 | 119.9906 | 0.9999 | -89.8925 | 0.10785 | 0.052563 | 0.11998 | 119.764 | 0.999 | 0.10663 | -0.22663 | -0.000899 | -0.001222 | -0.1889 | -0.0899 | -1.1328
pf_sweep | 2 | 3 | 17 | 3 | 120 | 1 | -0.899 | 60 | 119.9906 | 0.9999 | -89.8925 | 0.10785 | 0.052563 | 0.11998 | 119.764 | 0.999 | 0.10661 | -0.22663 | -0.000899 | -0.001247 | -0.1889 | -0.0899 | -1.156
pf_sweep | 2 | 3 | 18 | 3 | 120 | 1 | -0.899 | 60 | 119.9906 | 0.9999 | -89.8925 | 0.10785 | 0.052563 | 0.11998 | 119.764 | 1 | 0.10666 | -0.22663 | 0.000101 | -0.001189 | -0.1889 | 0.0101 | -1.1022
pf_sweep | 3 | 1 | 1 | 1 | 120 | 1 | -0.799 | 60 | 120.0131 | 0.99996 | -79.8893 | 0.095874 | 0.072181 | 0.12001 | 119.662 | 2.968 | 0.27719 | -0.35112 | 1.968 | 0.18132 | -0.2926 | -1.0628 | -3.6259
pf_sweep | 3 | 1 | 2 | 1 | 120 | 1 | -0.799 | 60 | 120.0131 | 0.99996 | -79.8893 | 0.095874 | 0.072181 | 0.12001 | 119.662 | 1.001 | 0.093887 | -0.35112 | 0.001039 | -0.001987 | -0.2926 | 0.1039 | -2.0724
pf_sweep | 3 | 1 | 3 | 1 | 120 | 1 | -0.799 | 60 | 120.0131 | 0.99996 | -79.8893 | 0.095874 | 0.072181 | 0.12001 | 119.662 | 1.001 | 0.093817 | -0.35112 | 0.001039 | -0.002057 | -0.2926 | 0.1039 | -2.1454
pf_sweep | 3 | 1 | 4 | 1 | 120 | 1 | -0.799 | 60 | 120.0131 | 0.99996 | -79.8893 | 0.095874 | 0.072181 | 0.12001 | 119.662 | 1 | 0.093851 | -0.35112 | 3.9e-05 | -0.002023 | -0.2926 | 0.0039 | -2.11
pf_sweep | 3 | 1 | 5 | 1 | 120 | 1 | -0.799 | 60 | 120.0131 | 0.99996 | -79.8893 | 0.095874 | 0.072181 | 0.12001 | 119.662 | 1.001 | 0.093919 | -0.35112 | 0.001039 | -0.001955 | -0.2926 | 0.1039 | -2.039
pf_sweep | 3 | 1 | 6 | 1 | 120 | 1 | -0.799 | 60 | 120.0131 | 0.99996 | -79.8893 | 0.095874 | 0.072181 | 0.12001 | 119.662 | 0.998 | 0.093901 | -0.35112 | -0.001961 | -0.001973 | -0.2926 | -0.1961 | -2.0578
pf_sweep | 3 | 1 | 7 | 2 | 120 | 1 | -0.799 | 60 | 119.9878 | 1.0002 | -79.8527 | 0.095834 | 0.072243 | 0.12001 | 119.778 | 2.978 | 0.27807 | -0.20981 | 1.9778 | 0.18224 | -0.1749 | -0.75416 | -3.2789
pf_sweep | 3 | 1 | 8 | 2 | 120 | 1 | -0.799 | 60 | 119.9878 | 1.0002 | -79.8527 | 0.095834 | 0.072243 | 0.12001 | 119.778 | 1 | 0.093941 | -0.20981 | -0.00021 | -0.001893 | -0.1749 | -0.021 | -1.9749
pf_sweep | 3 | 1 | 9 | 2 | 120 | 1 | -0.799 | 60 | 119.9878 | 1.0002 | -79.8527 | 0.095834 | 0.072243 | 0.12001 | 119.778 | 1.001 | 0.094053 | -0.20981 | 0.00079 | -0.001781 | -0.1749 | 0.079 | -1.858
pf_sweep | 3 | 1 | 10 | 2 | 120 | 1 | -0.799 | 60 | 119.9878 | 1.0002 | -79.8527 | 0.095834 | 0.072243 | 0.12001 | 119.778 | 1 | 0.093934 | -0.20981 | -0.00021 | -0.0019 | -0.1749 | -0.021 | -1.9822
pf_sweep | 3 | 1 | 11 | 2 | 120 | 1 | -0.799 | 60 | 119.9878 | 1.0002 | -79.8527 | 0.095834 | 0.072243 | 0.12001 | 119.778 | 1 | 0.094107 | -0.20981 | -0.00021 | -0.001727 | -0.1749 | -0.021 | -1.8017
pf_sweep | 3 | 1 | 12 | 2 | 120 | 1 | -0.799 | 60 | 119.9878 | 1.0002 | -79.8527 | 0.095834 | 0.072243 | 0.12001 | 119.778 | 1 | 0.093906 | -0.20981 | -0.00021 | -0.001928 | -0.1749 | -0.021 | -2.0114
pf_sweep | 3 | 1 | 13 | 3 | 120 | 1 | -0.799 | 60 | 119.9871 | 0.99992 | -79.8921 | 0.095852 | 0.072158 | 0.11998 | 119.759 | 1 | 0.093987 | -0.22811 | 8.4e-05 | -0.001865 | -0.1901 | 0.0084 | -1.9459
pf_sweep | 3 | 1 | 14 | 3 | 120 | 1 | -0.799 | 60 | 119.9871 | 0.99992 | -79.8921 | 0.095852 | 0.072158 | 0.11998 | 119.759 | 0.999 | 0.094089 | -0.22811 | -0.000916 | -0.001763 | -0.1901 | -0.0916 | -1.8395
pf_sweep | 3 | 1 | 15 | 3 | 120 | 1 | -0.799 | 60 | 119.9871 | 0.99992 | -79.8921 | 0.095852 | 0.072158 | 0.11998 | 119.759 | 1 | 0.094077 | -0.22811 | 8.4e-05 | -0.001775 | -0.1901 | 0.0084 | -1.852
pf_sweep | 3 | 1 | 16 | 3 | 120 | 1 | -0.799 | 60 | 119.9871 | 0.99992 | -79.8921 | 0.095852 | 0.072158 | 0.11998 | 119.759 | 0.999 | 0.094143 | -0.22811 | -0.000916 | -0.001709 | -0.1901 | -0.0916 | -1.7832
pf_sweep | 3 | 1 | 17 | 3 | 120 | 1 | -0.799 | 60 | 119.9871 | 0.99992 | -79.8921 | 0.095852 | 0.072158 | 0.11998 | 119.759 | 1 | 0.094061 | -0.22811 | 8.4e-05 | -0.001791 | -0.1901 | 0.0084 | -1.8687
pf_sweep | 3 | 1 | 18 | 3 | 120 | 1 | -0.799 | 60 | 119.9871 | 0.99992 | -79.8921 | 0.095852 | 0.072158 | 0.11998 | 119.759 | 0.999 | 0.094063 | -0.22811 | -0.000916 | -0.001789 | -0.1901 | -0.0916 | -1.8666
pf_sweep | 3 | 2 | 1 | 1 | 120 | 1 | -0.799 | 60 | 120.0084 | 0.99995 | -79.8896 | 0.095869 | 0.072177 | 0.12 | 119.732 | 2.968 | 0.27711 | -0.27644 | 1.9681 | 0.18124 | -0.2303 | -1.0617 | -3.6502
pf_sweep | 3 | 2 | 2 | 1 | 120 | 1 | -0.799 | 60 | 120.0084 | 0.99995 | -79.8896 | 0.095869 | 0.072177 | 0.12 | 119.732 | 1.001 | 0.093983 | -0.27644 | 0.00105 | -0.001886 | -0.2303 | 0.105 | -1.9677
pf_sweep | 3 | 2 | 3 | 1 | 120 | 1 | -0.799 | 60 | 120.0084 | 0.99995 | -79.8896 | 0.095869 | 0.072177 | 0.12 | 119.732 | 1.001 | 0.093939 | -0.27644 | 0.00105 | -0.00193 | -0.2303 | 0.105 | -2.0136
pf_sweep | 3 | 2 | 4 | 1 | 120 | 1 | -0.799 | 60 | 120.0084 | 0.99995 | -79.8896 | 0.095869 | 0.072177 | 0.12 | 119.732 | 1 | 0.093829 | -0.27644 | 5e-05 | -0.00204 | -0.2303 | 0.005 | -2.1283
pf_sweep | 3 | 2 | 5 | 1 | 120 | 1 | -0.799 | 60 | 120.0084 | 0.99995 | -79.8896 | 0.095869 | 0.072177 | 0.12 | 119.732 | 1.001 | 0.094015 | -0.27644 | 0.00105 | -0.001854 | -0.2303 | 0.105 | -1.9343
pf_sweep | 3 | 2 | 6 | 1 | 120 | 1 | -0.799 | 60 | 120.0084 | 0.99995 | -79.8896 | 0.095869 | 0.072177 | 0.12 | 119.732 | 0.998 | 0.094004 | -0.27644 | -0.00195 | -0.001865 | -0.2303 | -0.195 | -1.9458
pf_sweep | 3 | 2 | 7 | 2 | 120 | 1 | -0.799 | 60 | 119.9819 | 1.0002 | -79.8542 | 0.09583 | 0.072236 | 0.12001 | 119.757 | 2.978 | 0.27818 | -0.22486 | 1.9778 | 0.18235 | -0.1874 | -0.75334 | -3.24
pf_sweep | 3 | 2 | 8 | 2 | 120 | 1 | -0.799 | 60 | 119.9819 | 1.0002 | -79.8542 | 0.09583 | 0.072236 | 0.12001 | 119.757 | 1 | 0.093958 | -0.22486 | -0.000202 | -0.001872 | -0.1874 | -0.0202 | -1.9533
pf_sweep | 3 | 2 | 9 | 2 | 120 | 1 | -0.799 | 60 | 119.9819 | 1.0002 | -79.8542 | 0.09583 | 0.072236 | 0.12001 | 119.757 | 1.001 | 0.094007 | -0.22486 | 0.000798 | -0.001823 | -0.1874 | 0.0798 | -1.9022
pf_sweep | 3 | 2 | 10 | 2 | 120 | 1 | -0.799 | 60 | 119.9819 | 1.0002 | -79.8542 | 0.09583 | 0.072236 | 0.12001 | 119.757 | 1 | 0.093963 | -0.22486 | -0.000202 | -0.001867 | -0.1874 | -0.0202 | -1.9481
pf_sweep | 3 | 2 | 11 | 2 | 120 | 1 | -0.799 | 60 | 119.9819 | 1.0002 | -79.8542 | 0.09583 | 0.072236 | 0.12001 | 119.757 | 1 | 0.094114 | -0.22486 | -0.000202 | -0.001716 | -0.1874 | -0.0202 | -1.7905
pf_sweep | 3 | 2 | 12 | 2 | 120 | 1 | -0.799 | 60 | 119.9819 | 1.0002 | -79.8542 | 0.09583 | 0.072236 | 0.12001 | 119.757 | 1 | 0.09388 | -0.22486 | -0.000202 | -0.00195 | -0.1874 | -0.0202 | -2.0347
pf_sweep | 3 | 2 | 13 | 3 | 120 | 1 | -0.799 | 60 | 119.9903 | 0.9999 | -79.8905 | 0.095851 | 0.072162 | 0.11998 | 119.771 | 1 | 0.093979 | -0.21931 | 9.8e-05 | -0.001872 | -0.1828 | 0.0098 | -1.9535
pf_sweep | 3 | 2 | 14 | 3 | 120 | 1 | -0.799 | 60 | 119.9903 | 0.9999 | -79.8905 | 0.095851 | 0.072162 | 0.11998 | 119.771 | 0.999 | 0.094035 | -0.21931 | -0.000902 | -0.001816 | -0.1828 | -0.0902 | -1.8951
pf_sweep | 3 | 2 | 15 | 3 | 120 | 1 | -0.799 | 60 | 119.9903 | 0.9999 | -79.8905 | 0.095851 | 0.072162 | 0.11998 | 119.771 | 1 | 0.094096 | -0.21931 | 9.8e-05 | -0.001755 | -0.1828 | 0.0098 | -1.8315
pf_sweep | 3 | 2 | 16 | 3 | 120 | 1 | -0.799 | 60 | 119.9903 | 0.9999 | -79.8905 | 0.095851 | 0.072162 | 0.11998 | 119.771 | 0.999 | 0.094125 | -0.21931 | -0.000902 | -0.001726 | -0.1828 | -0.0902 | -1.8012
pf_sweep | 3 | 2 | 17 | 3 | 120 | 1 | -0.799 | 60 | 119.9903 | 0.9999 | -79.8905 | 0.095851 | 0.072162 | 0.11998 | 119.771 | 1 | 0.094036 | -0.21931 | 9.8e-05 | -0.001815 | -0.1828 | 0.0098 | -1.8941
pf_sweep | 3 | 2 | 18 | 3 | 120 | 1 | -0.799 | 60 | 119.9903 | 0.9999 | -79.8905 | 0.095851 | 0.072162 | 0.11998 | 119.771 | 1 | 0.094092 | -0.21931 | 9.8e-05 | -0.001759 | -0.1828 | 0.0098 | -1.8356
pf_sweep | 3 | 3 | 1 | 1 | 120 | 1 | -0.799 | 60 | 120.0051 | 0.99996 | -79.8893 | 0.095867 | 0.072177 | 0.12 | 119.714 | 2.968 | 0.27705 | -0.29111 | 1.968 | 0.18118 | -0.2426 | -1.0627 | -3.6701
pf_sweep | 3 | 3 | 2 | 1 | 120 | 1 | -0.799 | 60 | 120.0051 | 0.99996 | -79.8893 | 0.095867 | 0.072177 | 0.12 | 119.714 | 1.001 | 0.093932 | -0.29111 | 0.00104 | -0.001935 | -0.2426 | 0.104 | -2.0188
pf_sweep | 3 | 3 | 3 | 1 | 120 | 1 | -0.799 | 60 | 120.0051 | 0.99996 | -79.8893 | 0.095867 | 0.072177 | 0.12 | 119.714 | 1.001 | 0.093932 | -0.29111 | 0.00104 | -0.001935 | -0.2426 | 0.104 | -2.0188
pf_sweep | 3 | 3 | 4 | 1 | 120 | 1 | -0.799 | 60 | 120.0051 | 0.99996 | -79.8893 | 0.095867 | 0.072177 | 0.12 | 119.714 | 1 | 0.093818 | -0.29111 | 4e-05 | -0.002049 | -0.2426 | 0.004 | -2.1378
pf_sweep | 3 | 3 | 5 | 1 | 120 | 1 | -0.799 | 60 | 120.0051 | 0.99996 | -79.8893 | 0.095867 | 0.072177 | 0.12 | 119.714 | 1.001 | 0.093973 | -0.29111 | 0.00104 | -0.001894 | -0.2426 | 0.104 | -1.9761
pf_sweep | 3 | 3 | 6 | 1 | 120 | 1 | -0.799 | 60 | 120.0051 | 0.99996 | -79.8893 | 0.095867 | 0.072177 | 0.12 | 119.714 | 0.998 | 0.094006 | -0.29111 | -0.00196 | -0.001861 | -0.2426 | -0.196 | -1.9417
pf_sweep | 3 | 3 | 7 | 2 | 120 | 1 | -0.799 | 60 | 119.9876 | 1.0002 | -79.8498 | 0.095829 | 0.072247 | 0.12001 | 119.761 | 2.978 | 0.27812 | -0.22665 | 1.9778 | 0.18229 | -0.1889 | -0.75296 | -3.2581
pf_sweep | 3 | 3 | 8 | 2 | 120 | 1 | -0.799 | 60 | 119.9876 | 1.0002 | -79.8498 | 0.095829 | 0.072247 | 0.12001 | 119.761 | 1 | 0.09388 | -0.22665 | -0.000198 | -0.001949 | -0.1889 | -0.0198 | -2.0337
pf_sweep | 3 | 3 | 9 | 2 | 120 | 1 | -0.799 | 60 | 119.9876 | 1.0002 | -79.8498 | 0.095829 | 0.072247 | 0.12001 | 119.761 | 1 | 0.094007 | -0.22665 | -0.000198 | -0.001822 | -0.1889 | -0.0198 | -1.9012
pf_sweep | 3 | 3 | 10 | 2 | 120 | 1 | -0.799 | 60 | 119.9876 | 1.0002 | -79.8498 | 0.095829 | 0.072247 | 0.12001 | 119.761 | 1 | 0.093951 | -0.22665 | -0.000198 | -0.001878 | -0.1889 | -0.0198 | -1.9596
pf_sweep | 3 | 3 | 11 | 2 | 120 | 1 | -0.799 | 60 | 119.9876 | 1.0002 | -79.8498 | 0.095829 | 0.072247 | 0.12001 | 119.761 | 1 | 0.094084 | -0.22665 | -0.000198 | -0.001745 | -0.1889 | -0.0198 | -1.8208
pf_sweep | 3 | 3 | 12 | 2 | 120 | 1 | -0.799 | 60 | 119.9876 | 1.0002 | -79.8498 | 0.095829 | 0.072247 | 0.12001 | 119.761 | 1 | 0.093851 | -0.22665 | -0.000198 | -0.001978 | -0.1889 | -0.0198 | -2.064
pf_sweep | 3 | 3 | 13 | 3 | 120 | 1 | -0.799 | 60 | 119.99 | 0.99992 | -79.8921 | 0.095854 | 0.07216 | 0.11998 | 119.775 | 0.999 | 0.093976 | -0.21497 | -0.000916 | -0.001878 | -0.1792 | -0.0916 | -1.9597
pf_sweep | 3 | 3 | 14 | 3 | 120 | 1 | -0.799 | 60 | 119.99 | 0.99992 | -79.8921 | 0.095854 | 0.07216 | 0.11998 | 119.775 | 0.999 | 0.094045 | -0.21497 | -0.000916 | -0.001809 | -0.1792 | -0.0916 | -1.8877
pf_sweep | 3 | 3 | 15 | 3 | 120 | 1 | -0.799 | 60 | 119.99 | 0.99992 | -79.8921 | 0.095854 | 0.07216 | 0.11998 | 119.775 | 1 | 0.094092 | -0.21497 | 8.4e-05 | -0.001762 | -0.1792 | 0.0084 | -1.8387
pf_sweep | 3 | 3 | 16 | 3 | 120 | 1 | -0.799 | 60 | 119.99 | 0.99992 | -79.8921 | 0.095854 | 0.07216 | 0.11998 | 119.775 | 0.999 | 0.094131 | -0.21497 | -0.000916 | -0.001723 | -0.1792 | -0.0916 | -1.798
pf_sweep | 3 | 3 | 17 | 3 | 120 | 1 | -0.799 | 60 | 119.99 | 0.99992 | -79.8921 | 0.095854 | 0.07216 | 0.11998 | 119.775 | 1 | 0.094036 | -0.21497 | 8.4e-05 | -0.001818 | -0.1792 | 0.0084 | -1.8971
pf_sweep | 3 | 3 | 18 | 3 | 120 | 1 | -0.799 | 60 | 119.99 | 0.99992 | -79.8921 | 0.095854 | 0.07216 | 0.11998 | 119.775 | 0.999 | 0.094069 | -0.21497 | -0.000916 | -0.001785 | -0.1792 | -0.0916 | -1.8627
pf_sweep | 4 | 1 | 1 | 1 | 120 | 1 | -0.699 | 60 | 120.0131 | 0.99997 | -69.8875 | 0.083872 | 0.085836 | 0.12001 | 119.746 | 2.968 | 0.24008 | -0.26705 | 1.968 | 0.15621 | -0.2225 | -1.0637 | -4.5827
pf_sweep | 4 | 1 | 2 | 1 | 120 | 1 | -0.699 | 60 | 120.0131 | 0.99997 | -69.8875 | 0.083872 | 0.085836 | 0.12001 | 119.746 | 1.001 | 0.081515 | -0.26705 | 0.00103 | -0.002357 | -0.2225 | 0.103 | -2.8097
pf_sweep | 4 | 1 | 3 | 1 | 120 | 1 | -0.699 | 60 | 120.0131 | 0.99997 | -69.8875 | 0.083872 | 0.085836 | 0.12001 | 119.746 | 1.001 | 0.08148 | -0.26705 | 0.00103 | -0.002392 | -0.2225 | 0.103 | -2.8515
pf_sweep | 4 | 1 | 4 | 1 | 120 | 1 | -0.699 | 60 | 120.0131 | 0.99997 | -69.8875 | 0.083872 | 0.085836 | 0.12001 | 119.746 | 1 | 0.081433 | -0.26705 | 3e-05 | -0.002439 | -0.2225 | 0.003 | -2.9075
pf_sweep | 4 | 1 | 5 | 1 | 120 | 1 | -0.699 | 60 | 120.0131 | 0.99997 | -69.8875 | 0.083872 | 0.085836 | 0.12001 | 119.746 | 1.001 | 0.081575 | -0.26705 | 0.00103 | -0.002297 | -0.2225 | 0.103 | -2.7382
pf_sweep | 4 | 1 | 6 | 1 | 120 | 1 | -0.699 | 60 | 120.0131 | 0.99997 | -69.8875 | 0.083872 | 0.085836 | 0.12001 | 119.746 | 0.998 | 0.081656 | -0.26705 | -0.00197 | -0.002216 | -0.2225 | -0.197 | -2.6416
pf_sweep | 4 | 1 | 7 | 2 | 120 | 1 | -0.699 | 60 | 119.979 | 1.0002 | -69.8425 | 0.083815 | 0.085886 | 0.12001 | 119.749 | 2.979 | 0.24088 | -0.23003 | 1.9788 | 0.15706 | -0.1917 | -0.72167 | -4.2033
pf_sweep | 4 | 1 | 8 | 2 | 120 | 1 | -0.699 | 60 | 119.979 | 1.0002 | -69.8425 | 0.083815 | 0.085886 | 0.12001 | 119.749 | 1 | 0.08142 | -0.23003 | -0.000218 | -0.002395 | -0.1917 | -0.0218 | -2.857
pf_sweep | 4 | 1 | 9 | 2 | 120 | 1 | -0.699 | 60 | 119.979 | 1.0002 | -69.8425 | 0.083815 | 0.085886 | 0.12001 | 119.749 | 1.001 | 0.081593 | -0.23003 | 0.000782 | -0.002222 | -0.1917 | 0.0782 | -2.6506
pf_sweep | 4 | 1 | 10 | 2 | 120 | 1 | -0.699 | 60 | 119.979 | 1.0002 | -69.8425 | 0.083815 | 0.085886 | 0.12001 | 119.749 | 1 | 0.081534 | -0.23003 | -0.000218 | -0.002281 | -0.1917 | -0.0218 | -2.721
pf_sweep | 4 | 1 | 11 | 2 | 120 | 1 | -0.699 | 60 | 119.979 | 1.0002 | -69.8425 | 0.083815 | 0.085886 | 0.12001 | 119.749 | 1 | 0.081635 | -0.23003 | -0.000218 | -0.00218 | -0.1917 | -0.0218 | -2.6005
pf_sweep | 4 | 1 | 12 | 2 | 120 | 1 | -0.699 | 60 | 119.979 | 1.0002 | -69.8425 | 0.083815 | 0.085886 | 0.12001 | 119.749 | 1 | 0.081427 | -0.23003 | -0.000218 | -0.002388 | -0.1917 | -0.0218 | -2.8487
pf_sweep | 4 | 1 | 13 | 3 | 120 | 1 | -0.699 | 60 | 119.9871 | 0.9999 | -69.8906 | 0.083851 | 0.085808 | 0.11997 | 119.769 | 0.999 | 0.081535 | -0.21808 | -0.000899 | -0.002316 | -0.1818 | -0.0899 | -2.7623
pf_sweep | 4 | 1 | 14 | 3 | 120 | 1 | -0.699 | 60 | 119.9871 | 0.9999 | -69.8906 | 0.083851 | 0.085808 | 0.11997 | 119.769 | 0.999 | 0.081656 | -0.21808 | -0.000899 | -0.002195 | -0.1818 | -0.0899 | -2.618
pf_sweep | 4 | 1 | 15 | 3 | 120 | 1 | -0.699 | 60 | 119.9871 | 0.9999 | -69.8906 | 0.083851 | 0.085808 | 0.11997 | 119.769 | 1 | 0.081686 | -0.21808 | 0.000101 | -0.002165 | -0.1818 | 0.0101 | -2.5822
pf_sweep | 4 | 1 | 16 | 3 | 120 | 1 | -0.699 | 60 | 119.9871 | 0.9999 | -69.8906 | 0.083851 | 0.085808 | 0.11997 | 119.769 | 0.999 | 0.081713 | -0.21808 | -0.000899 | -0.002138 | -0.1818 | -0.0899 | -2.55
pf_sweep | 4 | 1 | 17 | 3 | 120 | 1 | -0.699 | 60 | 119.9871 | 0.9999 | -69.8906 | 0.083851 | 0.085808 | 0.11997 | 119.769 | 1 | 0.08164 | -0.21808 | 0.000101 | -0.002211 | -0.1818 | 0.0101 | -2.637
pf_sweep | 4 | 1 | 18 | 3 | 120 | 1 | -0.699 | 60 | 119.9871 | 0.9999 | -69.8906 | 0.083851 | 0.085808 | 0.11997 | 119.769 | 1 | 0.081687 | -0.21808 | 0.000101 | -0.002164 | -0.1818 | 0.0101 | -2.581
pf_sweep | 4 | 2 | 1 | 1 | 120 | 1 | -0.699 | 60 | 119.9943 | 0.99996 | -69.8878 | 0.083858 | 0.085821 | 0.11999 | 119.792 | 2.967 | 0.23987 | -0.20233 | 1.967 | 0.15601 | -0.1686 | -1.0961 | -4.6533
pf_sweep | 4 | 2 | 2 | 1 | 120 | 1 | -0.699 | 60 | 119.9943 | 0.99996 | -69.8878 | 0.083858 | 0.085821 | 0.11999 | 119.792 | 1.001 | 0.081565 | -0.20233 | 0.001039 | -0.002293 | -0.1686 | 0.1039 | -2.7346
pf_sweep | 4 | 2 | 3 | 1 | 120 | 1 | -0.699 | 60 | 119.9943 | 0.99996 | -69.8878 | 0.083858 | 0.085821 | 0.11999 | 119.792 | 1.001 | 0.081436 | -0.20233 | 0.001039 | -0.002422 | -0.1686 | 0.1039 | -2.8884
pf_sweep | 4 | 2 | 4 | 1 | 120 | 1 | -0.699 | 60 | 119.9943 | 0.99996 | -69.8878 | 0.083858 | 0.085821 | 0.11999 | 119.792 | 1 | 0.081353 | -0.20233 | 3.9e-05 | -0.002505 | -0.1686 | 0.0039 | -2.9874
pf_sweep | 4 | 2 | 5 | 1 | 120 | 1 | -0.699 | 60 | 119.9943 | 0.99996 | -69.8878 | 0.083858 | 0.085821 | 0.11999 | 119.792 | 1.001 | 0.081616 | -0.20233 | 0.001039 | -0.002242 | -0.1686 | 0.1039 | -2.6737
pf_sweep | 4 | 2 | 6 | 1 | 120 | 1 | -0.699 | 60 | 119.9943 | 0.99996 | -69.8878 | 0.083858 | 0.085821 | 0.11999 | 119.792 | 0.998 | 0.081573 | -0.20233 | -0.001961 | -0.002285 | -0.1686 | -0.1961 | -2.725
pf_sweep | 4 | 2 | 7 | 2 | 120 | 1 | -0.699 | 60 | 119.9829 | 1.0002 | -69.8464 | 0.083821 | 0.085883 | 0.12001 | 119.766 | 2.979 | 0.24089 | -0.21687 | 1.9788 | 0.15707 | -0.1808 | -0.72052 | -4.2059
pf_sweep | 4 | 2 | 8 | 2 | 120 | 1 | -0.699 | 60 | 119.9829 | 1.0002 | -69.8464 | 0.083821 | 0.085883 | 0.12001 | 119.766 | 1 | 0.081451 | -0.21687 | -0.000207 | -0.00237 | -0.1808 | -0.0207 | -2.8276
pf_sweep | 4 | 2 | 9 | 2 | 120 | 1 | -0.699 | 60 | 119.9829 | 1.0002 | -69.8464 | 0.083821 | 0.085883 | 0.12001 | 119.766 | 1.001 | 0.081561 | -0.21687 | 0.000793 | -0.00226 | -0.1808 | 0.0793 | -2.6963
pf_sweep | 4 | 2 | 10 | 2 | 120 | 1 | -0.699 | 60 | 119.9829 | 1.0002 | -69.8464 | 0.083821 | 0.085883 | 0.12001 | 119.766 | 1 | 0.081538 | -0.21687 | -0.000207 | -0.002283 | -0.1808 | -0.0207 | -2.7238
pf_sweep | 4 | 2 | 11 | 2 | 120 | 1 | -0.699 | 60 | 119.9829 | 1.0002 | -69.8464 | 0.083821 | 0.085883 | 0.12001 | 119.766 | 1 | 0.081656 | -0.21687 | -0.000207 | -0.002165 | -0.1808 | -0.0207 | -2.583
pf_sweep | 4 | 2 | 12 | 2 | 120 | 1 | -0.699 | 60 | 119.9829 | 1.0002 | -69.8464 | 0.083821 | 0.085883 | 0.12001 | 119.766 | 1 | 0.081405 | -0.21687 | -0.000207 | -0.002416 | -0.1808 | -0.0207 | -2.8824
pf_sweep | 4 | 2 | 13 | 3 | 120 | 1 | -0.699 | 60 | 119.988 | 0.99991 | -69.8877 | 0.08385 | 0.085813 | 0.11998 | 119.767 | 0.999 | 0.081526 | -0.22104 | -0.000913 | -0.002324 | -0.1842 | -0.0913 | -2.7712
pf_sweep | 4 | 2 | 14 | 3 | 120 | 1 | -0.699 | 60 | 119.988 | 0.99991 | -69.8877 | 0.08385 | 0.085813 | 0.11998 | 119.767 | 0.999 | 0.081665 | -0.22104 | -0.000913 | -0.002185 | -0.1842 | -0.0913 | -2.6054
pf_sweep | 4 | 2 | 15 | 3 | 120 | 1 | -0.699 | 60 | 119.988 | 0.99991 | -69.8877 | 0.08385 | 0.085813 | 0.11998 | 119.767 | 1.001 | 0.081709 | -0.22104 | 0.001087 | -0.002141 | -0.1842 | 0.1087 | -2.553
pf_sweep | 4 | 2 | 16 | 3 | 120 | 1 | -0.699 | 60 | 119.988 | 0.99991 | -69.8877 | 0.08385 | 0.085813 | 0.11998 | 119.767 | 0.999 | 0.081724 | -0.22104 | -0.000913 | -0.002126 | -0.1842 | -0.0913 | -2.5351
pf_sweep | 4 | 2 | 17 | 3 | 120 | 1 | -0.699 | 60 | 119.988 | 0.99991 | -69.8877 | 0.08385 | 0.085813 | 0.11998 | 119.767 | 1 | 0.081638 | -0.22104 | 8.7e-05 | -0.002212 | -0.1842 | 0.0087 | -2.6376
pf_sweep | 4 | 2 | 18 | 3 | 120 | 1 | -0.699 | 60 | 119.988 | 0.99991 | -69.8877 | 0.08385 | 0.085813 | 0.11998 | 119.767 | 1 | 0.081701 | -0.22104 | 8.7e-05 | -0.002149 | -0.1842 | 0.0087 | -2.5625
pf_sweep | 4 | 3 | 1 | 1 | 120 | 1 | -0.699 | 60 | 120.0043 | 0.99997 | -69.8888 | 0.083867 | 0.085828 | 0.12 | 119.692 | 2.968 | 0.24013 | -0.31229 | 1.968 | 0.15626 | -0.2602 | -1.0641 | -4.56
pf_sweep | 4 | 3 | 2 | 1 | 120 | 1 | -0.699 | 60 | 120.0043 | 0.99997 | -69.8888 | 0.083867 | 0.085828 | 0.12 | 119.692 | 1.001 | 0.081513 | -0.31229 | 0.001026 | -0.002354 | -0.2602 | 0.1026 | -2.8073
pf_sweep | 4 | 3 | 3 | 1 | 120 | 1 | -0.699 | 60 | 120.0043 | 0.99997 | -69.8888 | 0.083867 | 0.085828 | 0.12 | 119.692 | 1.001 | 0.08142 | -0.31229 | 0.001026 | -0.002447 | -0.2602 | 0.1026 | -2.9182
pf_sweep | 4 | 3 | 4 | 1 | 120 | 1 | -0.699 | 60 | 120.0043 | 0.99997 | -69.8888 | 0.083867 | 0.085828 | 0.12 | 119.692 | 1 | 0.081422 | -0.31229 | 2.6e-05 | -0.002445 | -0.2602 | 0.0026 | -2.9158
pf_sweep | 4 | 3 | 5 | 1 | 120 | 1 | -0.699 | 60 | 120.0043 | 0.99997 | -69.8888 | 0.083867 | 0.085828 | 0.12 | 119.692 | 1.001 | 0.081563 | -0.31229 | 0.001026 | -0.002304 | -0.2602 | 0.1026 | -2.7477
pf_sweep | 4 | 3 | 6 | 1 | 120 | 1 | -0.699 | 60 | 120.0043 | 0.99997 | -69.8888 | 0.083867 | 0.085828 | 0.12 | 119.692 | 0.998 | 0.081606 | -0.31229 | -0.001974 | -0.002261 | -0.2602 | -0.1974 | -2.6964
pf_sweep | 4 | 3 | 7 | 2 | 120 | 1 | -0.699 | 60 | 119.9867 | 1.0002 | -69.84 | 0.083815 | 0.085892 | 0.12001 | 119.773 | 2.978 | 0.24083 | -0.21367 | 1.9778 | 0.15702 | -0.1781 | -0.75226 | -4.2193
pf_sweep | 4 | 3 | 8 | 2 | 120 | 1 | -0.699 | 60 | 119.9867 | 1.0002 | -69.84 | 0.083815 | 0.085892 | 0.12001 | 119.773 | 1 | 0.081426 | -0.21367 | -0.000191 | -0.002389 | -0.1781 | -0.0191 | -2.85
pf_sweep | 4 | 3 | 9 | 2 | 120 | 1 | -0.699 | 60 | 119.9867 | 1.0002 | -69.84 | 0.083815 | 0.085892 | 0.12001 | 119.773 | 1 | 0.081576 | -0.21367 | -0.000191 | -0.002239 | -0.1781 | -0.0191 | -2.671
pf_sweep | 4 | 3 | 10 | 2 | 120 | 1 | -0.699 | 60 | 119.9867 | 1.0002 | -69.84 | 0.083815 | 0.085892 | 0.12001 | 119.773 | 1 | 0.081499 | -0.21367 | -0.000191 | -0.002316 | -0.1781 | -0.0191 | -2.7629
pf_sweep | 4 | 3 | 11 | 2 | 120 | 1 | -0.699 | 60 | 119.9867 | 1.0002 | -69.84 | 0.083815 | 0.085892 | 0.12001 | 119.773 | 1 | 0.081632 | -0.21367 | -0.000191 | -0.002183 | -0.1781 | -0.0191 | -2.6042
pf_sweep | 4 | 3 | 12 | 2 | 120 | 1 | -0.699 | 60 | 119.9867 | 1.0002 | -69.84 | 0.083815 | 0.085892 | 0.12001 | 119.773 | 1 | 0.081391 | -0.21367 | -0.000191 | -0.002424 | -0.1781 | -0.0191 | -2.8917
pf_sweep | 4 | 3 | 13 | 3 | 120 | 1 | -0.699 | 60 | 119.9921 | 0.9999 | -69.8904 | 0.083855 | 0.085812 | 0.11998 | 119.779 | 0.999 | 0.081529 | -0.2131 | -0.000904 | -0.002326 | -0.1776 | -0.0904 | -2.7738
pf_sweep | 4 | 3 | 14 | 3 | 120 | 1 | -0.699 | 60 | 119.9921 | 0.9999 | -69.8904 | 0.083855 | 0.085812 | 0.11998 | 119.779 | 0.999 | 0.081671 | -0.2131 | -0.000904 | -0.002184 | -0.1776 | -0.0904 | -2.6044
pf_sweep | 4 | 3 | 15 | 3 | 120 | 1 | -0.699 | 60 | 119.9921 | 0.9999 | -69.8904 | 0.083855 | 0.085812 | 0.11998 | 119.779 | 1.001 | 0.081725 | -0.2131 | 0.001096 | -0.00213 | -0.1776 | 0.1096 | -2.54
pf_sweep | 4 | 3 | 16 | 3 | 120 | 1 | -0.699 | 60 | 119.9921 | 0.9999 | -69.8904 | 0.083855 | 0.085812 | 0.11998 | 119.779 | 0.999 | 0.081718 | -0.2131 | -0.000904 | -0.002137 | -0.1776 | -0.0904 | -2.5484
pf_sweep | 4 | 3 | 17 | 3 | 120 | 1 | -0.699 | 60 | 119.9921 | 0.9999 | -69.8904 | 0.083855 | 0.085812 | 0.11998 | 119.779 | 1 | 0.081644 | -0.2131 | 9.6e-05 | -0.002211 | -0.1776 | 0.0096 | -2.6366
pf_sweep | 4 | 3 | 18 | 3 | 120 | 1 | -0.699 | 60 | 119.9921 | 0.9999 | -69.8904 | 0.083855 | 0.085812 | 0.11998 | 119.779 | 1 | 0.081717 | -0.2131 | 9.6e-05 | -0.002138 | -0.1776 | 0.0096 | -2.5496
pf_sweep | 5 | 1 | 1 | 1 | 120 | 1 | -0.599 | 60 | 120.0497 | 0.99996 | -59.8789 | 0.071881 | 0.096144 | 0.12004 | 119.716 | 2.968 | 0.20317 | -0.33372 | 1.968 | 0.13129 | -0.278 | -1.0624 | -5.7832
pf_sweep | 5 | 1 | 2 | 1 | 120 | 1 | -0.599 | 60 | 120.0497 | 0.99996 | -59.8789 | 0.071881 | 0.096144 | 0.12004 | 119.716 | 1.001 | 0.069134 | -0.33372 | 0.001043 | -0.002747 | -0.278 | 0.1043 | -3.8221
pf_sweep | 5 | 1 | 3 | 1 | 120 | 1 | -0.599 | 60 | 120.0497 | 0.99996 | -59.8789 | 0.071881 | 0.096144 | 0.12004 | 119.716 | 1.001 | 0.069107 | -0.33372 | 0.001043 | -0.002774 | -0.278 | 0.1043 | -3.8596
pf_sweep | 5 | 1 | 4 | 1 | 120 | 1 | -0.599 | 60 | 120.0497 | 0.99996 | -59.8789 | 0.071881 | 0.096144 | 0.12004 | 119.716 | 1 | 0.06903 | -0.33372 | 4.3e-05 | -0.002851 | -0.278 | 0.0043 | -3.9668
pf_sweep | 5 | 1 | 5 | 1 | 120 | 1 | -0.599 | 60 | 120.0497 | 0.99996 | -59.8789 | 0.071881 | 0.096144 | 0.12004 | 119.716 | 1.001 | 0.069185 | -0.33372 | 0.001043 | -0.002696 | -0.278 | 0.1043 | -3.7511
pf_sweep | 5 | 1 | 6 | 1 | 120 | 1 | -0.599 | 60 | 120.0497 | 0.99996 | -59.8789 | 0.071881 | 0.096144 | 0.12004 | 119.716 | 0.998 | 0.069345 | -0.33372 | -0.001957 | -0.002536 | -0.278 | -0.1957 | -3.5285
pf_sweep | 5 | 1 | 7 | 2 | 120 | 1 | -0.599 | 60 | 119.9795 | 1.0002 | -59.8334 | 0.071804 | 0.096154 | 0.12001 | 119.775 | 2.979 | 0.20398 | -0.20445 | 1.9788 | 0.13217 | -0.1704 | -0.72193 | -5.3084
pf_sweep | 5 | 1 | 8 | 2 | 120 | 1 | -0.599 | 60 | 119.9795 | 1.0002 | -59.8334 | 0.071804 | 0.096154 | 0.12001 | 119.775 | 1 | 0.06905 | -0.20445 | -0.000221 | -0.002754 | -0.1704 | -0.0221 | -3.835
pf_sweep | 5 | 1 | 9 | 2 | 120 | 1 | -0.599 | 60 | 119.9795 | 1.0002 | -59.8334 | 0.071804 | 0.096154 | 0.12001 | 119.775 | 1 | 0.069217 | -0.20445 | -0.000221 | -0.002587 | -0.1704 | -0.0221 | -3.6024
pf_sweep | 5 | 1 | 10 | 2 | 120 | 1 | -0.599 | 60 | 119.9795 | 1.0002 | -59.8334 | 0.071804 | 0.096154 | 0.12001 | 119.775 | 1 | 0.069184 | -0.20445 | -0.000221 | -0.00262 | -0.1704 | -0.0221 | -3.6483
pf_sweep | 5 | 1 | 11 | 2 | 120 | 1 | -0.599 | 60 | 119.9795 | 1.0002 | -59.8334 | 0.071804 | 0.096154 | 0.12001 | 119.775 | 1 | 0.069305 | -0.20445 | -0.000221 | -0.002499 | -0.1704 | -0.0221 | -3.4798
pf_sweep | 5 | 1 | 12 | 2 | 120 | 1 | -0.599 | 60 | 119.9795 | 1.0002 | -59.8334 | 0.071804 | 0.096154 | 0.12001 | 119.775 | 1 | 0.069042 | -0.20445 | -0.000221 | -0.002762 | -0.1704 | -0.0221 | -3.8461
pf_sweep | 5 | 1 | 13 | 3 | 120 | 1 | -0.599 | 60 | 119.9916 | 0.9999 | -59.8905 | 0.071857 | 0.096082 | 0.11998 | 119.764 | 1 | 0.069224 | -0.2276 | 9.5e-05 | -0.002633 | -0.1897 | 0.0095 | -3.6639
pf_sweep | 5 | 1 | 14 | 3 | 120 | 1 | -0.599 | 60 | 119.9916 | 0.9999 | -59.8905 | 0.071857 | 0.096082 | 0.11998 | 119.764 | 0.999 | 0.069403 | -0.2276 | -0.000905 | -0.002454 | -0.1897 | -0.0905 | -3.4148
pf_sweep | 5 | 1 | 15 | 3 | 120 | 1 | -0.599 | 60 | 119.9916 | 0.9999 | -59.8905 | 0.071857 | 0.096082 | 0.11998 | 119.764 | 1 | 0.069382 | -0.2276 | 9.5e-05 | -0.002475 | -0.1897 | 0.0095 | -3.444
pf_sweep | 5 | 1 | 16 | 3 | 120 | 1 | -0.599 | 60 | 119.9916 | 0.9999 | -59.8905 | 0.071857 | 0.096082 | 0.11998 | 119.764 | 0.999 | 0.069437 | -0.2276 | -0.000905 | -0.00242 | -0.1897 | -0.0905 | -3.3674
pf_sweep | 5 | 1 | 17 | 3 | 120 | 1 | -0.599 | 60 | 119.9916 | 0.9999 | -59.8905 | 0.071857 | 0.096082 | 0.11998 | 119.764 | 1 | 0.06936 | -0.2276 | 9.5e-05 | -0.002497 | -0.1897 | 0.0095 | -3.4746
pf_sweep | 5 | 1 | 18 | 3 | 120 | 1 | -0.599 | 60 | 119.9916 | 0.9999 | -59.8905 | 0.071857 | 0.096082 | 0.11998 | 119.764 | 1 | 0.069409 | -0.2276 | 9.5e-05 | -0.002448 | -0.1897 | 0.0095 | -3.4064
pf_sweep | 5 | 2 | 1 | 1 | 120 | 1 | -0.599 | 60 | 120.012 | 0.99996 | -59.8865 | 0.071868 | 0.096108 | 0.12001 | 119.741 | 2.968 | 0.20331 | -0.27099 | 1.968 | 0.13145 | -0.2258 | -1.0631 | -5.7013
pf_sweep | 5 | 2 | 2 | 1 | 120 | 1 | -0.599 | 60 | 120.012 | 0.99996 | -59.8865 | 0.071868 | 0.096108 | 0.12001 | 119.741 | 1.001 | 0.069119 | -0.27099 | 0.001036 | -0.002749 | -0.2258 | 0.1036 | -3.8256
pf_sweep | 5 | 2 | 3 | 1 | 120 | 1 | -0.599 | 60 | 120.012 | 0.99996 | -59.8865 | 0.071868 | 0.096108 | 0.12001 | 119.741 | 1.001 | 0.069146 | -0.27099 | 0.001036 | -0.002722 | -0.2258 | 0.1036 | -3.788
pf_sweep | 5 | 2 | 4 | 1 | 120 | 1 | -0.599 | 60 | 120.012 | 0.99996 | -59.8865 | 0.071868 | 0.096108 | 0.12001 | 119.741 | 1 | 0.069064 | -0.27099 | 3.6e-05 | -0.002804 | -0.2258 | 0.0036 | -3.9021
pf_sweep | 5 | 2 | 5 | 1 | 120 | 1 | -0.599 | 60 | 120.012 | 0.99996 | -59.8865 | 0.071868 | 0.096108 | 0.12001 | 119.741 | 1.001 | 0.069189 | -0.27099 | 0.001036 | -0.002679 | -0.2258 | 0.1036 | -3.7282
pf_sweep | 5 | 2 | 6 | 1 | 120 | 1 | -0.599 | 60 | 120.012 | 0.99996 | -59.8865 | 0.071868 | 0.096108 | 0.12001 | 119.741 | 0.998 | 0.069376 | -0.27099 | -0.001964 | -0.002492 | -0.2258 | -0.1964 | -3.468
pf_sweep | 5 | 2 | 7 | 2 | 120 | 1 | -0.599 | 60 | 119.984 | 1.0002 | -59.8353 | 0.071807 | 0.096155 | 0.12001 | 119.765 | 2.979 | 0.20394 | -0.21905 | 1.9788 | 0.13213 | -0.1826 | -0.72016 | -5.3323
pf_sweep | 5 | 2 | 8 | 2 | 120 | 1 | -0.599 | 60 | 119.984 | 1.0002 | -59.8353 | 0.071807 | 0.096155 | 0.12001 | 119.765 | 1 | 0.069089 | -0.21905 | -0.000203 | -0.002718 | -0.1826 | -0.0203 | -3.7856
pf_sweep | 5 | 2 | 9 | 2 | 120 | 1 | -0.599 | 60 | 119.984 | 1.0002 | -59.8353 | 0.071807 | 0.096155 | 0.12001 | 119.765 | 1.001 | 0.069231 | -0.21905 | 0.000797 | -0.002576 | -0.1826 | 0.0797 | -3.5878
pf_sweep | 5 | 2 | 10 | 2 | 120 | 1 | -0.599 | 60 | 119.984 | 1.0002 | -59.8353 | 0.071807 | 0.096155 | 0.12001 | 119.765 | 1 | 0.069168 | -0.21905 | -0.000203 | -0.002639 | -0.1826 | -0.0203 | -3.6756
pf_sweep | 5 | 2 | 11 | 2 | 120 | 1 | -0.599 | 60 | 119.984 | 1.0002 | -59.8353 | 0.071807 | 0.096155 | 0.12001 | 119.765 | 1 | 0.069334 | -0.21905 | -0.000203 | -0.002473 | -0.1826 | -0.0203 | -3.4444
pf_sweep | 5 | 2 | 12 | 2 | 120 | 1 | -0.599 | 60 | 119.984 | 1.0002 | -59.8353 | 0.071807 | 0.096155 | 0.12001 | 119.765 | 1 | 0.069066 | -0.21905 | -0.000203 | -0.002741 | -0.1826 | -0.0203 | -3.8176
pf_sweep | 5 | 2 | 13 | 3 | 120 | 1 | -0.599 | 60 | 119.9876 | 0.99992 | -59.8879 | 0.071852 | 0.096083 | 0.11998 | 119.768 | 0.999 | 0.069217 | -0.21956 | -0.00092 | -0.002635 | -0.183 | -0.092 | -3.6676
pf_sweep | 5 | 2 | 14 | 3 | 120 | 1 | -0.599 | 60 | 119.9876 | 0.99992 | -59.8879 | 0.071852 | 0.096083 | 0.11998 | 119.768 | 0.999 | 0.069398 | -0.21956 | -0.00092 | -0.002454 | -0.183 | -0.092 | -3.4157
pf_sweep | 5 | 2 | 15 | 3 | 120 | 1 | -0.599 | 60 | 119.9876 | 0.99992 | -59.8879 | 0.071852 | 0.096083 | 0.11998 | 119.768 | 1 | 0.069347 | -0.21956 | 8e-05 | -0.002505 | -0.183 | 0.008 | -3.4867
pf_sweep | 5 | 2 | 16 | 3 | 120 | 1 | -0.599 | 60 | 119.9876 | 0.99992 | -59.8879 | 0.071852 | 0.096083 | 0.11998 | 119.768 | 0.999 | 0.06943 | -0.21956 | -0.00092 | -0.002422 | -0.183 | -0.092 | -3.3712
pf_sweep | 5 | 2 | 17 | 3 | 120 | 1 | -0.599 | 60 | 119.9876 | 0.99992 | -59.8879 | 0.071852 | 0.096083 | 0.11998 | 119.768 | 1 | 0.069356 | -0.21956 | 8e-05 | -0.002496 | -0.183 | 0.008 | -3.4742
pf_sweep | 5 | 2 | 18 | 3 | 120 | 1 | -0.599 | 60 | 119.9876 | 0.99992 | -59.8879 | 0.071852 | 0.096083 | 0.11998 | 119.768 | 1 | 0.069403 | -0.21956 | 8e-05 | -0.002449 | -0.183 | 0.008 | -3.4087
pf_sweep | 5 | 3 | 1 | 1 | 120 | 1 | -0.599 | 60 | 120.0074 | 0.99994 | -59.8864 | 0.071864 | 0.096102 | 0.12 | 119.697 | 2.968 | 0.20338 | -0.31039 | 1.9681 | 0.13151 | -0.2586 | -1.0612 | -5.6669
pf_sweep | 5 | 3 | 2 | 1 | 120 | 1 | -0.599 | 60 | 120.0074 | 0.99994 | -59.8864 | 0.071864 | 0.096102 | 0.12 | 119.697 | 1.001 | 0.069107 | -0.31039 | 0.001055 | -0.002757 | -0.2586 | 0.1055 | -3.8366
pf_sweep | 5 | 3 | 3 | 1 | 120 | 1 | -0.599 | 60 | 120.0074 | 0.99994 | -59.8864 | 0.071864 | 0.096102 | 0.12 | 119.697 | 1.001 | 0.069203 | -0.31039 | 0.001055 | -0.002661 | -0.2586 | 0.1055 | -3.703
pf_sweep | 5 | 3 | 4 | 1 | 120 | 1 | -0.599 | 60 | 120.0074 | 0.99994 | -59.8864 | 0.071864 | 0.096102 | 0.12 | 119.697 | 1 | 0.069097 | -0.31039 | 5.5e-05 | -0.002767 | -0.2586 | 0.0055 | -3.8505
pf_sweep | 5 | 3 | 5 | 1 | 120 | 1 | -0.599 | 60 | 120.0074 | 0.99994 | -59.8864 | 0.071864 | 0.096102 | 0.12 | 119.697 | 1.001 | 0.069157 | -0.31039 | 0.001055 | -0.002707 | -0.2586 | 0.1055 | -3.767
pf_sweep | 5 | 3 | 6 | 1 | 120 | 1 | -0.599 | 60 | 120.0074 | 0.99994 | -59.8864 | 0.071864 | 0.096102 | 0.12 | 119.697 | 0.998 | 0.069452 | -0.31039 | -0.001945 | -0.002412 | -0.2586 | -0.1945 | -3.3565
pf_sweep | 5 | 3 | 7 | 2 | 120 | 1 | -0.599 | 60 | 119.9806 | 1.0002 | -59.8369 | 0.071809 | 0.096152 | 0.12001 | 119.764 | 2.978 | 0.20383 | -0.21658 | 1.9778 | 0.13202 | -0.1805 | -0.75568 | -5.3821
pf_sweep | 5 | 3 | 8 | 2 | 120 | 1 | -0.599 | 60 | 119.9806 | 1.0002 | -59.8369 | 0.071809 | 0.096152 | 0.12001 | 119.764 | 1 | 0.069075 | -0.21658 | -0.000225 | -0.002734 | -0.1805 | -0.0225 | -3.8071
pf_sweep | 5 | 3 | 9 | 2 | 120 | 1 | -0.599 | 60 | 119.9806 | 1.0002 | -59.8369 | 0.071809 | 0.096152 | 0.12001 | 119.764 | 1.001 | 0.06922 | -0.21658 | 0.000775 | -0.002589 | -0.1805 | 0.0775 | -3.6051
pf_sweep | 5 | 3 | 10 | 2 | 120 | 1 | -0.599 | 60 | 119.9806 | 1.0002 | -59.8369 | 0.071809 | 0.096152 | 0.12001 | 119.764 | 1 | 0.069147 | -0.21658 | -0.000225 | -0.002662 | -0.1805 | -0.0225 | -3.7068
pf_sweep | 5 | 3 | 11 | 2 | 120 | 1 | -0.599 | 60 | 119.9806 | 1.0002 | -59.8369 | 0.071809 | 0.096152 | 0.12001 | 119.764 | 1 | 0.069339 | -0.21658 | -0.000225 | -0.00247 | -0.1805 | -0.0225 | -3.4394
pf_sweep | 5 | 3 | 12 | 2 | 120 | 1 | -0.599 | 60 | 119.9806 | 1.0002 | -59.8369 | 0.071809 | 0.096152 | 0.12001 | 119.764 | 1 | 0.069035 | -0.21658 | -0.000225 | -0.002774 | -0.1805 | -0.0225 | -3.8628
pf_sweep | 5 | 3 | 13 | 3 | 120 | 1 | -0.599 | 60 | 119.9921 | 0.99992 | -59.8879 | 0.071855 | 0.096087 | 0.11998 | 119.778 | 0.999 | 0.069242 | -0.21406 | -0.000924 | -0.002613 | -0.1784 | -0.0924 | -3.6369
pf_sweep | 5 | 3 | 14 | 3 | 120 | 1 | -0.599 | 60 | 119.9921 | 0.99992 | -59.8879 | 0.071855 | 0.096087 | 0.11998 | 119.778 | 0.999 | 0.069401 | -0.21406 | -0.000924 | -0.002454 | -0.1784 | -0.0924 | -3.4156
pf_sweep | 5 | 3 | 15 | 3 | 120 | 1 | -0.599 | 60 | 119.9921 | 0.99992 | -59.8879 | 0.071855 | 0.096087 | 0.11998 | 119.778 | 1 | 0.069342 | -0.21406 | 7.6e-05 | -0.002513 | -0.1784 | 0.0076 | -3.4977
pf_sweep | 5 | 3 | 16 | 3 | 120 | 1 | -0.599 | 60 | 119.9921 | 0.99992 | -59.8879 | 0.071855 | 0.096087 | 0.11998 | 119.778 | 0.999 | 0.069445 | -0.21406 | -0.000924 | -0.00241 | -0.1784 | -0.0924 | -3.3544
pf_sweep | 5 | 3 | 17 | 3 | 120 | 1 | -0.599 | 60 | 119.9921 | 0.99992 | -59.8879 | 0.071855 | 0.096087 | 0.11998 | 119.778 | 1 | 0.069359 | -0.21406 | 7.6e-05 | -0.002496 | -0.1784 | 0.0076 | -3.474
pf_sweep | 5 | 3 | 18 | 3 | 120 | 1 | -0.599 | 60 | 119.9921 | 0.99992 | -59.8879 | 0.071855 | 0.096087 | 0.11998 | 119.778 | 1 | 0.069379 | -0.21406 | 7.6e-05 | -0.002476 | -0.1784 | 0.0076 | -3.4462
pf_sweep | 6 | 1 | 1 | 1 | 120 | 1 | -0.5 | 60 | 120.0187 | 0.99996 | -49.983 | 0.059986 | 0.10395 | 0.12001 | 119.764 | 2.968 | 0.16713 | -0.25469 | 1.968 | 0.10715 | -0.2122 | -1.0624 | -7.1272
pf_sweep | 6 | 1 | 2 | 1 | 120 | 1 | -0.5 | 60 | 120.0187 | 0.99996 | -49.983 | 0.059986 | 0.10395 | 0.12001 | 119.764 | 1.001 | 0.056984 | -0.25469 | 0.001044 | -0.003002 | -0.2122 | 0.1044 | -5.0051
pf_sweep | 6 | 1 | 3 | 1 | 120 | 1 | -0.5 | 60 | 120.0187 | 0.99996 | -49.983 | 0.059986 | 0.10395 | 0.12001 | 119.764 | 1.001 | 0.056972 | -0.25469 | 0.001044 | -0.003014 | -0.2122 | 0.1044 | -5.0251
pf_sweep | 6 | 1 | 4 | 1 | 120 | 1 | -0.5 | 60 | 120.0187 | 0.99996 | -49.983 | 0.059986 | 0.10395 | 0.12001 | 119.764 | 1 | 0.056908 | -0.25469 | 4.4e-05 | -0.003078 | -0.2122 | 0.0044 | -5.1318
pf_sweep | 6 | 1 | 5 | 1 | 120 | 1 | -0.5 | 60 | 120.0187 | 0.99996 | -49.983 | 0.059986 | 0.10395 | 0.12001 | 119.764 | 1.001 | 0.057042 | -0.25469 | 0.001044 | -0.002944 | -0.2122 | 0.1044 | -4.9084
pf_sweep | 6 | 1 | 6 | 1 | 120 | 1 | -0.5 | 60 | 120.0187 | 0.99996 | -49.983 | 0.059986 | 0.10395 | 0.12001 | 119.764 | 0.998 | 0.057332 | -0.25469 | -0.001956 | -0.002654 | -0.2122 | -0.1957 | -4.425
pf_sweep | 6 | 1 | 7 | 2 | 120 | 1 | -0.5 | 60 | 119.9799 | 1.0002 | -49.928 | 0.059917 | 0.10398 | 0.12001 | 119.766 | 2.978 | 0.16752 | -0.21387 | 1.9778 | 0.1076 | -0.1783 | -0.75541 | -6.8047
pf_sweep | 6 | 1 | 8 | 2 | 120 | 1 | -0.5 | 60 | 119.9799 | 1.0002 | -49.928 | 0.059917 | 0.10398 | 0.12001 | 119.766 | 1 | 0.056928 | -0.21387 | -0.000222 | -0.002989 | -0.1783 | -0.0222 | -4.9883
pf_sweep | 6 | 1 | 9 | 2 | 120 | 1 | -0.5 | 60 | 119.9799 | 1.0002 | -49.928 | 0.059917 | 0.10398 | 0.12001 | 119.766 | 1.001 | 0.057082 | -0.21387 | 0.000778 | -0.002835 | -0.1783 | 0.0777 | -4.7313
pf_sweep | 6 | 1 | 10 | 2 | 120 | 1 | -0.5 | 60 | 119.9799 | 1.0002 | -49.928 | 0.059917 | 0.10398 | 0.12001 | 119.766 | 1 | 0.056982 | -0.21387 | -0.000222 | -0.002935 | -0.1783 | -0.0222 | -4.8982
pf_sweep | 6 | 1 | 11 | 2 | 120 | 1 | -0.5 | 60 | 119.9799 | 1.0002 | -49.928 | 0.059917 | 0.10398 | 0.12001 | 119.766 | 1 | 0.057213 | -0.21387 | -0.000222 | -0.002704 | -0.1783 | -0.0222 | -4.5126
pf_sweep | 6 | 1 | 12 | 2 | 120 | 1 | -0.5 | 60 | 119.9799 | 1.0002 | -49.928 | 0.059917 | 0.10398 | 0.12001 | 119.766 | 1 | 0.056898 | -0.21387 | -0.000222 | -0.003019 | -0.1783 | -0.0222 | -5.0384
pf_sweep | 6 | 1 | 13 | 3 | 120 | 1 | -0.5 | 60 | 119.99 | 0.99991 | -49.9885 | 0.059976 | 0.10391 | 0.11998 | 119.762 | 1 | 0.057139 | -0.22796 | 8.7e-05 | -0.002837 | -0.19 | 0.0087 | -4.7301
pf_sweep | 6 | 1 | 14 | 3 | 120 | 1 | -0.5 | 60 | 119.99 | 0.99991 | -49.9885 | 0.059976 | 0.10391 | 0.11998 | 119.762 | 0.999 | 0.057291 | -0.22796 | -0.000913 | -0.002685 | -0.19 | -0.0913 | -4.4767
pf_sweep | 6 | 1 | 15 | 3 | 120 | 1 | -0.5 | 60 | 119.99 | 0.99991 | -49.9885 | 0.059976 | 0.10391 | 0.11998 | 119.762 | 1.001 | 0.057224 | -0.22796 | 0.001087 | -0.002752 | -0.19 | 0.1087 | -4.5884
pf_sweep | 6 | 1 | 16 | 3 | 120 | 1 | -0.5 | 60 | 119.99 | 0.99991 | -49.9885 | 0.059976 | 0.10391 | 0.11998 | 119.762 | 0.999 | 0.057349 | -0.22796 | -0.000913 | -0.002627 | -0.19 | -0.0913 | -4.38
pf_sweep | 6 | 1 | 17 | 3 | 120 | 1 | -0.5 | 60 | 119.99 | 0.99991 | -49.9885 | 0.059976 | 0.10391 | 0.11998 | 119.762 | 1 | 0.057215 | -0.22796 | 8.7e-05 | -0.002761 | -0.19 | 0.0087 | -4.6034
pf_sweep | 6 | 1 | 18 | 3 | 120 | 1 | -0.5 | 60 | 119.99 | 0.99991 | -49.9885 | 0.059976 | 0.10391 | 0.11998 | 119.762 | 1 | 0.057258 | -0.22796 | 8.7e-05 | -0.002718 | -0.19 | 0.0087 | -4.5317
pf_sweep | 6 | 2 | 1 | 1 | 120 | 1 | -0.5 | 60 | 119.9755 | 0.99997 | -49.9885 | 0.059972 | 0.10391 | 0.11997 | 119.754 | 2.968 | 0.16709 | -0.22152 | 1.968 | 0.10712 | -0.1846 | -1.0638 | -7.127
pf_sweep | 6 | 2 | 2 | 1 | 120 | 1 | -0.5 | 60 | 119.9755 | 0.99997 | -49.9885 | 0.059972 | 0.10391 | 0.11997 | 119.754 | 1.001 | 0.056996 | -0.22152 | 0.001029 | -0.002976 | -0.1846 | 0.1029 | -4.9626
pf_sweep | 6 | 2 | 3 | 1 | 120 | 1 | -0.5 | 60 | 119.9755 | 0.99997 | -49.9885 | 0.059972 | 0.10391 | 0.11997 | 119.754 | 1.001 | 0.057006 | -0.22152 | 0.001029 | -0.002966 | -0.1846 | 0.1029 | -4.946
pf_sweep | 6 | 2 | 4 | 1 | 120 | 1 | -0.5 | 60 | 119.9755 | 0.99997 | -49.9885 | 0.059972 | 0.10391 | 0.11997 | 119.754 | 1 | 0.056902 | -0.22152 | 2.9e-05 | -0.00307 | -0.1846 | 0.0029 | -5.1194
pf_sweep | 6 | 2 | 5 | 1 | 120 | 1 | -0.5 | 60 | 119.9755 | 0.99997 | -49.9885 | 0.059972 | 0.10391 | 0.11997 | 119.754 | 1.001 | 0.057045 | -0.22152 | 0.001029 | -0.002927 | -0.1846 | 0.1029 | -4.8809
pf_sweep | 6 | 2 | 6 | 1 | 120 | 1 | -0.5 | 60 | 119.9755 | 0.99997 | -49.9885 | 0.059972 | 0.10391 | 0.11997 | 119.754 | 0.998 | 0.057318 | -0.22152 | -0.001971 | -0.002654 | -0.1846 | -0.1971 | -4.4257
pf_sweep | 6 | 2 | 7 | 2 | 120 | 1 | -0.5 | 60 | 119.9822 | 1.0002 | -49.9272 | 0.059918 | 0.10398 | 0.12001 | 119.76 | 2.979 | 0.16755 | -0.22219 | 1.9788 | 0.10763 | -0.1852 | -0.72333 | -6.7907
pf_sweep | 6 | 2 | 8 | 2 | 120 | 1 | -0.5 | 60 | 119.9822 | 1.0002 | -49.9272 | 0.059918 | 0.10398 | 0.12001 | 119.76 | 1 | 0.056939 | -0.22219 | -0.000235 | -0.002979 | -0.1852 | -0.0235 | -4.9715
pf_sweep | 6 | 2 | 9 | 2 | 120 | 1 | -0.5 | 60 | 119.9822 | 1.0002 | -49.9272 | 0.059918 | 0.10398 | 0.12001 | 119.76 | 1.001 | 0.057069 | -0.22219 | 0.000765 | -0.002849 | -0.1852 | 0.0765 | -4.7546
pf_sweep | 6 | 2 | 10 | 2 | 120 | 1 | -0.5 | 60 | 119.9822 | 1.0002 | -49.9272 | 0.059918 | 0.10398 | 0.12001 | 119.76 | 1 | 0.056993 | -0.22219 | -0.000235 | -0.002925 | -0.1852 | -0.0235 | -4.8814
pf_sweep | 6 | 2 | 11 | 2 | 120 | 1 | -0.5 | 60 | 119.9822 | 1.0002 | -49.9272 | 0.059918 | 0.10398 | 0.12001 | 119.76 | 1 | 0.057214 | -0.22219 | -0.000235 | -0.002704 | -0.1852 | -0.0235 | -4.5126
pf_sweep | 6 | 2 | 12 | 2 | 120 | 1 | -0.5 | 60 | 119.9822 | 1.0002 | -49.9272 | 0.059918 | 0.10398 | 0.12001 | 119.76 | 1 | 0.056905 | -0.22219 | -0.000235 | -0.003013 | -0.1852 | -0.0235 | -5.0283
pf_sweep | 6 | 2 | 13 | 3 | 120 | 1 | -0.5 | 60 | 119.9872 | 0.99991 | -49.9852 | 0.059971 | 0.10391 | 0.11998 | 119.766 | 1 | 0.057135 | -0.22116 | 8.8e-05 | -0.002836 | -0.1843 | 0.0088 | -4.7283
pf_sweep | 6 | 2 | 14 | 3 | 120 | 1 | -0.5 | 60 | 119.9872 | 0.99991 | -49.9852 | 0.059971 | 0.10391 | 0.11998 | 119.766 | 0.999 | 0.057269 | -0.22116 | -0.000912 | -0.002702 | -0.1843 | -0.0912 | -4.5048
pf_sweep | 6 | 2 | 15 | 3 | 120 | 1 | -0.5 | 60 | 119.9872 | 0.99991 | -49.9852 | 0.059971 | 0.10391 | 0.11998 | 119.766 | 1 | 0.057217 | -0.22116 | 8.8e-05 | -0.002754 | -0.1843 | 0.0088 | -4.5915
pf_sweep | 6 | 2 | 16 | 3 | 120 | 1 | -0.5 | 60 | 119.9872 | 0.99991 | -49.9852 | 0.059971 | 0.10391 | 0.11998 | 119.766 | 0.999 | 0.057354 | -0.22116 | -0.000912 | -0.002617 | -0.1843 | -0.0912 | -4.3631
pf_sweep | 6 | 2 | 17 | 3 | 120 | 1 | -0.5 | 60 | 119.9872 | 0.99991 | -49.9852 | 0.059971 | 0.10391 | 0.11998 | 119.766 | 1 | 0.057202 | -0.22116 | 8.8e-05 | -0.002769 | -0.1843 | 0.0088 | -4.6165
pf_sweep | 6 | 2 | 18 | 3 | 120 | 1 | -0.5 | 60 | 119.9872 | 0.99991 | -49.9852 | 0.059971 | 0.10391 | 0.11998 | 119.766 | 0.999 | 0.05725 | -0.22116 | -0.000912 | -0.002721 | -0.1843 | -0.0912 | -4.5365
pf_sweep | 6 | 3 | 1 | 1 | 120 | 1 | -0.5 | 60 | 119.9961 | 0.99997 | -49.9859 | 0.059979 | 0.10393 | 0.11999 | 119.689 | 2.968 | 0.16709 | -0.30707 | 1.968 | 0.10711 | -0.2559 | -1.0633 | -7.1416
pf_sweep | 6 | 3 | 2 | 1 | 120 | 1 | -0.5 | 60 | 119.9961 | 0.99997 | -49.9859 | 0.059979 | 0.10393 | 0.11999 | 119.689 | 1.001 | 0.056995 | -0.30707 | 0.001034 | -0.002984 | -0.2559 | 0.1034 | -4.9753
pf_sweep | 6 | 3 | 3 | 1 | 120 | 1 | -0.5 | 60 | 119.9961 | 0.99997 | -49.9859 | 0.059979 | 0.10393 | 0.11999 | 119.689 | 1.001 | 0.056979 | -0.30707 | 0.001034 | -0.003 | -0.2559 | 0.1034 | -5.002
pf_sweep | 6 | 3 | 4 | 1 | 120 | 1 | -0.5 | 60 | 119.9961 | 0.99997 | -49.9859 | 0.059979 | 0.10393 | 0.11999 | 119.689 | 1 | 0.056912 | -0.30707 | 3.4e-05 | -0.003067 | -0.2559 | 0.0034 | -5.1137
pf_sweep | 6 | 3 | 5 | 1 | 120 | 1 | -0.5 | 60 | 119.9961 | 0.99997 | -49.9859 | 0.059979 | 0.10393 | 0.11999 | 119.689 | 1.001 | 0.057053 | -0.30707 | 0.001034 | -0.002926 | -0.2559 | 0.1034 | -4.8786
pf_sweep | 6 | 3 | 6 | 1 | 120 | 1 | -0.5 | 60 | 119.9961 | 0.99997 | -49.9859 | 0.059979 | 0.10393 | 0.11999 | 119.689 | 0.998 | 0.057281 | -0.30707 | -0.001966 | -0.002698 | -0.2559 | -0.1966 | -4.4985
pf_sweep | 6 | 3 | 7 | 2 | 120 | 1 | -0.5 | 60 | 119.982 | 1.0002 | -49.9321 | 0.059922 | 0.10398 | 0.12001 | 119.751 | 2.978 | 0.1675 | -0.23099 | 1.9778 | 0.10758 | -0.1925 | -0.7543 | -6.8225
pf_sweep | 6 | 3 | 8 | 2 | 120 | 1 | -0.5 | 60 | 119.982 | 1.0002 | -49.9321 | 0.059922 | 0.10398 | 0.12001 | 119.751 | 1 | 0.056922 | -0.23099 | -0.000211 | -0.003 | -0.1925 | -0.0211 | -5.0068
pf_sweep | 6 | 3 | 9 | 2 | 120 | 1 | -0.5 | 60 | 119.982 | 1.0002 | -49.9321 | 0.059922 | 0.10398 | 0.12001 | 119.751 | 1.001 | 0.057059 | -0.23099 | 0.000789 | -0.002863 | -0.1925 | 0.0789 | -4.7782
pf_sweep | 6 | 3 | 10 | 2 | 120 | 1 | -0.5 | 60 | 119.982 | 1.0002 | -49.9321 | 0.059922 | 0.10398 | 0.12001 | 119.751 | 1 | 0.056976 | -0.23099 | -0.000211 | -0.002946 | -0.1925 | -0.0211 | -4.9167
pf_sweep | 6 | 3 | 11 | 2 | 120 | 1 | -0.5 | 60 | 119.982 | 1.0002 | -49.9321 | 0.059922 | 0.10398 | 0.12001 | 119.751 | 1 | 0.057206 | -0.23099 | -0.000211 | -0.002716 | -0.1925 | -0.0211 | -4.5328
pf_sweep | 6 | 3 | 12 | 2 | 120 | 1 | -0.5 | 60 | 119.982 | 1.0002 | -49.9321 | 0.059922 | 0.10398 | 0.12001 | 119.751 | 1 | 0.056875 | -0.23099 | -0.000211 | -0.003047 | -0.1925 | -0.0211 | -5.0852
pf_sweep | 6 | 3 | 13 | 3 | 120 | 1 | -0.5 | 60 | 119.9914 | 0.99991 | -49.9858 | 0.059973 | 0.10392 | 0.11998 | 119.78 | 0.999 | 0.05712 | -0.21137 | -0.000913 | -0.002853 | -0.1762 | -0.0914 | -4.7579
pf_sweep | 6 | 3 | 14 | 3 | 120 | 1 | -0.5 | 60 | 119.9914 | 0.99991 | -49.9858 | 0.059973 | 0.10392 | 0.11998 | 119.78 | 0.999 | 0.057268 | -0.21137 | -0.000913 | -0.002705 | -0.1762 | -0.0914 | -4.5111
pf_sweep | 6 | 3 | 15 | 3 | 120 | 1 | -0.5 | 60 | 119.9914 | 0.99991 | -49.9858 | 0.059973 | 0.10392 | 0.11998 | 119.78 | 1 | 0.05721 | -0.21137 | 8.7e-05 | -0.002763 | -0.1762 | 0.0087 | -4.6078
pf_sweep | 6 | 3 | 16 | 3 | 120 | 1 | -0.5 | 60 | 119.9914 | 0.99991 | -49.9858 | 0.059973 | 0.10392 | 0.11998 | 119.78 | 0.999 | 0.057358 | -0.21137 | -0.000913 | -0.002615 | -0.1762 | -0.0914 | -4.3611
pf_sweep | 6 | 3 | 17 | 3 | 120 | 1 | -0.5 | 60 | 119.9914 | 0.99991 | -49.9858 | 0.059973 | 0.10392 | 0.11998 | 119.78 | 1 | 0.057182 | -0.21137 | 8.7e-05 | -0.002791 | -0.1762 | 0.0087 | -4.6545
pf_sweep | 6 | 3 | 18 | 3 | 120 | 1 | -0.5 | 60 | 119.9914 | 0.99991 | -49.9858 | 0.059973 | 0.10392 | 0.11998 | 119.78 | 1 | 0.057263 | -0.21137 | 8.7e-05 | -0.00271 | -0.1762 | 0.0087 | -4.5195
pf_sweep | 7 | 1 | 1 | 1 | 120 | 1 | -0.4 | 60 | 120.0262 | 0.99997 | -39.9841 | 0.04799 | 0.11001 | 0.12002 | 119.741 | 2.968 | 0.13076 | -0.28522 | 1.968 | 0.082773 | -0.2376 | -1.0639 | -9.1737
pf_sweep | 7 | 1 | 2 | 1 | 120 | 1 | -0.4 | 60 | 120.0262 | 0.99997 | -39.9841 | 0.04799 | 0.11001 | 0.12002 | 119.741 | 1.001 | 0.044762 | -0.28522 | 0.001028 | -0.003228 | -0.2376 | 0.1028 | -6.7266
pf_sweep | 7 | 1 | 3 | 1 | 120 | 1 | -0.4 | 60 | 120.0262 | 0.99997 | -39.9841 | 0.04799 | 0.11001 | 0.12002 | 119.741 | 1.001 | 0.044768 | -0.28522 | 0.001028 | -0.003222 | -0.2376 | 0.1028 | -6.7141
pf_sweep | 7 | 1 | 4 | 1 | 120 | 1 | -0.4 | 60 | 120.0262 | 0.99997 | -39.9841 | 0.04799 | 0.11001 | 0.12002 | 119.741 | 1 | 0.044686 | -0.28522 | 2.8e-05 | -0.003304 | -0.2376 | 0.0028 | -6.885
pf_sweep | 7 | 1 | 5 | 1 | 120 | 1 | -0.4 | 60 | 120.0262 | 0.99997 | -39.9841 | 0.04799 | 0.11001 | 0.12002 | 119.741 | 1.001 | 0.044838 | -0.28522 | 0.001028 | -0.003152 | -0.2376 | 0.1028 | -6.5683
pf_sweep | 7 | 1 | 6 | 1 | 120 | 1 | -0.4 | 60 | 120.0262 | 0.99997 | -39.9841 | 0.04799 | 0.11001 | 0.12002 | 119.741 | 0.998 | 0.045181 | -0.28522 | -0.001972 | -0.002809 | -0.2376 | -0.1972 | -5.8535
pf_sweep | 7 | 1 | 7 | 2 | 120 | 1 | -0.4 | 60 | 119.9814 | 1.0002 | -39.9265 | 0.047915 | 0.11003 | 0.12001 | 119.771 | 2.979 | 0.13104 | -0.21041 | 1.9788 | 0.083126 | -0.1754 | -0.72154 | -8.8374
pf_sweep | 7 | 1 | 8 | 2 | 120 | 1 | -0.4 | 60 | 119.9814 | 1.0002 | -39.9265 | 0.047915 | 0.11003 | 0.12001 | 119.771 | 1 | 0.044697 | -0.21041 | -0.000217 | -0.003218 | -0.1754 | -0.0217 | -6.7156
pf_sweep | 7 | 1 | 9 | 2 | 120 | 1 | -0.4 | 60 | 119.9814 | 1.0002 | -39.9265 | 0.047915 | 0.11003 | 0.12001 | 119.771 | 1.001 | 0.044821 | -0.21041 | 0.000783 | -0.003094 | -0.1754 | 0.0783 | -6.4568
pf_sweep | 7 | 1 | 10 | 2 | 120 | 1 | -0.4 | 60 | 119.9814 | 1.0002 | -39.9265 | 0.047915 | 0.11003 | 0.12001 | 119.771 | 1 | 0.044752 | -0.21041 | -0.000217 | -0.003163 | -0.1754 | -0.0217 | -6.6008
pf_sweep | 7 | 1 | 11 | 2 | 120 | 1 | -0.4 | 60 | 119.9814 | 1.0002 | -39.9265 | 0.047915 | 0.11003 | 0.12001 | 119.771 | 1 | 0.045021 | -0.21041 | -0.000217 | -0.002894 | -0.1754 | -0.0217 | -6.0394
pf_sweep | 7 | 1 | 12 | 2 | 120 | 1 | -0.4 | 60 | 119.9814 | 1.0002 | -39.9265 | 0.047915 | 0.11003 | 0.12001 | 119.771 | 1 | 0.044638 | -0.21041 | -0.000217 | -0.003277 | -0.1754 | -0.0217 | -6.8387
pf_sweep | 7 | 1 | 13 | 3 | 120 | 1 | -0.4 | 60 | 119.9885 | 0.99992 | -39.9863 | 0.047975 | 0.10997 | 0.11998 | 119.765 | 1 | 0.044927 | -0.22352 | 7.9e-05 | -0.003048 | -0.1863 | 0.0079 | -6.3536
pf_sweep | 7 | 1 | 14 | 3 | 120 | 1 | -0.4 | 60 | 119.9885 | 0.99992 | -39.9863 | 0.047975 | 0.10997 | 0.11998 | 119.765 | 0.999 | 0.045087 | -0.22352 | -0.000921 | -0.002888 | -0.1863 | -0.0921 | -6.0201
pf_sweep | 7 | 1 | 15 | 3 | 120 | 1 | -0.4 | 60 | 119.9885 | 0.99992 | -39.9863 | 0.047975 | 0.10997 | 0.11998 | 119.765 | 1 | 0.04503 | -0.22352 | 7.9e-05 | -0.002945 | -0.1863 | 0.0079 | -6.1389
pf_sweep | 7 | 1 | 16 | 3 | 120 | 1 | -0.4 | 60 | 119.9885 | 0.99992 | -39.9863 | 0.047975 | 0.10997 | 0.11998 | 119.765 | 0.999 | 0.045181 | -0.22352 | -0.000921 | -0.002794 | -0.1863 | -0.0921 | -5.8242
pf_sweep | 7 | 1 | 17 | 3 | 120 | 1 | -0.4 | 60 | 119.9885 | 0.99992 | -39.9863 | 0.047975 | 0.10997 | 0.11998 | 119.765 | 1 | 0.045004 | -0.22352 | 7.9e-05 | -0.002971 | -0.1863 | 0.0079 | -6.1931
pf_sweep | 7 | 1 | 18 | 3 | 120 | 1 | -0.4 | 60 | 119.9885 | 0.99992 | -39.9863 | 0.047975 | 0.10997 | 0.11998 | 119.765 | 1 | 0.045099 | -0.22352 | 7.9e-05 | -0.002876 | -0.1863 | 0.0079 | -5.9951
pf_sweep | 7 | 2 | 1 | 1 | 120 | 1 | -0.4 | 60 | 120.0118 | 0.99995 | -39.9837 | 0.047983 | 0.11 | 0.12001 | 119.745 | 2.968 | 0.13082 | -0.26679 | 1.9681 | 0.082836 | -0.2223 | -1.0614 | -9.1206
pf_sweep | 7 | 2 | 2 | 1 | 120 | 1 | -0.4 | 60 | 120.0118 | 0.99995 | -39.9837 | 0.047983 | 0.11 | 0.12001 | 119.745 | 1.001 | 0.044783 | -0.26679 | 0.001053 | -0.0032 | -0.2223 | 0.1053 | -6.6684
pf_sweep | 7 | 2 | 3 | 1 | 120 | 1 | -0.4 | 60 | 120.0118 | 0.99995 | -39.9837 | 0.047983 | 0.11 | 0.12001 | 119.745 | 1.001 | 0.044742 | -0.26679 | 0.001053 | -0.003241 | -0.2223 | 0.1053 | -6.7538
pf_sweep | 7 | 2 | 4 | 1 | 120 | 1 | -0.4 | 60 | 120.0118 | 0.99995 | -39.9837 | 0.047983 | 0.11 | 0.12001 | 119.745 | 1 | 0.044696 | -0.26679 | 5.3e-05 | -0.003287 | -0.2223 | 0.0053 | -6.8497
pf_sweep | 7 | 2 | 5 | 1 | 120 | 1 | -0.4 | 60 | 120.0118 | 0.99995 | -39.9837 | 0.047983 | 0.11 | 0.12001 | 119.745 | 1.001 | 0.044859 | -0.26679 | 0.001053 | -0.003124 | -0.2223 | 0.1053 | -6.51
pf_sweep | 7 | 2 | 6 | 1 | 120 | 1 | -0.4 | 60 | 120.0118 | 0.99995 | -39.9837 | 0.047983 | 0.11 | 0.12001 | 119.745 | 0.998 | 0.045136 | -0.26679 | -0.001947 | -0.002847 | -0.2223 | -0.1947 | -5.9327
pf_sweep | 7 | 2 | 7 | 2 | 120 | 1 | -0.4 | 60 | 119.9839 | 1.0002 | -39.9256 | 0.047915 | 0.11003 | 0.12001 | 119.767 | 2.978 | 0.13112 | -0.21695 | 1.9778 | 0.083202 | -0.1808 | -0.75464 | -8.7842
pf_sweep | 7 | 2 | 8 | 2 | 120 | 1 | -0.4 | 60 | 119.9839 | 1.0002 | -39.9256 | 0.047915 | 0.11003 | 0.12001 | 119.767 | 1 | 0.044695 | -0.21695 | -0.000215 | -0.00322 | -0.1808 | -0.0215 | -6.7194
pf_sweep | 7 | 2 | 9 | 2 | 120 | 1 | -0.4 | 60 | 119.9839 | 1.0002 | -39.9256 | 0.047915 | 0.11003 | 0.12001 | 119.767 | 1.001 | 0.044813 | -0.21695 | 0.000785 | -0.003102 | -0.1808 | 0.0785 | -6.4732
pf_sweep | 7 | 2 | 10 | 2 | 120 | 1 | -0.4 | 60 | 119.9839 | 1.0002 | -39.9256 | 0.047915 | 0.11003 | 0.12001 | 119.767 | 1 | 0.044791 | -0.21695 | -0.000215 | -0.003124 | -0.1808 | -0.0215 | -6.5191
pf_sweep | 7 | 2 | 11 | 2 | 120 | 1 | -0.4 | 60 | 119.9839 | 1.0002 | -39.9256 | 0.047915 | 0.11003 | 0.12001 | 119.767 | 1 | 0.04501 | -0.21695 | -0.000215 | -0.002905 | -0.1808 | -0.0215 | -6.062
pf_sweep | 7 | 2 | 12 | 2 | 120 | 1 | -0.4 | 60 | 119.9839 | 1.0002 | -39.9256 | 0.047915 | 0.11003 | 0.12001 | 119.767 | 1 | 0.044649 | -0.21695 | -0.000215 | -0.003266 | -0.1808 | -0.0215 | -6.8154
pf_sweep | 7 | 2 | 13 | 3 | 120 | 1 | -0.4 | 60 | 119.9867 | 0.9999 | -39.9843 | 0.047971 | 0.10997 | 0.11997 | 119.762 | 1 | 0.044914 | -0.22466 | 0.000103 | -0.003057 | -0.1872 | 0.0103 | -6.3723
pf_sweep | 7 | 2 | 14 | 3 | 120 | 1 | -0.4 | 60 | 119.9867 | 0.9999 | -39.9843 | 0.047971 | 0.10997 | 0.11997 | 119.762 | 0.999 | 0.045076 | -0.22466 | -0.000897 | -0.002895 | -0.1872 | -0.0897 | -6.0346
pf_sweep | 7 | 2 | 15 | 3 | 120 | 1 | -0.4 | 60 | 119.9867 | 0.9999 | -39.9843 | 0.047971 | 0.10997 | 0.11997 | 119.762 | 1 | 0.045047 | -0.22466 | 0.000103 | -0.002924 | -0.1872 | 0.0103 | -6.0951
pf_sweep | 7 | 2 | 16 | 3 | 120 | 1 | -0.4 | 60 | 119.9867 | 0.9999 | -39.9843 | 0.047971 | 0.10997 | 0.11997 | 119.762 | 0.999 | 0.045169 | -0.22466 | -0.000897 | -0.002802 | -0.1872 | -0.0897 | -5.8408
pf_sweep | 7 | 2 | 17 | 3 | 120 | 1 | -0.4 | 60 | 119.9867 | 0.9999 | -39.9843 | 0.047971 | 0.10997 | 0.11997 | 119.762 | 1 | 0.045003 | -0.22466 | 0.000103 | -0.002968 | -0.1872 | 0.0103 | -6.1868
pf_sweep | 7 | 2 | 18 | 3 | 120 | 1 | -0.4 | 60 | 119.9867 | 0.9999 | -39.9843 | 0.047971 | 0.10997 | 0.11997 | 119.762 | 1 | 0.045106 | -0.22466 | 0.000103 | -0.002865 | -0.1872 | 0.0103 | -5.9721
pf_sweep | 7 | 3 | 1 | 1 | 120 | 1 | -0.4 | 60 | 120.0154 | 0.99996 | -39.9825 | 0.047983 | 0.11 | 0.12001 | 119.736 | 2.968 | 0.13081 | -0.27945 | 1.968 | 0.08283 | -0.2328 | -1.0625 | -9.1259
pf_sweep | 7 | 3 | 2 | 1 | 120 | 1 | -0.4 | 60 | 120.0154 | 0.99996 | -39.9825 | 0.047983 | 0.11 | 0.12001 | 119.736 | 1.001 | 0.044764 | -0.27945 | 0.001042 | -0.003219 | -0.2328 | 0.1042 | -6.709
pf_sweep | 7 | 3 | 3 | 1 | 120 | 1 | -0.4 | 60 | 120.0154 | 0.99996 | -39.9825 | 0.047983 | 0.11 | 0.12001 | 119.736 | 1.001 | 0.044776 | -0.27945 | 0.001042 | -0.003207 | -0.2328 | 0.1042 | -6.684
pf_sweep | 7 | 3 | 4 | 1 | 120 | 1 | -0.4 | 60 | 120.0154 | 0.99996 | -39.9825 | 0.047983 | 0.11 | 0.12001 | 119.736 | 1 | 0.044712 | -0.27945 | 4.2e-05 | -0.003271 | -0.2328 | 0.0042 | -6.8174
pf_sweep | 7 | 3 | 5 | 1 | 120 | 1 | -0.4 | 60 | 120.0154 | 0.99996 | -39.9825 | 0.047983 | 0.11 | 0.12001 | 119.736 | 1.001 | 0.044831 | -0.27945 | 0.001042 | -0.003152 | -0.2328 | 0.1042 | -6.5694
pf_sweep | 7 | 3 | 6 | 1 | 120 | 1 | -0.4 | 60 | 120.0154 | 0.99996 | -39.9825 | 0.047983 | 0.11 | 0.12001 | 119.736 | 0.998 | 0.045151 | -0.27945 | -0.001958 | -0.002832 | -0.2328 | -0.1958 | -5.9025
pf_sweep | 7 | 3 | 7 | 2 | 120 | 1 | -0.4 | 60 | 119.9812 | 1.0002 | -39.9273 | 0.047916 | 0.11003 | 0.12001 | 119.764 | 2.978 | 0.1311 | -0.21724 | 1.9778 | 0.083189 | -0.1811 | -0.75644 | -8.7961
pf_sweep | 7 | 3 | 8 | 2 | 120 | 1 | -0.4 | 60 | 119.9812 | 1.0002 | -39.9273 | 0.047916 | 0.11003 | 0.12001 | 119.764 | 1 | 0.044691 | -0.21724 | -0.000233 | -0.003225 | -0.1811 | -0.0233 | -6.7314
pf_sweep | 7 | 3 | 9 | 2 | 120 | 1 | -0.4 | 60 | 119.9812 | 1.0002 | -39.9273 | 0.047916 | 0.11003 | 0.12001 | 119.764 | 1.001 | 0.044831 | -0.21724 | 0.000767 | -0.003085 | -0.1811 | 0.0767 | -6.4393
pf_sweep | 7 | 3 | 10 | 2 | 120 | 1 | -0.4 | 60 | 119.9812 | 1.0002 | -39.9273 | 0.047916 | 0.11003 | 0.12001 | 119.764 | 1 | 0.044777 | -0.21724 | -0.000233 | -0.003139 | -0.1811 | -0.0233 | -6.5519
pf_sweep | 7 | 3 | 11 | 2 | 120 | 1 | -0.4 | 60 | 119.9812 | 1.0002 | -39.9273 | 0.047916 | 0.11003 | 0.12001 | 119.764 | 1 | 0.045025 | -0.21724 | -0.000233 | -0.002891 | -0.1811 | -0.0233 | -6.0344
pf_sweep | 7 | 3 | 12 | 2 | 120 | 1 | -0.4 | 60 | 119.9812 | 1.0002 | -39.9273 | 0.047916 | 0.11003 | 0.12001 | 119.764 | 1 | 0.044648 | -0.21724 | -0.000233 | -0.003268 | -0.1811 | -0.0233 | -6.8212
pf_sweep | 7 | 3 | 13 | 3 | 120 | 1 | -0.4 | 60 | 119.9915 | 0.99991 | -39.9859 | 0.047975 | 0.10997 | 0.11998 | 119.78 | 1 | 0.044914 | -0.21152 | 8.9e-05 | -0.003061 | -0.1763 | 0.0089 | -6.3811
pf_sweep | 7 | 3 | 14 | 3 | 120 | 1 | -0.4 | 60 | 119.9915 | 0.99991 | -39.9859 | 0.047975 | 0.10997 | 0.11998 | 119.78 | 0.999 | 0.04509 | -0.21152 | -0.000911 | -0.002885 | -0.1763 | -0.0911 | -6.0143
pf_sweep | 7 | 3 | 15 | 3 | 120 | 1 | -0.4 | 60 | 119.9915 | 0.99991 | -39.9859 | 0.047975 | 0.10997 | 0.11998 | 119.78 | 1 | 0.045001 | -0.21152 | 8.9e-05 | -0.002974 | -0.1763 | 0.0089 | -6.1998
pf_sweep | 7 | 3 | 16 | 3 | 120 | 1 | -0.4 | 60 | 119.9915 | 0.99991 | -39.9859 | 0.047975 | 0.10997 | 0.11998 | 119.78 | 0.999 | 0.045159 | -0.21152 | -0.000911 | -0.002816 | -0.1763 | -0.0911 | -5.8704
pf_sweep | 7 | 3 | 17 | 3 | 120 | 1 | -0.4 | 60 | 119.9915 | 0.99991 | -39.9859 | 0.047975 | 0.10997 | 0.11998 | 119.78 | 1 | 0.045008 | -0.21152 | 8.9e-05 | -0.002967 | -0.1763 | 0.0089 | -6.1852
pf_sweep | 7 | 3 | 18 | 3 | 120 | 1 | -0.4 | 60 | 119.9915 | 0.99991 | -39.9859 | 0.047975 | 0.10997 | 0.11998 | 119.78 | 0.999 | 0.04507 | -0.21152 | -0.000911 | -0.002905 | -0.1763 | -0.0911 | -6.056
pf_sweep | 8 | 1 | 1 | 1 | 120 | 1 | -0.3 | 60 | 120.0006 | 0.99994 | -29.9839 | 0.035979 | 0.11447 | 0.11999 | 119.774 | 2.968 | 0.094757 | -0.22662 | 1.9681 | 0.058778 | -0.1888 | -1.0607 | -12.21
pf_sweep | 8 | 1 | 2 | 1 | 120 | 1 | -0.3 | 60 | 120.0006 | 0.99994 | -29.9839 | 0.035979 | 0.11447 | 0.11999 | 119.774 | 1.001 | 0.032605 | -0.22662 | 0.00106 | -0.003374 | -0.1888 | 0.1061 | -9.3768
pf_sweep | 8 | 1 | 3 | 1 | 120 | 1 | -0.3 | 60 | 120.0006 | 0.99994 | -29.9839 | 0.035979 | 0.11447 | 0.11999 | 119.774 | 1.001 | 0.032571 | -0.22662 | 0.00106 | -0.003408 | -0.1888 | 0.1061 | -9.4713
pf_sweep | 8 | 1 | 4 | 1 | 120 | 1 | -0.3 | 60 | 120.0006 | 0.99994 | -29.9839 | 0.035979 | 0.11447 | 0.11999 | 119.774 | 1 | 0.032556 | -0.22662 | 6e-05 | -0.003423 | -0.1888 | 0.0061 | -9.513
pf_sweep | 8 | 1 | 5 | 1 | 120 | 1 | -0.3 | 60 | 120.0006 | 0.99994 | -29.9839 | 0.035979 | 0.11447 | 0.11999 | 119.774 | 1.001 | 0.032669 | -0.22662 | 0.00106 | -0.00331 | -0.1888 | 0.1061 | -9.1989
pf_sweep | 8 | 1 | 6 | 1 | 120 | 1 | -0.3 | 60 | 120.0006 | 0.99994 | -29.9839 | 0.035979 | 0.11447 | 0.11999 | 119.774 | 0.998 | 0.033048 | -0.22662 | -0.00194 | -0.002931 | -0.1888 | -0.194 | -8.1455
pf_sweep | 8 | 1 | 7 | 2 | 120 | 1 | -0.3 | 60 | 119.9798 | 1.0002 | -29.9246 | 0.035912 | 0.11451 | 0.12001 | 119.767 | 2.979 | 0.094793 | -0.2128 | 1.9788 | 0.058881 | -0.1774 | -0.72407 | -12.014
pf_sweep | 8 | 1 | 8 | 2 | 120 | 1 | -0.3 | 60 | 119.9798 | 1.0002 | -29.9246 | 0.035912 | 0.11451 | 0.12001 | 119.767 | 1 | 0.032494 | -0.2128 | -0.000242 | -0.003418 | -0.1774 | -0.0242 | -9.5181
pf_sweep | 8 | 1 | 9 | 2 | 120 | 1 | -0.3 | 60 | 119.9798 | 1.0002 | -29.9246 | 0.035912 | 0.11451 | 0.12001 | 119.767 | 1.001 | 0.03266 | -0.2128 | 0.000758 | -0.003252 | -0.1774 | 0.0757 | -9.0558
pf_sweep | 8 | 1 | 10 | 2 | 120 | 1 | -0.3 | 60 | 119.9798 | 1.0002 | -29.9246 | 0.035912 | 0.11451 | 0.12001 | 119.767 | 1 | 0.032623 | -0.2128 | -0.000242 | -0.003289 | -0.1774 | -0.0242 | -9.1589
pf_sweep | 8 | 1 | 11 | 2 | 120 | 1 | -0.3 | 60 | 119.9798 | 1.0002 | -29.9246 | 0.035912 | 0.11451 | 0.12001 | 119.767 | 1 | 0.032819 | -0.2128 | -0.000242 | -0.003093 | -0.1774 | -0.0242 | -8.6131
pf_sweep | 8 | 1 | 12 | 2 | 120 | 1 | -0.3 | 60 | 119.9798 | 1.0002 | -29.9246 | 0.035912 | 0.11451 | 0.12001 | 119.767 | 1 | 0.032488 | -0.2128 | -0.000242 | -0.003424 | -0.1774 | -0.0242 | -9.5348
pf_sweep | 8 | 1 | 13 | 3 | 120 | 1 | -0.3 | 60 | 119.9913 | 0.99993 | -29.9845 | 0.035976 | 0.11446 | 0.11998 | 119.764 | 0.999 | 0.032744 | -0.22726 | -0.000927 | -0.003232 | -0.1894 | -0.0927 | -8.9843
pf_sweep | 8 | 1 | 14 | 3 | 120 | 1 | -0.3 | 60 | 119.9913 | 0.99993 | -29.9845 | 0.035976 | 0.11446 | 0.11998 | 119.764 | 0.999 | 0.032955 | -0.22726 | -0.000927 | -0.003021 | -0.1894 | -0.0927 | -8.3978
pf_sweep | 8 | 1 | 15 | 3 | 120 | 1 | -0.3 | 60 | 119.9913 | 0.99993 | -29.9845 | 0.035976 | 0.11446 | 0.11998 | 119.764 | 1.001 | 0.032872 | -0.22726 | 0.001073 | -0.003104 | -0.1894 | 0.1073 | -8.6285
pf_sweep | 8 | 1 | 16 | 3 | 120 | 1 | -0.3 | 60 | 119.9913 | 0.99993 | -29.9845 | 0.035976 | 0.11446 | 0.11998 | 119.764 | 0.999 | 0.032994 | -0.22726 | -0.000927 | -0.002982 | -0.1894 | -0.0927 | -8.2894
pf_sweep | 8 | 1 | 17 | 3 | 120 | 1 | -0.3 | 60 | 119.9913 | 0.99993 | -29.9845 | 0.035976 | 0.11446 | 0.11998 | 119.764 | 1 | 0.032887 | -0.22726 | 7.3e-05 | -0.003089 | -0.1894 | 0.0073 | -8.5868
pf_sweep | 8 | 1 | 18 | 3 | 120 | 1 | -0.3 | 60 | 119.9913 | 0.99993 | -29.9845 | 0.035976 | 0.11446 | 0.11998 | 119.764 | 1 | 0.032937 | -0.22726 | 7.3e-05 | -0.003039 | -0.1894 | 0.0073 | -8.4478
pf_sweep | 8 | 2 | 1 | 1 | 120 | 1 | -0.3 | 60 | 120.021 | 0.99997 | -29.9811 | 0.035983 | 0.1145 | 0.12002 | 119.741 | 2.968 | 0.094731 | -0.28 | 1.968 | 0.058748 | -0.2333 | -1.0638 | -12.2436
pf_sweep | 8 | 2 | 2 | 1 | 120 | 1 | -0.3 | 60 | 120.021 | 0.99997 | -29.9811 | 0.035983 | 0.1145 | 0.12002 | 119.741 | 1.001 | 0.032598 | -0.28 | 0.001029 | -0.003385 | -0.2333 | 0.1029 | -9.4061
pf_sweep | 8 | 2 | 3 | 1 | 120 | 1 | -0.3 | 60 | 120.021 | 0.99997 | -29.9811 | 0.035983 | 0.1145 | 0.12002 | 119.741 | 1.001 | 0.03258 | -0.28 | 0.001029 | -0.003403 | -0.2333 | 0.1029 | -9.4561
pf_sweep | 8 | 2 | 4 | 1 | 120 | 1 | -0.3 | 60 | 120.021 | 0.99997 | -29.9811 | 0.035983 | 0.1145 | 0.12002 | 119.741 | 1 | 0.032557 | -0.28 | 2.9e-05 | -0.003426 | -0.2333 | 0.0029 | -9.52
pf_sweep | 8 | 2 | 5 | 1 | 120 | 1 | -0.3 | 60 | 120.021 | 0.99997 | -29.9811 | 0.035983 | 0.1145 | 0.12002 | 119.741 | 1.001 | 0.032672 | -0.28 | 0.001029 | -0.003311 | -0.2333 | 0.1029 | -9.2004
pf_sweep | 8 | 2 | 6 | 1 | 120 | 1 | -0.3 | 60 | 120.021 | 0.99997 | -29.9811 | 0.035983 | 0.1145 | 0.12002 | 119.741 | 0.998 | 0.033008 | -0.28 | -0.001971 | -0.002975 | -0.2333 | -0.1971 | -8.2667
pf_sweep | 8 | 2 | 7 | 2 | 120 | 1 | -0.3 | 60 | 119.9792 | 1.0002 | -29.9216 | 0.035908 | 0.11451 | 0.12001 | 119.762 | 2.979 | 0.09479 | -0.21716 | 1.9788 | 0.058882 | -0.181 | -0.72247 | -12.0062
pf_sweep | 8 | 2 | 8 | 2 | 120 | 1 | -0.3 | 60 | 119.9792 | 1.0002 | -29.9216 | 0.035908 | 0.11451 | 0.12001 | 119.762 | 1 | 0.032515 | -0.21716 | -0.000226 | -0.003393 | -0.181 | -0.0226 | -9.4487
pf_sweep | 8 | 2 | 9 | 2 | 120 | 1 | -0.3 | 60 | 119.9792 | 1.0002 | -29.9216 | 0.035908 | 0.11451 | 0.12001 | 119.762 | 1.001 | 0.032671 | -0.21716 | 0.000774 | -0.003237 | -0.181 | 0.0773 | -9.0143
pf_sweep | 8 | 2 | 10 | 2 | 120 | 1 | -0.3 | 60 | 119.9792 | 1.0002 | -29.9216 | 0.035908 | 0.11451 | 0.12001 | 119.762 | 1 | 0.032602 | -0.21716 | -0.000226 | -0.003306 | -0.181 | -0.0226 | -9.2064
pf_sweep | 8 | 2 | 11 | 2 | 120 | 1 | -0.3 | 60 | 119.9792 | 1.0002 | -29.9216 | 0.035908 | 0.11451 | 0.12001 | 119.762 | 1 | 0.032859 | -0.21716 | -0.000226 | -0.003049 | -0.181 | -0.0226 | -8.4907
pf_sweep | 8 | 2 | 12 | 2 | 120 | 1 | -0.3 | 60 | 119.9792 | 1.0002 | -29.9216 | 0.035908 | 0.11451 | 0.12001 | 119.762 | 1 | 0.032499 | -0.21716 | -0.000226 | -0.003409 | -0.181 | -0.0226 | -9.4933
pf_sweep | 8 | 2 | 13 | 3 | 120 | 1 | -0.3 | 60 | 119.9881 | 0.99991 | -29.9882 | 0.035979 | 0.11446 | 0.11998 | 119.766 | 1 | 0.032725 | -0.22211 | 8.7e-05 | -0.003254 | -0.1851 | 0.0087 | -9.0444
pf_sweep | 8 | 2 | 14 | 3 | 120 | 1 | -0.3 | 60 | 119.9881 | 0.99991 | -29.9882 | 0.035979 | 0.11446 | 0.11998 | 119.766 | 0.999 | 0.032942 | -0.22211 | -0.000913 | -0.003037 | -0.1851 | -0.0913 | -8.4413
pf_sweep | 8 | 2 | 15 | 3 | 120 | 1 | -0.3 | 60 | 119.9881 | 0.99991 | -29.9882 | 0.035979 | 0.11446 | 0.11998 | 119.766 | 1.001 | 0.032862 | -0.22211 | 0.001087 | -0.003117 | -0.1851 | 0.1087 | -8.6636
pf_sweep | 8 | 2 | 16 | 3 | 120 | 1 | -0.3 | 60 | 119.9881 | 0.99991 | -29.9882 | 0.035979 | 0.11446 | 0.11998 | 119.766 | 0.999 | 0.032995 | -0.22211 | -0.000913 | -0.002984 | -0.1851 | -0.0913 | -8.294
pf_sweep | 8 | 2 | 17 | 3 | 120 | 1 | -0.3 | 60 | 119.9881 | 0.99991 | -29.9882 | 0.035979 | 0.11446 | 0.11998 | 119.766 | 1 | 0.032854 | -0.22211 | 8.7e-05 | -0.003125 | -0.1851 | 0.0087 | -8.6859
pf_sweep | 8 | 2 | 18 | 3 | 120 | 1 | -0.3 | 60 | 119.9881 | 0.99991 | -29.9882 | 0.035979 | 0.11446 | 0.11998 | 119.766 | 1 | 0.032937 | -0.22211 | 8.7e-05 | -0.003042 | -0.1851 | 0.0087 | -8.4552
pf_sweep | 8 | 3 | 1 | 1 | 120 | 1 | -0.3 | 60 | 120.0099 | 0.99997 | -29.984 | 0.035983 | 0.11448 | 0.12001 | 119.765 | 2.968 | 0.094799 | -0.24493 | 1.968 | 0.058816 | -0.2041 | -1.0638 | -12.181
pf_sweep | 8 | 3 | 2 | 1 | 120 | 1 | -0.3 | 60 | 120.0099 | 0.99997 | -29.984 | 0.035983 | 0.11448 | 0.12001 | 119.765 | 1.001 | 0.032599 | -0.24493 | 0.001029 | -0.003384 | -0.2041 | 0.1029 | -9.4037
pf_sweep | 8 | 3 | 3 | 1 | 120 | 1 | -0.3 | 60 | 120.0099 | 0.99997 | -29.984 | 0.035983 | 0.11448 | 0.12001 | 119.765 | 1.001 | 0.032636 | -0.24493 | 0.001029 | -0.003347 | -0.2041 | 0.1029 | -9.3009
pf_sweep | 8 | 3 | 4 | 1 | 120 | 1 | -0.3 | 60 | 120.0099 | 0.99997 | -29.984 | 0.035983 | 0.11448 | 0.12001 | 119.765 | 1 | 0.032593 | -0.24493 | 2.9e-05 | -0.00339 | -0.2041 | 0.0029 | -9.4204
pf_sweep | 8 | 3 | 5 | 1 | 120 | 1 | -0.3 | 60 | 120.0099 | 0.99997 | -29.984 | 0.035983 | 0.11448 | 0.12001 | 119.765 | 1.001 | 0.032673 | -0.24493 | 0.001029 | -0.00331 | -0.2041 | 0.1029 | -9.1981
pf_sweep | 8 | 3 | 6 | 1 | 120 | 1 | -0.3 | 60 | 120.0099 | 0.99997 | -29.984 | 0.035983 | 0.11448 | 0.12001 | 119.765 | 0.998 | 0.033045 | -0.24493 | -0.001971 | -0.002938 | -0.2041 | -0.1971 | -8.1643
pf_sweep | 8 | 3 | 7 | 2 | 120 | 1 | -0.3 | 60 | 119.9845 | 1.0002 | -29.9241 | 0.035912 | 0.11451 | 0.12001 | 119.764 | 2.978 | 0.094693 | -0.22054 | 1.9778 | 0.058781 | -0.1838 | -0.75451 | -12.1063
pf_sweep | 8 | 3 | 8 | 2 | 120 | 1 | -0.3 | 60 | 119.9845 | 1.0002 | -29.9241 | 0.035912 | 0.11451 | 0.12001 | 119.764 | 1 | 0.032515 | -0.22054 | -0.000213 | -0.003397 | -0.1838 | -0.0213 | -9.4591
pf_sweep | 8 | 3 | 9 | 2 | 120 | 1 | -0.3 | 60 | 119.9845 | 1.0002 | -29.9241 | 0.035912 | 0.11451 | 0.12001 | 119.764 | 1.001 | 0.032656 | -0.22054 | 0.000787 | -0.003256 | -0.1838 | 0.0786 | -9.0664
pf_sweep | 8 | 3 | 10 | 2 | 120 | 1 | -0.3 | 60 | 119.9845 | 1.0002 | -29.9241 | 0.035912 | 0.11451 | 0.12001 | 119.764 | 1 | 0.032583 | -0.22054 | -0.000213 | -0.003329 | -0.1838 | -0.0213 | -9.2697
pf_sweep | 8 | 3 | 11 | 2 | 120 | 1 | -0.3 | 60 | 119.9845 | 1.0002 | -29.9241 | 0.035912 | 0.11451 | 0.12001 | 119.764 | 1 | 0.032849 | -0.22054 | -0.000213 | -0.003063 | -0.1838 | -0.0213 | -8.529
pf_sweep | 8 | 3 | 12 | 2 | 120 | 1 | -0.3 | 60 | 119.9845 | 1.0002 | -29.9241 | 0.035912 | 0.11451 | 0.12001 | 119.764 | 1 | 0.032483 | -0.22054 | -0.000213 | -0.003429 | -0.1838 | -0.0213 | -9.5482
pf_sweep | 8 | 3 | 13 | 3 | 120 | 1 | -0.3 | 60 | 119.9878 | 0.99991 | -29.9835 | 0.035973 | 0.11446 | 0.11998 | 119.781 | 1 | 0.032758 | -0.20682 | 9.2e-05 | -0.003215 | -0.1724 | 0.0092 | -8.9377
pf_sweep | 8 | 3 | 14 | 3 | 120 | 1 | -0.3 | 60 | 119.9878 | 0.99991 | -29.9835 | 0.035973 | 0.11446 | 0.11998 | 119.781 | 0.999 | 0.032933 | -0.20682 | -0.000908 | -0.00304 | -0.1724 | -0.0908 | -8.4512
pf_sweep | 8 | 3 | 15 | 3 | 120 | 1 | -0.3 | 60 | 119.9878 | 0.99991 | -29.9835 | 0.035973 | 0.11446 | 0.11998 | 119.781 | 1 | 0.032855 | -0.20682 | 9.2e-05 | -0.003118 | -0.1724 | 0.0092 | -8.6681
pf_sweep | 8 | 3 | 16 | 3 | 120 | 1 | -0.3 | 60 | 119.9878 | 0.99991 | -29.9835 | 0.035973 | 0.11446 | 0.11998 | 119.781 | 0.999 | 0.033028 | -0.20682 | -0.000908 | -0.002945 | -0.1724 | -0.0908 | -8.1871
pf_sweep | 8 | 3 | 17 | 3 | 120 | 1 | -0.3 | 60 | 119.9878 | 0.99991 | -29.9835 | 0.035973 | 0.11446 | 0.11998 | 119.781 | 1 | 0.032855 | -0.20682 | 9.2e-05 | -0.003118 | -0.1724 | 0.0092 | -8.6681
pf_sweep | 8 | 3 | 18 | 3 | 120 | 1 | -0.3 | 60 | 119.9878 | 0.99991 | -29.9835 | 0.035973 | 0.11446 | 0.11998 | 119.781 | 1 | 0.03292 | -0.20682 | 9.2e-05 | -0.003053 | -0.1724 | 0.0092 | -8.4874
pf_sweep | 9 | 1 | 1 | 1 | 120 | 1 | -0.2 | 60 | 120.0125 | 0.99996 | -19.981 | 0.023979 | 0.11759 | 0.12001 | 119.73 | 2.968 | 0.058779 | -0.28251 | 1.968 | 0.0348 | -0.2354 | -1.0628 | -18.2901
pf_sweep | 9 | 1 | 2 | 1 | 120 | 1 | -0.2 | 60 | 120.0125 | 0.99996 | -19.981 | 0.023979 | 0.11759 | 0.12001 | 119.73 | 1.001 | 0.02051 | -0.28251 | 0.001039 | -0.003469 | -0.2354 | 0.1039 | -14.4659
pf_sweep | 9 | 1 | 3 | 1 | 120 | 1 | -0.2 | 60 | 120.0125 | 0.99996 | -19.981 | 0.023979 | 0.11759 | 0.12001 | 119.73 | 1.001 | 0.020448 | -0.28251 | 0.001039 | -0.003531 | -0.2354 | 0.1039 | -14.7244
pf_sweep | 9 | 1 | 4 | 1 | 120 | 1 | -0.2 | 60 | 120.0125 | 0.99996 | -19.981 | 0.023979 | 0.11759 | 0.12001 | 119.73 | 1 | 0.020451 | -0.28251 | 3.9e-05 | -0.003528 | -0.2354 | 0.0039 | -14.7119
pf_sweep | 9 | 1 | 5 | 1 | 120 | 1 | -0.2 | 60 | 120.0125 | 0.99996 | -19.981 | 0.023979 | 0.11759 | 0.12001 | 119.73 | 1.001 | 0.020573 | -0.28251 | 0.001039 | -0.003406 | -0.2354 | 0.1039 | -14.2031
pf_sweep | 9 | 1 | 6 | 1 | 120 | 1 | -0.2 | 60 | 120.0125 | 0.99996 | -19.981 | 0.023979 | 0.11759 | 0.12001 | 119.73 | 0.998 | 0.02094 | -0.28251 | -0.001961 | -0.003039 | -0.2354 | -0.1961 | -12.6726
pf_sweep | 9 | 1 | 7 | 2 | 120 | 1 | -0.2 | 60 | 119.9804 | 1.0002 | -19.9225 | 0.023909 | 0.1176 | 0.12001 | 119.77 | 2.978 | 0.05857 | -0.2104 | 1.9778 | 0.034661 | -0.1754 | -0.75601 | -18.3416
pf_sweep | 9 | 1 | 8 | 2 | 120 | 1 | -0.2 | 60 | 119.9804 | 1.0002 | -19.9225 | 0.023909 | 0.1176 | 0.12001 | 119.77 | 1 | 0.020342 | -0.2104 | -0.000229 | -0.003567 | -0.1754 | -0.0228 | -14.9175
pf_sweep | 9 | 1 | 9 | 2 | 120 | 1 | -0.2 | 60 | 119.9804 | 1.0002 | -19.9225 | 0.023909 | 0.1176 | 0.12001 | 119.77 | 1.001 | 0.020507 | -0.2104 | 0.000771 | -0.003402 | -0.1754 | 0.0771 | -14.2273
pf_sweep | 9 | 1 | 10 | 2 | 120 | 1 | -0.2 | 60 | 119.9804 | 1.0002 | -19.9225 | 0.023909 | 0.1176 | 0.12001 | 119.77 | 1 | 0.020465 | -0.2104 | -0.000229 | -0.003444 | -0.1754 | -0.0228 | -14.403
pf_sweep | 9 | 1 | 11 | 2 | 120 | 1 | -0.2 | 60 | 119.9804 | 1.0002 | -19.9225 | 0.023909 | 0.1176 | 0.12001 | 119.77 | 1 | 0.020706 | -0.2104 | -0.000229 | -0.003203 | -0.1754 | -0.0228 | -13.395
pf_sweep | 9 | 1 | 12 | 2 | 120 | 1 | -0.2 | 60 | 119.9804 | 1.0002 | -19.9225 | 0.023909 | 0.1176 | 0.12001 | 119.77 | 1 | 0.020335 | -0.2104 | -0.000229 | -0.003574 | -0.1754 | -0.0228 | -14.9467
pf_sweep | 9 | 1 | 13 | 3 | 120 | 1 | -0.2 | 60 | 119.9906 | 0.99991 | -19.9836 | 0.023976 | 0.11756 | 0.11998 | 119.774 | 0.999 | 0.020634 | -0.21665 | -0.000905 | -0.003342 | -0.1806 | -0.0905 | -13.9396
pf_sweep | 9 | 1 | 14 | 3 | 120 | 1 | -0.2 | 60 | 119.9906 | 0.99991 | -19.9836 | 0.023976 | 0.11756 | 0.11998 | 119.774 | 0.999 | 0.020859 | -0.21665 | -0.000905 | -0.003117 | -0.1806 | -0.0905 | -13.0012
pf_sweep | 9 | 1 | 15 | 3 | 120 | 1 | -0.2 | 60 | 119.9906 | 0.99991 | -19.9836 | 0.023976 | 0.11756 | 0.11998 | 119.774 | 1 | 0.02075 | -0.21665 | 9.5e-05 | -0.003226 | -0.1806 | 0.0095 | -13.4558
pf_sweep | 9 | 1 | 16 | 3 | 120 | 1 | -0.2 | 60 | 119.9906 | 0.99991 | -19.9836 | 0.023976 | 0.11756 | 0.11998 | 119.774 | 0.999 | 0.0209 | -0.21665 | -0.000905 | -0.003076 | -0.1806 | -0.0905 | -12.8302
pf_sweep | 9 | 1 | 17 | 3 | 120 | 1 | -0.2 | 60 | 119.9906 | 0.99991 | -19.9836 | 0.023976 | 0.11756 | 0.11998 | 119.774 | 1 | 0.020756 | -0.21665 | 9.5e-05 | -0.00322 | -0.1806 | 0.0095 | -13.4308
pf_sweep | 9 | 1 | 18 | 3 | 120 | 1 | -0.2 | 60 | 119.9906 | 0.99991 | -19.9836 | 0.023976 | 0.11756 | 0.11998 | 119.774 | 1 | 0.02084 | -0.21665 | 9.5e-05 | -0.003136 | -0.1806 | 0.0095 | -13.0804
pf_sweep | 9 | 2 | 1 | 1 | 120 | 1 | -0.2 | 60 | 120.0252 | 0.99996 | -19.9781 | 0.023978 | 0.1176 | 0.12002 | 119.765 | 2.968 | 0.058761 | -0.26017 | 1.968 | 0.034783 | -0.2168 | -1.0628 | -18.3118
pf_sweep | 9 | 2 | 2 | 1 | 120 | 1 | -0.2 | 60 | 120.0252 | 0.99996 | -19.9781 | 0.023978 | 0.1176 | 0.12002 | 119.765 | 1.001 | 0.020456 | -0.26017 | 0.001039 | -0.003522 | -0.2168 | 0.1039 | -14.6876
pf_sweep | 9 | 2 | 3 | 1 | 120 | 1 | -0.2 | 60 | 120.0252 | 0.99996 | -19.9781 | 0.023978 | 0.1176 | 0.12002 | 119.765 | 1.001 | 0.020445 | -0.26017 | 0.001039 | -0.003533 | -0.2168 | 0.1039 | -14.7335
pf_sweep | 9 | 2 | 4 | 1 | 120 | 1 | -0.2 | 60 | 120.0252 | 0.99996 | -19.9781 | 0.023978 | 0.1176 | 0.12002 | 119.765 | 1.001 | 0.020436 | -0.26017 | 0.001039 | -0.003542 | -0.2168 | 0.1039 | -14.771
pf_sweep | 9 | 2 | 5 | 1 | 120 | 1 | -0.2 | 60 | 120.0252 | 0.99996 | -19.9781 | 0.023978 | 0.1176 | 0.12002 | 119.765 | 1.001 | 0.020539 | -0.26017 | 0.001039 | -0.003439 | -0.2168 | 0.1039 | -14.3415
pf_sweep | 9 | 2 | 6 | 1 | 120 | 1 | -0.2 | 60 | 120.0252 | 0.99996 | -19.9781 | 0.023978 | 0.1176 | 0.12002 | 119.765 | 0.998 | 0.020955 | -0.26017 | -0.001961 | -0.003023 | -0.2168 | -0.1961 | -12.6065
pf_sweep | 9 | 2 | 7 | 2 | 120 | 1 | -0.2 | 60 | 119.9791 | 1.0002 | -19.9213 | 0.023906 | 0.1176 | 0.12 | 119.768 | 2.978 | 0.058524 | -0.21106 | 1.9778 | 0.034618 | -0.1759 | -0.75432 | -18.3984
pf_sweep | 9 | 2 | 8 | 2 | 120 | 1 | -0.2 | 60 | 119.9791 | 1.0002 | -19.9213 | 0.023906 | 0.1176 | 0.12 | 119.768 | 1 | 0.020344 | -0.21106 | -0.000211 | -0.003562 | -0.1759 | -0.0211 | -14.9015
pf_sweep | 9 | 2 | 9 | 2 | 120 | 1 | -0.2 | 60 | 119.9791 | 1.0002 | -19.9213 | 0.023906 | 0.1176 | 0.12 | 119.768 | 1.001 | 0.020535 | -0.21106 | 0.000789 | -0.003371 | -0.1759 | 0.0788 | -14.1025
pf_sweep | 9 | 2 | 10 | 2 | 120 | 1 | -0.2 | 60 | 119.9791 | 1.0002 | -19.9213 | 0.023906 | 0.1176 | 0.12 | 119.768 | 1 | 0.020456 | -0.21106 | -0.000211 | -0.00345 | -0.1759 | -0.0211 | -14.433
pf_sweep | 9 | 2 | 11 | 2 | 120 | 1 | -0.2 | 60 | 119.9791 | 1.0002 | -19.9213 | 0.023906 | 0.1176 | 0.12 | 119.768 | 1 | 0.020689 | -0.21106 | -0.000211 | -0.003217 | -0.1759 | -0.0211 | -13.4583
pf_sweep | 9 | 2 | 12 | 2 | 120 | 1 | -0.2 | 60 | 119.9791 | 1.0002 | -19.9213 | 0.023906 | 0.1176 | 0.12 | 119.768 | 1 | 0.020344 | -0.21106 | -0.000211 | -0.003562 | -0.1759 | -0.0211 | -14.9015
pf_sweep | 9 | 2 | 13 | 3 | 120 | 1 | -0.2 | 60 | 119.9866 | 0.99992 | -19.9867 | 0.023979 | 0.11756 | 0.11998 | 119.761 | 1 | 0.020625 | -0.22563 | 8.4e-05 | -0.003354 | -0.188 | 0.0084 | -13.9883
pf_sweep | 9 | 2 | 14 | 3 | 120 | 1 | -0.2 | 60 | 119.9866 | 0.99992 | -19.9867 | 0.023979 | 0.11756 | 0.11998 | 119.761 | 0.999 | 0.020848 | -0.22563 | -0.000916 | -0.003131 | -0.188 | -0.0916 | -13.0584
pf_sweep | 9 | 2 | 15 | 3 | 120 | 1 | -0.2 | 60 | 119.9866 | 0.99992 | -19.9867 | 0.023979 | 0.11756 | 0.11998 | 119.761 | 1 | 0.020748 | -0.22563 | 8.4e-05 | -0.003231 | -0.188 | 0.0084 | -13.4754
pf_sweep | 9 | 2 | 16 | 3 | 120 | 1 | -0.2 | 60 | 119.9866 | 0.99992 | -19.9867 | 0.023979 | 0.11756 | 0.11998 | 119.761 | 0.999 | 0.020902 | -0.22563 | -0.000916 | -0.003077 | -0.188 | -0.0916 | -12.8332
pf_sweep | 9 | 2 | 17 | 3 | 120 | 1 | -0.2 | 60 | 119.9866 | 0.99992 | -19.9867 | 0.023979 | 0.11756 | 0.11998 | 119.761 | 1 | 0.020764 | -0.22563 | 8.4e-05 | -0.003215 | -0.188 | 0.0084 | -13.4087
pf_sweep | 9 | 2 | 18 | 3 | 120 | 1 | -0.2 | 60 | 119.9866 | 0.99992 | -19.9867 | 0.023979 | 0.11756 | 0.11998 | 119.761 | 1 | 0.020828 | -0.22563 | 8.4e-05 | -0.003151 | -0.188 | 0.0084 | -13.1418
pf_sweep | 9 | 3 | 1 | 1 | 120 | 1 | -0.2 | 60 | 120.0114 | 0.99994 | -19.9855 | 0.023984 | 0.11758 | 0.12 | 119.759 | 2.968 | 0.058742 | -0.25239 | 1.9681 | 0.034758 | -0.2103 | -1.0611 | -18.3578
pf_sweep | 9 | 3 | 2 | 1 | 120 | 1 | -0.2 | 60 | 120.0114 | 0.99994 | -19.9855 | 0.023984 | 0.11758 | 0.12 | 119.759 | 1.001 | 0.020478 | -0.25239 | 0.001056 | -0.003506 | -0.2103 | 0.1056 | -14.6164
pf_sweep | 9 | 3 | 3 | 1 | 120 | 1 | -0.2 | 60 | 120.0114 | 0.99994 | -19.9855 | 0.023984 | 0.11758 | 0.12 | 119.759 | 1.001 | 0.020443 | -0.25239 | 0.001056 | -0.003541 | -0.2103 | 0.1056 | -14.7623
pf_sweep | 9 | 3 | 4 | 1 | 120 | 1 | -0.2 | 60 | 120.0114 | 0.99994 | -19.9855 | 0.023984 | 0.11758 | 0.12 | 119.759 | 1.001 | 0.020428 | -0.25239 | 0.001056 | -0.003556 | -0.2103 | 0.1056 | -14.8248
pf_sweep | 9 | 3 | 5 | 1 | 120 | 1 | -0.2 | 60 | 120.0114 | 0.99994 | -19.9855 | 0.023984 | 0.11758 | 0.12 | 119.759 | 1.001 | 0.020531 | -0.25239 | 0.001056 | -0.003453 | -0.2103 | 0.1056 | -14.3954
pf_sweep | 9 | 3 | 6 | 1 | 120 | 1 | -0.2 | 60 | 120.0114 | 0.99994 | -19.9855 | 0.023984 | 0.11758 | 0.12 | 119.759 | 0.998 | 0.020964 | -0.25239 | -0.001944 | -0.00302 | -0.2103 | -0.1944 | -12.59
pf_sweep | 9 | 3 | 7 | 2 | 120 | 1 | -0.2 | 60 | 119.9822 | 1.0002 | -19.9213 | 0.023907 | 0.1176 | 0.12001 | 119.765 | 2.979 | 0.058524 | -0.21722 | 1.9788 | 0.034617 | -0.181 | -0.72227 | -18.4016
pf_sweep | 9 | 3 | 8 | 2 | 120 | 1 | -0.2 | 60 | 119.9822 | 1.0002 | -19.9213 | 0.023907 | 0.1176 | 0.12001 | 119.765 | 1 | 0.020363 | -0.21722 | -0.000224 | -0.003544 | -0.181 | -0.0224 | -14.8253
pf_sweep | 9 | 3 | 9 | 2 | 120 | 1 | -0.2 | 60 | 119.9822 | 1.0002 | -19.9213 | 0.023907 | 0.1176 | 0.12001 | 119.765 | 1.001 | 0.020535 | -0.21722 | 0.000776 | -0.003372 | -0.181 | 0.0775 | -14.1058
pf_sweep | 9 | 3 | 10 | 2 | 120 | 1 | -0.2 | 60 | 119.9822 | 1.0002 | -19.9213 | 0.023907 | 0.1176 | 0.12001 | 119.765 | 1 | 0.020466 | -0.21722 | -0.000224 | -0.003441 | -0.181 | -0.0224 | -14.3944
pf_sweep | 9 | 3 | 11 | 2 | 120 | 1 | -0.2 | 60 | 119.9822 | 1.0002 | -19.9213 | 0.023907 | 0.1176 | 0.12001 | 119.765 | 1 | 0.020718 | -0.21722 | -0.000224 | -0.003189 | -0.181 | -0.0224 | -13.3404
pf_sweep | 9 | 3 | 12 | 2 | 120 | 1 | -0.2 | 60 | 119.9822 | 1.0002 | -19.9213 | 0.023907 | 0.1176 | 0.12001 | 119.765 | 1 | 0.020353 | -0.21722 | -0.000224 | -0.003554 | -0.181 | -0.0224 | -14.8671
pf_sweep | 9 | 3 | 13 | 3 | 120 | 1 | -0.2 | 60 | 119.9875 | 0.99991 | -19.9833 | 0.023975 | 0.11756 | 0.11998 | 119.766 | 1 | 0.020654 | -0.22148 | 8.6e-05 | -0.003321 | -0.1846 | 0.0086 | -13.8536
pf_sweep | 9 | 3 | 14 | 3 | 120 | 1 | -0.2 | 60 | 119.9875 | 0.99991 | -19.9833 | 0.023975 | 0.11756 | 0.11998 | 119.766 | 0.999 | 0.020847 | -0.22148 | -0.000914 | -0.003128 | -0.1846 | -0.0914 | -13.0486
pf_sweep | 9 | 3 | 15 | 3 | 120 | 1 | -0.2 | 60 | 119.9875 | 0.99991 | -19.9833 | 0.023975 | 0.11756 | 0.11998 | 119.766 | 1 | 0.020748 | -0.22148 | 8.6e-05 | -0.003227 | -0.1846 | 0.0086 | -13.4615
pf_sweep | 9 | 3 | 16 | 3 | 120 | 1 | -0.2 | 60 | 119.9875 | 0.99991 | -19.9833 | 0.023975 | 0.11756 | 0.11998 | 119.766 | 0.999 | 0.020931 | -0.22148 | -0.000914 | -0.003044 | -0.1846 | -0.0914 | -12.6982
pf_sweep | 9 | 3 | 17 | 3 | 120 | 1 | -0.2 | 60 | 119.9875 | 0.99991 | -19.9833 | 0.023975 | 0.11756 | 0.11998 | 119.766 | 1 | 0.020764 | -0.22148 | 8.6e-05 | -0.003211 | -0.1846 | 0.0086 | -13.3948
pf_sweep | 9 | 3 | 18 | 3 | 120 | 1 | -0.2 | 60 | 119.9875 | 0.99991 | -19.9833 | 0.023975 | 0.11756 | 0.11998 | 119.766 | 1 | 0.02082 | -0.22148 | 8.6e-05 | -0.003155 | -0.1846 | 0.0086 | -13.1612
pf_sweep | 10 | 1 | 1 | 1 | 120 | 1 | -0.1 | 60 | 120.0101 | 0.99996 | -9.9832 | 0.01198 | 0.11941 | 0.12001 | 119.722 | 0 | 0 | -0.28809 | -0.99996 | -0.01198 | -0.2401 | -100 | -100
pf_sweep | 10 | 1 | 2 | 1 | 120 | 1 | -0.1 | 60 | 120.0101 | 0.99996 | -9.9832 | 0.01198 | 0.11941 | 0.12001 | 119.722 | 0 | 0 | -0.28809 | -0.99996 | -0.01198 | -0.2401 | -100 | -100
pf_sweep | 10 | 1 | 3 | 1 | 120 | 1 | -0.1 | 60 | 120.0101 | 0.99996 | -9.9832 | 0.01198 | 0.11941 | 0.12001 | 119.722 | 0 | 0 | -0.28809 | -0.99996 | -0.01198 | -0.2401 | -100 | -100
pf_sweep | 10 | 1 | 4 | 1 | 120 | 1 | -0.1 | 60 | 120.0101 | 0.99996 | -9.9832 | 0.01198 | 0.11941 | 0.12001 | 119.722 | 0 | 0 | -0.28809 | -0.99996 | -0.01198 | -0.2401 | -100 | -100
pf_sweep | 10 | 1 | 5 | 1 | 120 | 1 | -0.1 | 60 | 120.0101 | 0.99996 | -9.9832 | 0.01198 | 0.11941 | 0.12001 | 119.722 | 0 | 0 | -0.28809 | -0.99996 | -0.01198 | -0.2401 | -100 | -100
pf_sweep | 10 | 1 | 6 | 1 | 120 | 1 | -0.1 | 60 | 120.0101 | 0.99996 | -9.9832 | 0.01198 | 0.11941 | 0.12001 | 119.722 | 0 | 0 | -0.28809 | -0.99996 | -0.01198 | -0.2401 | -100 | -100
pf_sweep | 10 | 1 | 7 | 2 | 120 | 1 | -0.1 | 60 | 119.9837 | 1.0002 | -9.9199 | 0.011905 | 0.11942 | 0.12001 | 119.765 | 0 | 0 | -0.21866 | -1.0002 | -0.011905 | -0.1822 | -100 | -100
pf_sweep | 10 | 1 | 8 | 2 | 120 | 1 | -0.1 | 60 | 119.9837 | 1.0002 | -9.9199 | 0.011905 | 0.11942 | 0.12001 | 119.765 | 0 | 0 | -0.21866 | -1.0002 | -0.011905 | -0.1822 | -100 | -100
pf_sweep | 10 | 1 | 9 | 2 | 120 | 1 | -0.1 | 60 | 119.9837 | 1.0002 | -9.9199 | 0.011905 | 0.11942 | 0.12001 | 119.765 | 0 | 0 | -0.21866 | -1.0002 | -0.011905 | -0.1822 | -100 | -100
pf_sweep | 10 | 1 | 10 | 2 | 120 | 1 | -0.1 | 60 | 119.9837 | 1.0002 | -9.9199 | 0.011905 | 0.11942 | 0.12001 | 119.765 | 0 | 0 | -0.21866 | -1.0002 | -0.011905 | -0.1822 | -100 | -100
pf_sweep | 10 | 1 | 11 | 2 | 120 | 1 | -0.1 | 60 | 119.9837 | 1.0002 | -9.9199 | 0.011905 | 0.11942 | 0.12001 | 119.765 | 0 | 0 | -0.21866 | -1.0002 | -0.011905 | -0.1822 | -100 | -100
pf_sweep | 10 | 1 | 12 | 2 | 120 | 1 | -0.1 | 60 | 119.9837 | 1.0002 | -9.9199 | 0.011905 | 0.11942 | 0.12001 | 119.765 | 0 | 0 | -0.21866 | -1.0002 | -0.011905 | -0.1822 | -100 | -100
pf_sweep | 10 | 1 | 13 | 3 | 120 | 1 | -0.1 | 60 | 119.9876 | 0.99992 | -9.9829 | 0.011977 | 0.11938 | 0.11998 | 119.767 | 0 | 0 | -0.22058 | -0.99992 | -0.011977 | -0.1838 | -100 | -100
pf_sweep | 10 | 1 | 14 | 3 | 120 | 1 | -0.1 | 60 | 119.9876 | 0.99992 | -9.9829 | 0.011977 | 0.11938 | 0.11998 | 119.767 | 0 | 0 | -0.22058 | -0.99992 | -0.011977 | -0.1838 | -100 | -100
pf_sweep | 10 | 1 | 15 | 3 | 120 | 1 | -0.1 | 60 | 119.9876 | 0.99992 | -9.9829 | 0.011977 | 0.11938 | 0.11998 | 119.767 | 0 | 0 | -0.22058 | -0.99992 | -0.011977 | -0.1838 | -100 | -100
pf_sweep | 10 | 1 | 16 | 3 | 120 | 1 | -0.1 | 60 | 119.9876 | 0.99992 | -9.9829 | 0.011977 | 0.11938 | 0.11998 | 119.767 | 0 | 0 | -0.22058 | -0.99992 | -0.011977 | -0.1838 | -100 | -100
pf_sweep | 10 | 1 | 17 | 3 | 120 | 1 | -0.1 | 60 | 119.9876 | 0.99992 | -9.9829 | 0.011977 | 0.11938 | 0.11998 | 119.767 | 0 | 0 | -0.22058 | -0.99992 | -0.011977 | -0.1838 | -100 | -100
pf_sweep | 10 | 1 | 18 | 3 | 120 | 1 | -0.1 | 60 | 119.9876 | 0.99992 | -9.9829 | 0.011977 | 0.11938 | 0.11998 | 119.767 | 0 | 0 | -0.22058 | -0.99992 | -0.011977 | -0.1838 | -100 | -100
pf_sweep | 10 | 2 | 1 | 1 | 120 | 1 | -0.1 | 60 | 120.0133 | 0.99995 | -9.9796 | 0.011976 | 0.11941 | 0.12001 | 119.727 | 0 | 0 | -0.2863 | -0.99994 | -0.011976 | -0.2386 | -100 | -100
pf_sweep | 10 | 2 | 2 | 1 | 120 | 1 | -0.1 | 60 | 120.0133 | 0.99995 | -9.9796 | 0.011976 | 0.11941 | 0.12001 | 119.727 | 0 | 0 | -0.2863 | -0.99994 | -0.011976 | -0.2386 | -100 | -100
pf_sweep | 10 | 2 | 3 | 1 | 120 | 1 | -0.1 | 60 | 120.0133 | 0.99995 | -9.9796 | 0.011976 | 0.11941 | 0.12001 | 119.727 | 0 | 0 | -0.2863 | -0.99994 | -0.011976 | -0.2386 | -100 | -100
pf_sweep | 10 | 2 | 4 | 1 | 120 | 1 | -0.1 | 60 | 120.0133 | 0.99995 | -9.9796 | 0.011976 | 0.11941 | 0.12001 | 119.727 | 0 | 0 | -0.2863 | -0.99994 | -0.011976 | -0.2386 | -100 | -100
pf_sweep | 10 | 2 | 5 | 1 | 120 | 1 | -0.1 | 60 | 120.0133 | 0.99995 | -9.9796 | 0.011976 | 0.11941 | 0.12001 | 119.727 | 0 | 0 | -0.2863 | -0.99994 | -0.011976 | -0.2386 | -100 | -100
pf_sweep | 10 | 2 | 6 | 1 | 120 | 1 | -0.1 | 60 | 120.0133 | 0.99995 | -9.9796 | 0.011976 | 0.11941 | 0.12001 | 119.727 | 0 | 0 | -0.2863 | -0.99994 | -0.011976 | -0.2386 | -100 | -100
pf_sweep | 10 | 2 | 7 | 2 | 120 | 1 | -0.1 | 60 | 119.9818 | 1.0002 | -9.9265 | 0.011913 | 0.11942 | 0.12001 | 119.773 | 0 | 0 | -0.20879 | -1.0002 | -0.011913 | -0.174 | -100 | -100
pf_sweep | 10 | 2 | 8 | 2 | 120 | 1 | -0.1 | 60 | 119.9818 | 1.0002 | -9.9265 | 0.011913 | 0.11942 | 0.12001 | 119.773 | 0 | 0 | -0.20879 | -1.0002 | -0.011913 | -0.174 | -100 | -100
pf_sweep | 10 | 2 | 9 | 2 | 120 | 1 | -0.1 | 60 | 119.9818 | 1.0002 | -9.9265 | 0.011913 | 0.11942 | 0.12001 | 119.773 | 0 | 0 | -0.20879 | -1.0002 | -0.011913 | -0.174 | -100 | -100
pf_sweep | 10 | 2 | 10 | 2 | 120 | 1 | -0.1 | 60 | 119.9818 | 1.0002 | -9.9265 | 0.011913 | 0.11942 | 0.12001 | 119.773 | 0 | 0 | -0.20879 | -1.0002 | -0.011913 | -0.174 | -100 | -100
pf_sweep | 10 | 2 | 11 | 2 | 120 | 1 | -0.1 | 60 | 119.9818 | 1.0002 | -9.9265 | 0.011913 | 0.11942 | 0.12001 | 119.773 | 0 | 0 | -0.20879 | -1.0002 | -0.011913 | -0.174 | -100 | -100
pf_sweep | 10 | 2 | 12 | 2 | 120 | 1 | -0.1 | 60 | 119.9818 | 1.0002 | -9.9265 | 0.011913 | 0.11942 | 0.12001 | 119.773 | 0 | 0 | -0.20879 | -1.0002 | -0.011913 | -0.174 | -100 | -100
pf_sweep | 10 | 2 | 13 | 3 | 120 | 1 | -0.1 | 60 | 119.9909 | 0.99993 | -9.9881 | 0.011984 | 0.11938 | 0.11998 | 119.776 | 0 | 0 | -0.21494 | -0.99993 | -0.011984 | -0.1791 | -100 | -100
pf_sweep | 10 | 2 | 14 | 3 | 120 | 1 | -0.1 | 60 | 119.9909 | 0.99993 | -9.9881 | 0.011984 | 0.11938 | 0.11998 | 119.776 | 0 | 0 | -0.21494 | -0.99993 | -0.011984 | -0.1791 | -100 | -100
pf_sweep | 10 | 2 | 15 | 3 | 120 | 1 | -0.1 | 60 | 119.9909 | 0.99993 | -9.9881 | 0.011984 | 0.11938 | 0.11998 | 119.776 | 0 | 0 | -0.21494 | -0.99993 | -0.011984 | -0.1791 | -100 | -100
pf_sweep | 10 | 2 | 16 | 3 | 120 | 1 | -0.1 | 60 | 119.9909 | 0.99993 | -9.9881 | 0.011984 | 0.11938 | 0.11998 | 119.776 | 0 | 0 | -0.21494 | -0.99993 | -0.011984 | -0.1791 | -100 | -100
pf_sweep | 10 | 2 | 17 | 3 | 120 | 1 | -0.1 | 60 | 119.9909 | 0.99993 | -9.9881 | 0.011984 | 0.11938 | 0.11998 | 119.776 | 0 | 0 | -0.21494 | -0.99993 | -0.011984 | -0.1791 | -100 | -100
pf_sweep | 10 | 2 | 18 | 3 | 120 | 1 | -0.1 | 60 | 119.9909 | 0.99993 | -9.9881 | 0.011984 | 0.11938 | 0.11998 | 119.776 | 0 | 0 | -0.21494 | -0.99993 | -0.011984 | -0.1791 | -100 | -100
pf_sweep | 10 | 3 | 1 | 1 | 120 | 1 | -0.1 | 60 | 120.0254 | 0.99997 | -9.9798 | 0.011978 | 0.11942 | 0.12002 | 119.697 | 0 | 0 | -0.3284 | -0.99997 | -0.011978 | -0.2736 | -100 | -100
pf_sweep | 10 | 3 | 2 | 1 | 120 | 1 | -0.1 | 60 | 120.0254 | 0.99997 | -9.9798 | 0.011978 | 0.11942 | 0.12002 | 119.697 | 0 | 0 | -0.3284 | -0.99997 | -0.011978 | -0.2736 | -100 | -100
pf_sweep | 10 | 3 | 3 | 1 | 120 | 1 | -0.1 | 60 | 120.0254 | 0.99997 | -9.9798 | 0.011978 | 0.11942 | 0.12002 | 119.697 | 0 | 0 | -0.3284 | -0.99997 | -0.011978 | -0.2736 | -100 | -100
pf_sweep | 10 | 3 | 4 | 1 | 120 | 1 | -0.1 | 60 | 120.0254 | 0.99997 | -9.9798 | 0.011978 | 0.11942 | 0.12002 | 119.697 | 0 | 0 | -0.3284 | -0.99997 | -0.011978 | -0.2736 | -100 | -100
pf_sweep | 10 | 3 | 5 | 1 | 120 | 1 | -0.1 | 60 | 120.0254 | 0.99997 | -9.9798 | 0.011978 | 0.11942 | 0.12002 | 119.697 | 0 | 0 | -0.3284 | -0.99997 | -0.011978 | -0.2736 | -100 | -100
pf_sweep | 10 | 3 | 6 | 1 | 120 | 1 | -0.1 | 60 | 120.0254 | 0.99997 | -9.9798 | 0.011978 | 0.11942 | 0.12002 | 119.697 | 0 | 0 | -0.3284 | -0.99997 | -0.011978 | -0.2736 | -100 | -100
pf_sweep | 10 | 3 | 7 | 2 | 120 | 1 | -0.1 | 60 | 119.9806 | 1.0002 | -9.9195 | 0.011904 | 0.11942 | 0.12001 | 119.769 | 0 | 0 | -0.21158 | -1.0002 | -0.011904 | -0.1763 | -100 | -100
pf_sweep | 10 | 3 | 8 | 2 | 120 | 1 | -0.1 | 60 | 119.9806 | 1.0002 | -9.9195 | 0.011904 | 0.11942 | 0.12001 | 119.769 | 0 | 0 | -0.21158 | -1.0002 | -0.011904 | -0.1763 | -100 | -100
pf_sweep | 10 | 3 | 9 | 2 | 120 | 1 | -0.1 | 60 | 119.9806 | 1.0002 | -9.9195 | 0.011904 | 0.11942 | 0.12001 | 119.769 | 0 | 0 | -0.21158 | -1.0002 | -0.011904 | -0.1763 | -100 | -100
pf_sweep | 10 | 3 | 10 | 2 | 120 | 1 | -0.1 | 60 | 119.9806 | 1.0002 | -9.9195 | 0.011904 | 0.11942 | 0.12001 | 119.769 | 0 | 0 | -0.21158 | -1.0002 | -0.011904 | -0.1763 | -100 | -100
pf_sweep | 10 | 3 | 11 | 2 | 120 | 1 | -0.1 | 60 | 119.9806 | 1.0002 | -9.9195 | 0.011904 | 0.11942 | 0.12001 | 119.769 | 0 | 0 | -0.21158 | -1.0002 | -0.011904 | -0.1763 | -100 | -100
pf_sweep | 10 | 3 | 12 | 2 | 120 | 1 | -0.1 | 60 | 119.9806 | 1.0002 | -9.9195 | 0.011904 | 0.11942 | 0.12001 | 119.769 | 0 | 0 | -0.21158 | -1.0002 | -0.011904 | -0.1763 | -100 | -100
pf_sweep | 10 | 3 | 13 | 3 | 120 | 1 | -0.1 | 60 | 119.9877 | 0.99992 | -9.986 | 0.011981 | 0.11938 | 0.11998 | 119.775 | 0 | 0 | -0.21269 | -0.99992 | -0.011981 | -0.1773 | -100 | -100
pf_sweep | 10 | 3 | 14 | 3 | 120 | 1 | -0.1 | 60 | 119.9877 | 0.99992 | -9.986 | 0.011981 | 0.11938 | 0.11998 | 119.775 | 0 | 0 | -0.21269 | -0.99992 | -0.011981 | -0.1773 | -100 | -100
pf_sweep | 10 | 3 | 15 | 3 | 120 | 1 | -0.1 | 60 | 119.9877 | 0.99992 | -9.986 | 0.011981 | 0.11938 | 0.11998 | 119.775 | 0 | 0 | -0.21269 | -0.99992 | -0.011981 | -0.1773 | -100 | -100
pf_sweep | 10 | 3 | 16 | 3 | 120 | 1 | -0.1 | 60 | 119.9877 | 0.99992 | -9.986 | 0.011981 | 0.11938 | 0.11998 | 119.775 | 0 | 0 | -0.21269 | -0.99992 | -0.011981 | -0.1773 | -100 | -100
pf_sweep | 10 | 3 | 17 | 3 | 120 | 1 | -0.1 | 60 | 119.9877 | 0.99992 | -9.986 | 0.011981 | 0.11938 | 0.11998 | 119.775 | 0 | 0 | -0.21269 | -0.99992 | -0.011981 | -0.1773 | -100 | -100
pf_sweep | 10 | 3 | 18 | 3 | 120 | 1 | -0.1 | 60 | 119.9877 | 0.99992 | -9.986 | 0.011981 | 0.11938 | 0.11998 | 119.775 | 0 | 0 | -0.21269 | -0.99992 | -0.011981 | -0.1773 | -100 | -100
pf_sweep | 11 | 1 | 1 | 1 | 120 | 1 | 0.1 | 60 | 119.99 | 0.99997 | 10.0177 | 0.01202 | -0.11938 | 0.11999 | 119.71 | 2.968 | 0.048427 | -0.28001 | 1.968 | 0.036407 | -0.2334 | -1.0637 | 34.2969
pf_sweep | 11 | 1 | 2 | 1 | 120 | 1 | 0.1 | 60 | 119.99 | 0.99997 | 10.0177 | 0.01202 | -0.11938 | 0.11999 | 119.71 | 1.001 | 0.015669 | -0.28001 | 0.00103 | 0.003649 | -0.2334 | 0.103 | 30.359
pf_sweep | 11 | 1 | 3 | 1 | 120 | 1 | 0.1 | 60 | 119.99 | 0.99997 | 10.0177 | 0.01202 | -0.11938 | 0.11999 | 119.71 | 1.001 | 0.015709 | -0.28001 | 0.00103 | 0.003689 | -0.2334 | 0.103 | 30.6917
pf_sweep | 11 | 1 | 4 | 1 | 120 | 1 | 0.1 | 60 | 119.99 | 0.99997 | 10.0177 | 0.01202 | -0.11938 | 0.11999 | 119.71 | 1 | 0.015714 | -0.28001 | 3e-05 | 0.003694 | -0.2334 | 0.003 | 30.7333
pf_sweep | 11 | 1 | 5 | 1 | 120 | 1 | 0.1 | 60 | 119.99 | 0.99997 | 10.0177 | 0.01202 | -0.11938 | 0.11999 | 119.71 | 1.001 | 0.015594 | -0.28001 | 0.00103 | 0.003574 | -0.2334 | 0.103 | 29.735
pf_sweep | 11 | 1 | 6 | 1 | 120 | 1 | 0.1 | 60 | 119.99 | 0.99997 | 10.0177 | 0.01202 | -0.11938 | 0.11999 | 119.71 | 0.998 | 0.015076 | -0.28001 | -0.00197 | 0.003056 | -0.2334 | -0.197 | 25.4255
pf_sweep | 11 | 1 | 7 | 2 | 120 | 1 | 0.1 | 60 | 119.9805 | 1.0001 | 10.0748 | 0.012089 | -0.11938 | 0.11999 | 119.773 | 2.977 | 0.049229 | -0.20752 | 1.9769 | 0.03714 | -0.173 | -0.77639 | 35.7406
pf_sweep | 11 | 1 | 8 | 2 | 120 | 1 | 0.1 | 60 | 119.9805 | 1.0001 | 10.0748 | 0.012089 | -0.11938 | 0.11999 | 119.773 | 1.001 | 0.015605 | -0.20752 | 0.000902 | 0.003516 | -0.173 | 0.0902 | 29.0844
pf_sweep | 11 | 1 | 9 | 2 | 120 | 1 | 0.1 | 60 | 119.9805 | 1.0001 | 10.0748 | 0.012089 | -0.11938 | 0.11999 | 119.773 | 1.001 | 0.015673 | -0.20752 | 0.000902 | 0.003584 | -0.173 | 0.0902 | 29.6469
pf_sweep | 11 | 1 | 10 | 2 | 120 | 1 | 0.1 | 60 | 119.9805 | 1.0001 | 10.0748 | 0.012089 | -0.11938 | 0.11999 | 119.773 | 1 | 0.015713 | -0.20752 | -9.8e-05 | 0.003624 | -0.173 | -0.0098 | 29.9778
pf_sweep | 11 | 1 | 11 | 2 | 120 | 1 | 0.1 | 60 | 119.9805 | 1.0001 | 10.0748 | 0.012089 | -0.11938 | 0.11999 | 119.773 | 1 | 0.015473 | -0.20752 | -9.8e-05 | 0.003384 | -0.173 | -0.0098 | 27.9925
pf_sweep | 11 | 1 | 12 | 2 | 120 | 1 | 0.1 | 60 | 119.9805 | 1.0001 | 10.0748 | 0.012089 | -0.11938 | 0.11999 | 119.773 | 1 | 0.015851 | -0.20752 | -9.8e-05 | 0.003762 | -0.173 | -0.0098 | 31.1193
pf_sweep | 11 | 1 | 13 | 3 | 120 | 1 | 0.1 | 60 | 119.9939 | 0.99991 | 10.0154 | 0.012017 | -0.11938 | 0.11998 | 119.769 | 1 | 0.015491 | -0.2249 | 8.8e-05 | 0.003474 | -0.1874 | 0.0088 | 28.9116
pf_sweep | 11 | 1 | 14 | 3 | 120 | 1 | 0.1 | 60 | 119.9939 | 0.99991 | 10.0154 | 0.012017 | -0.11938 | 0.11998 | 119.769 | 0.999 | 0.015297 | -0.2249 | -0.000912 | 0.00328 | -0.1874 | -0.0912 | 27.2972
pf_sweep | 11 | 1 | 15 | 3 | 120 | 1 | 0.1 | 60 | 119.9939 | 0.99991 | 10.0154 | 0.012017 | -0.11938 | 0.11998 | 119.769 | 1 | 0.01543 | -0.2249 | 8.8e-05 | 0.003413 | -0.1874 | 0.0088 | 28.404
pf_sweep | 11 | 1 | 16 | 3 | 120 | 1 | 0.1 | 60 | 119.9939 | 0.99991 | 10.0154 | 0.012017 | -0.11938 | 0.11998 | 119.769 | 0.999 | 0.015195 | -0.2249 | -0.000912 | 0.003178 | -0.1874 | -0.0912 | 26.4484
pf_sweep | 11 | 1 | 17 | 3 | 120 | 1 | 0.1 | 60 | 119.9939 | 0.99991 | 10.0154 | 0.012017 | -0.11938 | 0.11998 | 119.769 | 1 | 0.015395 | -0.2249 | 8.8e-05 | 0.003378 | -0.1874 | 0.0088 | 28.1127
pf_sweep | 11 | 1 | 18 | 3 | 120 | 1 | 0.1 | 60 | 119.9939 | 0.99991 | 10.0154 | 0.012017 | -0.11938 | 0.11998 | 119.769 | 1 | 0.015291 | -0.2249 | 8.8e-05 | 0.003274 | -0.1874 | 0.0088 | 27.2473
pf_sweep | 11 | 2 | 1 | 1 | 120 | 1 | 0.1 | 60 | 120.0185 | 0.99996 | 10.0246 | 0.012031 | -0.11941 | 0.12001 | 119.731 | 2.968 | 0.048485 | -0.28752 | 1.968 | 0.036454 | -0.2396 | -1.0623 | 34.3352
pf_sweep | 11 | 2 | 2 | 1 | 120 | 1 | 0.1 | 60 | 120.0185 | 0.99996 | 10.0246 | 0.012031 | -0.11941 | 0.12001 | 119.731 | 1.001 | 0.01564 | -0.28752 | 0.001044 | 0.003609 | -0.2396 | 0.1044 | 29.9991
pf_sweep | 11 | 2 | 3 | 1 | 120 | 1 | 0.1 | 60 | 120.0185 | 0.99996 | 10.0246 | 0.012031 | -0.11941 | 0.12001 | 119.731 | 1.001 | 0.015707 | -0.28752 | 0.001044 | 0.003676 | -0.2396 | 0.1044 | 30.556
pf_sweep | 11 | 2 | 4 | 1 | 120 | 1 | 0.1 | 60 | 120.0185 | 0.99996 | 10.0246 | 0.012031 | -0.11941 | 0.12001 | 119.731 | 1 | 0.015698 | -0.28752 | 4.4e-05 | 0.003667 | -0.2396 | 0.0044 | 30.4812
pf_sweep | 11 | 2 | 5 | 1 | 120 | 1 | 0.1 | 60 | 120.0185 | 0.99996 | 10.0246 | 0.012031 | -0.11941 | 0.12001 | 119.731 | 1.001 | 0.015574 | -0.28752 | 0.001044 | 0.003543 | -0.2396 | 0.1044 | 29.4505
pf_sweep | 11 | 2 | 6 | 1 | 120 | 1 | 0.1 | 60 | 120.0185 | 0.99996 | 10.0246 | 0.012031 | -0.11941 | 0.12001 | 119.731 | 0.998 | 0.015094 | -0.28752 | -0.001956 | 0.003063 | -0.2396 | -0.1956 | 25.4608
pf_sweep | 11 | 2 | 7 | 2 | 120 | 1 | 0.1 | 60 | 119.9796 | 1.0001 | 10.0778 | 0.012093 | -0.11938 | 0.11999 | 119.76 | 2.978 | 0.049151 | -0.21958 | 1.9779 | 0.037058 | -0.183 | -0.74371 | 35.4861
pf_sweep | 11 | 2 | 8 | 2 | 120 | 1 | 0.1 | 60 | 119.9796 | 1.0001 | 10.0778 | 0.012093 | -0.11938 | 0.11999 | 119.76 | 1.001 | 0.015646 | -0.21958 | 0.000895 | 0.003553 | -0.183 | 0.0895 | 29.3859
pf_sweep | 11 | 2 | 9 | 2 | 120 | 1 | 0.1 | 60 | 119.9796 | 1.0001 | 10.0778 | 0.012093 | -0.11938 | 0.11999 | 119.76 | 1 | 0.015655 | -0.21958 | -0.000105 | 0.003562 | -0.183 | -0.0105 | 29.4603
pf_sweep | 11 | 2 | 10 | 2 | 120 | 1 | 0.1 | 60 | 119.9796 | 1.0001 | 10.0778 | 0.012093 | -0.11938 | 0.11999 | 119.76 | 1 | 0.015711 | -0.21958 | -0.000105 | 0.003618 | -0.183 | -0.0105 | 29.9234
pf_sweep | 11 | 2 | 11 | 2 | 120 | 1 | 0.1 | 60 | 119.9796 | 1.0001 | 10.0778 | 0.012093 | -0.11938 | 0.11999 | 119.76 | 0.999 | 0.015466 | -0.21958 | -0.001105 | 0.003373 | -0.183 | -0.1104 | 27.8974
pf_sweep | 11 | 2 | 12 | 2 | 120 | 1 | 0.1 | 60 | 119.9796 | 1.0001 | 10.0778 | 0.012093 | -0.11938 | 0.11999 | 119.76 | 1 | 0.015833 | -0.21958 | -0.000105 | 0.00374 | -0.183 | -0.0105 | 30.9323
pf_sweep | 11 | 2 | 13 | 3 | 120 | 1 | 0.1 | 60 | 119.9941 | 0.9999 | 10.0126 | 0.012013 | -0.11938 | 0.11998 | 119.77 | 1 | 0.015501 | -0.22409 | 9.8e-05 | 0.003488 | -0.1868 | 0.0098 | 29.0317
pf_sweep | 11 | 2 | 14 | 3 | 120 | 1 | 0.1 | 60 | 119.9941 | 0.9999 | 10.0126 | 0.012013 | -0.11938 | 0.11998 | 119.77 | 0.999 | 0.015287 | -0.22409 | -0.000902 | 0.003274 | -0.1868 | -0.0902 | 27.2503
pf_sweep | 11 | 2 | 15 | 3 | 120 | 1 | 0.1 | 60 | 119.9941 | 0.9999 | 10.0126 | 0.012013 | -0.11938 | 0.11998 | 119.77 | 1.001 | 0.015418 | -0.22409 | 0.001098 | 0.003405 | -0.1868 | 0.1098 | 28.3408
pf_sweep | 11 | 2 | 16 | 3 | 120 | 1 | 0.1 | 60 | 119.9941 | 0.9999 | 10.0126 | 0.012013 | -0.11938 | 0.11998 | 119.77 | 0.999 | 0.015185 | -0.22409 | -0.000902 | 0.003172 | -0.1868 | -0.0902 | 26.4013
pf_sweep | 11 | 2 | 17 | 3 | 120 | 1 | 0.1 | 60 | 119.9941 | 0.9999 | 10.0126 | 0.012013 | -0.11938 | 0.11998 | 119.77 | 1 | 0.015394 | -0.22409 | 9.8e-05 | 0.003381 | -0.1868 | 0.0098 | 28.141
pf_sweep | 11 | 2 | 18 | 3 | 120 | 1 | 0.1 | 60 | 119.9941 | 0.9999 | 10.0126 | 0.012013 | -0.11938 | 0.11998 | 119.77 | 1 | 0.01529 | -0.22409 | 9.8e-05 | 0.003277 | -0.1868 | 0.0098 | 27.2753
pf_sweep | 11 | 3 | 1 | 1 | 120 | 1 | 0.1 | 60 | 120.0121 | 0.99994 | 10.0162 | 0.01202 | -0.1194 | 0.12001 | 119.706 | 2.968 | 0.048451 | -0.30608 | 1.9681 | 0.036431 | -0.255 | -1.061 | 34.3623
pf_sweep | 11 | 3 | 2 | 1 | 120 | 1 | 0.1 | 60 | 120.0121 | 0.99994 | 10.0162 | 0.01202 | -0.1194 | 0.12001 | 119.706 | 1.001 | 0.015672 | -0.30608 | 0.001057 | 0.003652 | -0.255 | 0.1057 | 30.3828
pf_sweep | 11 | 3 | 3 | 1 | 120 | 1 | 0.1 | 60 | 120.0121 | 0.99994 | 10.0162 | 0.01202 | -0.1194 | 0.12001 | 119.706 | 1.001 | 0.015718 | -0.30608 | 0.001057 | 0.003698 | -0.255 | 0.1057 | 30.7655
pf_sweep | 11 | 3 | 4 | 1 | 120 | 1 | 0.1 | 60 | 120.0121 | 0.99994 | 10.0162 | 0.01202 | -0.1194 | 0.12001 | 119.706 | 1.001 | 0.015693 | -0.30608 | 0.001057 | 0.003673 | -0.255 | 0.1057 | 30.5575
pf_sweep | 11 | 3 | 5 | 1 | 120 | 1 | 0.1 | 60 | 120.0121 | 0.99994 | 10.0162 | 0.01202 | -0.1194 | 0.12001 | 119.706 | 1.001 | 0.015588 | -0.30608 | 0.001057 | 0.003568 | -0.255 | 0.1057 | 29.684
pf_sweep | 11 | 3 | 6 | 1 | 120 | 1 | 0.1 | 60 | 120.0121 | 0.99994 | 10.0162 | 0.01202 | -0.1194 | 0.12001 | 119.706 | 0.998 | 0.015124 | -0.30608 | -0.001943 | 0.003104 | -0.255 | -0.1943 | 25.8238
pf_sweep | 11 | 3 | 7 | 2 | 120 | 1 | 0.1 | 60 | 119.9856 | 1.0001 | 10.0807 | 0.012097 | -0.11939 | 0.12 | 119.771 | 2.978 | 0.0492 | -0.2146 | 1.9779 | 0.037103 | -0.1789 | -0.74278 | 35.5759
pf_sweep | 11 | 3 | 8 | 2 | 120 | 1 | 0.1 | 60 | 119.9856 | 1.0001 | 10.0807 | 0.012097 | -0.11939 | 0.12 | 119.771 | 1.001 | 0.015607 | -0.2146 | 0.000905 | 0.00351 | -0.1789 | 0.0905 | 29.0203
pf_sweep | 11 | 3 | 9 | 2 | 120 | 1 | 0.1 | 60 | 119.9856 | 1.0001 | 10.0807 | 0.012097 | -0.11939 | 0.12 | 119.771 | 1.001 | 0.015643 | -0.2146 | 0.000905 | 0.003546 | -0.1789 | 0.0905 | 29.3179
pf_sweep | 11 | 3 | 10 | 2 | 120 | 1 | 0.1 | 60 | 119.9856 | 1.0001 | 10.0807 | 0.012097 | -0.11939 | 0.12 | 119.771 | 1 | 0.015721 | -0.2146 | -9.5e-05 | 0.003624 | -0.1789 | -0.0095 | 29.9628
pf_sweep | 11 | 3 | 11 | 2 | 120 | 1 | 0.1 | 60 | 119.9856 | 1.0001 | 10.0807 | 0.012097 | -0.11939 | 0.12 | 119.771 | 1 | 0.015475 | -0.2146 | -9.5e-05 | 0.003378 | -0.1789 | -0.0095 | 27.9291
pf_sweep | 11 | 3 | 12 | 2 | 120 | 1 | 0.1 | 60 | 119.9856 | 1.0001 | 10.0807 | 0.012097 | -0.11939 | 0.12 | 119.771 | 1 | 0.01583 | -0.2146 | -9.5e-05 | 0.003733 | -0.1789 | -0.0095 | 30.8638
pf_sweep | 11 | 3 | 13 | 3 | 120 | 1 | 0.1 | 60 | 119.9889 | 0.99991 | 10.0146 | 0.012015 | -0.11937 | 0.11998 | 119.781 | 1 | 0.015493 | -0.20791 | 9.1e-05 | 0.003478 | -0.1733 | 0.0091 | 28.9434
pf_sweep | 11 | 3 | 14 | 3 | 120 | 1 | 0.1 | 60 | 119.9889 | 0.99991 | 10.0146 | 0.012015 | -0.11937 | 0.11998 | 119.781 | 0.999 | 0.015268 | -0.20791 | -0.000909 | 0.003253 | -0.1733 | -0.0909 | 27.0708
pf_sweep | 11 | 3 | 15 | 3 | 120 | 1 | 0.1 | 60 | 119.9889 | 0.99991 | 10.0146 | 0.012015 | -0.11937 | 0.11998 | 119.781 | 1.001 | 0.015419 | -0.20791 | 0.001091 | 0.003404 | -0.1733 | 0.1091 | 28.3276
pf_sweep | 11 | 3 | 16 | 3 | 120 | 1 | 0.1 | 60 | 119.9889 | 0.99991 | 10.0146 | 0.012015 | -0.11937 | 0.11998 | 119.781 | 0.999 | 0.015206 | -0.20791 | -0.000909 | 0.003191 | -0.1733 | -0.0909 | 26.5548
pf_sweep | 11 | 3 | 17 | 3 | 120 | 1 | 0.1 | 60 | 119.9889 | 0.99991 | 10.0146 | 0.012015 | -0.11937 | 0.11998 | 119.781 | 1 | 0.015375 | -0.20791 | 9.1e-05 | 0.00336 | -0.1733 | 0.0091 | 27.9614
pf_sweep | 11 | 3 | 18 | 3 | 120 | 1 | 0.1 | 60 | 119.9889 | 0.99991 | 10.0146 | 0.012015 | -0.11937 | 0.11998 | 119.781 | 1 | 0.015329 | -0.20791 | 9.1e-05 | 0.003314 | -0.1733 | 0.0091 | 27.5785
pf_sweep | 12 | 1 | 1 | 1 | 120 | 1 | 0.2 | 60 | 120.0163 | 0.99996 | 20.0201 | 0.024026 | -0.11758 | 0.12001 | 119.73 | 2.968 | 0.083888 | -0.28634 | 1.968 | 0.059862 | -0.2386 | -1.0629 | 16.3827
pf_sweep | 12 | 1 | 2 | 1 | 120 | 1 | 0.2 | 60 | 120.0163 | 0.99996 | 20.0201 | 0.024026 | -0.11758 | 0.12001 | 119.73 | 1.001 | 0.027637 | -0.28634 | 0.001038 | 0.003611 | -0.2386 | 0.1038 | 15.0272
pf_sweep | 12 | 1 | 3 | 1 | 120 | 1 | 0.2 | 60 | 120.0163 | 0.99996 | 20.0201 | 0.024026 | -0.11758 | 0.12001 | 119.73 | 1.001 | 0.02768 | -0.28634 | 0.001038 | 0.003654 | -0.2386 | 0.1038 | 15.2062
pf_sweep | 12 | 1 | 4 | 1 | 120 | 1 | 0.2 | 60 | 120.0163 | 0.99996 | 20.0201 | 0.024026 | -0.11758 | 0.12001 | 119.73 | 1 | 0.027674 | -0.28634 | 3.8e-05 | 0.003648 | -0.2386 | 0.0038 | 15.1812
pf_sweep | 12 | 1 | 5 | 1 | 120 | 1 | 0.2 | 60 | 120.0163 | 0.99996 | 20.0201 | 0.024026 | -0.11758 | 0.12001 | 119.73 | 1.001 | 0.027564 | -0.28634 | 0.001038 | 0.003538 | -0.2386 | 0.1038 | 14.7234
pf_sweep | 12 | 1 | 6 | 1 | 120 | 1 | 0.2 | 60 | 120.0163 | 0.99996 | 20.0201 | 0.024026 | -0.11758 | 0.12001 | 119.73 | 0.998 | 0.027052 | -0.28634 | -0.001962 | 0.003026 | -0.2386 | -0.1962 | 12.5924
pf_sweep | 12 | 1 | 7 | 2 | 120 | 1 | 0.2 | 60 | 119.9834 | 1.0001 | 20.0746 | 0.024088 | -0.11755 | 0.11999 | 119.763 | 2.978 | 0.08479 | -0.22039 | 1.9779 | 0.060702 | -0.1837 | -0.74086 | 17.3334
pf_sweep | 12 | 1 | 8 | 2 | 120 | 1 | 0.2 | 60 | 119.9834 | 1.0001 | 20.0746 | 0.024088 | -0.11755 | 0.11999 | 119.763 | 1.001 | 0.027573 | -0.22039 | 0.000924 | 0.003485 | -0.1837 | 0.0924 | 14.4675
pf_sweep | 12 | 1 | 9 | 2 | 120 | 1 | 0.2 | 60 | 119.9834 | 1.0001 | 20.0746 | 0.024088 | -0.11755 | 0.11999 | 119.763 | 1.001 | 0.027624 | -0.22039 | 0.000924 | 0.003536 | -0.1837 | 0.0924 | 14.6793
pf_sweep | 12 | 1 | 10 | 2 | 120 | 1 | 0.2 | 60 | 119.9834 | 1.0001 | 20.0746 | 0.024088 | -0.11755 | 0.11999 | 119.763 | 1 | 0.027686 | -0.22039 | -7.6e-05 | 0.003598 | -0.1837 | -0.0076 | 14.9367
pf_sweep | 12 | 1 | 11 | 2 | 120 | 1 | 0.2 | 60 | 119.9834 | 1.0001 | 20.0746 | 0.024088 | -0.11755 | 0.11999 | 119.763 | 1 | 0.027411 | -0.22039 | -7.6e-05 | 0.003323 | -0.1837 | -0.0076 | 13.795
pf_sweep | 12 | 1 | 12 | 2 | 120 | 1 | 0.2 | 60 | 119.9834 | 1.0001 | 20.0746 | 0.024088 | -0.11755 | 0.11999 | 119.763 | 1 | 0.027789 | -0.22039 | -7.6e-05 | 0.003701 | -0.1837 | -0.0076 | 15.3643
pf_sweep | 12 | 1 | 13 | 3 | 120 | 1 | 0.2 | 60 | 119.9939 | 0.9999 | 20.015 | 0.024014 | -0.11755 | 0.11998 | 119.772 | 1 | 0.027468 | -0.22185 | 9.6e-05 | 0.003454 | -0.1849 | 0.0096 | 14.3812
pf_sweep | 12 | 1 | 14 | 3 | 120 | 1 | 0.2 | 60 | 119.9939 | 0.9999 | 20.015 | 0.024014 | -0.11755 | 0.11998 | 119.772 | 0.999 | 0.027276 | -0.22185 | -0.000904 | 0.003262 | -0.1849 | -0.0904 | 13.5817
pf_sweep | 12 | 1 | 15 | 3 | 120 | 1 | 0.2 | 60 | 119.9939 | 0.9999 | 20.015 | 0.024014 | -0.11755 | 0.11998 | 119.772 | 1.001 | 0.027427 | -0.22185 | 0.001096 | 0.003413 | -0.1849 | 0.1096 | 14.2104
pf_sweep | 12 | 1 | 16 | 3 | 120 | 1 | 0.2 | 60 | 119.9939 | 0.9999 | 20.015 | 0.024014 | -0.11755 | 0.11998 | 119.772 | 0.999 | 0.027174 | -0.22185 | -0.000904 | 0.00316 | -0.1849 | -0.0904 | 13.1569
pf_sweep | 12 | 1 | 17 | 3 | 120 | 1 | 0.2 | 60 | 119.9939 | 0.9999 | 20.015 | 0.024014 | -0.11755 | 0.11998 | 119.772 | 1 | 0.027388 | -0.22185 | 9.6e-05 | 0.003374 | -0.1849 | 0.0096 | 14.048
pf_sweep | 12 | 1 | 18 | 3 | 120 | 1 | 0.2 | 60 | 119.9939 | 0.9999 | 20.015 | 0.024014 | -0.11755 | 0.11998 | 119.772 | 1 | 0.027292 | -0.22185 | 9.6e-05 | 0.003278 | -0.1849 | 0.0096 | 13.6483
pf_sweep | 12 | 2 | 1 | 1 | 120 | 1 | 0.2 | 60 | 120.0146 | 0.99996 | 20.0238 | 0.02403 | -0.11758 | 0.12001 | 119.756 | 2.968 | 0.083983 | -0.25864 | 1.968 | 0.059953 | -0.2155 | -1.0623 | 16.4952
pf_sweep | 12 | 2 | 2 | 1 | 120 | 1 | 0.2 | 60 | 120.0146 | 0.99996 | 20.0238 | 0.02403 | -0.11758 | 0.12001 | 119.756 | 1.001 | 0.027644 | -0.25864 | 0.001044 | 0.003614 | -0.2155 | 0.1044 | 15.0374
pf_sweep | 12 | 2 | 3 | 1 | 120 | 1 | 0.2 | 60 | 120.0146 | 0.99996 | 20.0238 | 0.02403 | -0.11758 | 0.12001 | 119.756 | 1.001 | 0.027677 | -0.25864 | 0.001044 | 0.003647 | -0.2155 | 0.1044 | 15.1747
pf_sweep | 12 | 2 | 4 | 1 | 120 | 1 | 0.2 | 60 | 120.0146 | 0.99996 | 20.0238 | 0.02403 | -0.11758 | 0.12001 | 119.756 | 1 | 0.0277 | -0.25864 | 4.4e-05 | 0.00367 | -0.2155 | 0.0044 | 15.2704
pf_sweep | 12 | 2 | 5 | 1 | 120 | 1 | 0.2 | 60 | 120.0146 | 0.99996 | 20.0238 | 0.02403 | -0.11758 | 0.12001 | 119.756 | 1.001 | 0.027572 | -0.25864 | 0.001044 | 0.003542 | -0.2155 | 0.1044 | 14.7378
pf_sweep | 12 | 2 | 6 | 1 | 120 | 1 | 0.2 | 60 | 120.0146 | 0.99996 | 20.0238 | 0.02403 | -0.11758 | 0.12001 | 119.756 | 0.998 | 0.027048 | -0.25864 | -0.001956 | 0.003018 | -0.2155 | -0.1956 | 12.5572
pf_sweep | 12 | 2 | 7 | 2 | 120 | 1 | 0.2 | 60 | 119.9823 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.758 | 2.977 | 0.084791 | -0.22427 | 1.9769 | 0.060702 | -0.1869 | -0.77599 | 17.3296
pf_sweep | 12 | 2 | 8 | 2 | 120 | 1 | 0.2 | 60 | 119.9823 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.758 | 1.001 | 0.027592 | -0.22427 | 0.000906 | 0.003503 | -0.1869 | 0.0906 | 14.5414
pf_sweep | 12 | 2 | 9 | 2 | 120 | 1 | 0.2 | 60 | 119.9823 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.758 | 1.001 | 0.027651 | -0.22427 | 0.000906 | 0.003562 | -0.1869 | 0.0906 | 14.7863
pf_sweep | 12 | 2 | 10 | 2 | 120 | 1 | 0.2 | 60 | 119.9823 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.758 | 1 | 0.027686 | -0.22427 | -9.4e-05 | 0.003597 | -0.1869 | -0.0094 | 14.9316
pf_sweep | 12 | 2 | 11 | 2 | 120 | 1 | 0.2 | 60 | 119.9823 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.758 | 1 | 0.027411 | -0.22427 | -9.4e-05 | 0.003322 | -0.1869 | -0.0094 | 13.79
pf_sweep | 12 | 2 | 12 | 2 | 120 | 1 | 0.2 | 60 | 119.9823 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.758 | 1 | 0.027817 | -0.22427 | -9.4e-05 | 0.003728 | -0.1869 | -0.0094 | 15.4754
pf_sweep | 12 | 2 | 13 | 3 | 120 | 1 | 0.2 | 60 | 119.9938 | 0.99993 | 20.0139 | 0.024014 | -0.11756 | 0.11999 | 119.783 | 1 | 0.027449 | -0.21079 | 6.8e-05 | 0.003435 | -0.1757 | 0.0068 | 14.3049
pf_sweep | 12 | 2 | 14 | 3 | 120 | 1 | 0.2 | 60 | 119.9938 | 0.99993 | 20.0139 | 0.024014 | -0.11756 | 0.11999 | 119.783 | 0.999 | 0.027267 | -0.21079 | -0.000932 | 0.003253 | -0.1757 | -0.0932 | 13.547
pf_sweep | 12 | 2 | 15 | 3 | 120 | 1 | 0.2 | 60 | 119.9938 | 0.99993 | 20.0139 | 0.024014 | -0.11756 | 0.11999 | 119.783 | 1.001 | 0.027439 | -0.21079 | 0.001068 | 0.003425 | -0.1757 | 0.1068 | 14.2633
pf_sweep | 12 | 2 | 16 | 3 | 120 | 1 | 0.2 | 60 | 119.9938 | 0.99993 | 20.0139 | 0.024014 | -0.11756 | 0.11999 | 119.783 | 0.999 | 0.027165 | -0.21079 | -0.000932 | 0.003151 | -0.1757 | -0.0932 | 13.1223
pf_sweep | 12 | 2 | 17 | 3 | 120 | 1 | 0.2 | 60 | 119.9938 | 0.99993 | 20.0139 | 0.024014 | -0.11756 | 0.11999 | 119.783 | 1 | 0.027369 | -0.21079 | 6.8e-05 | 0.003355 | -0.1757 | 0.0068 | 13.9718
pf_sweep | 12 | 2 | 18 | 3 | 120 | 1 | 0.2 | 60 | 119.9938 | 0.99993 | 20.0139 | 0.024014 | -0.11756 | 0.11999 | 119.783 | 1 | 0.027314 | -0.21079 | 6.8e-05 | 0.0033 | -0.1757 | 0.0068 | 13.7428
pf_sweep | 12 | 3 | 1 | 1 | 120 | 1 | 0.2 | 60 | 119.9892 | 0.99997 | 20.0198 | 0.024021 | -0.11756 | 0.11999 | 119.738 | 2.968 | 0.083784 | -0.25121 | 1.968 | 0.059763 | -0.2094 | -1.0638 | 16.2654
pf_sweep | 12 | 3 | 2 | 1 | 120 | 1 | 0.2 | 60 | 119.9892 | 0.99997 | 20.0198 | 0.024021 | -0.11756 | 0.11999 | 119.738 | 1.001 | 0.027643 | -0.25121 | 0.001029 | 0.003622 | -0.2094 | 0.1029 | 15.0789
pf_sweep | 12 | 3 | 3 | 1 | 120 | 1 | 0.2 | 60 | 119.9892 | 0.99997 | 20.0198 | 0.024021 | -0.11756 | 0.11999 | 119.738 | 1.001 | 0.027672 | -0.25121 | 0.001029 | 0.003651 | -0.2094 | 0.1029 | 15.1996
pf_sweep | 12 | 3 | 4 | 1 | 120 | 1 | 0.2 | 60 | 119.9892 | 0.99997 | 20.0198 | 0.024021 | -0.11756 | 0.11999 | 119.738 | 1 | 0.027634 | -0.25121 | 2.9e-05 | 0.003613 | -0.2094 | 0.0029 | 15.0414
pf_sweep | 12 | 3 | 5 | 1 | 120 | 1 | 0.2 | 60 | 119.9892 | 0.99997 | 20.0198 | 0.024021 | -0.11756 | 0.11999 | 119.738 | 1.001 | 0.027552 | -0.25121 | 0.001029 | 0.003531 | -0.2094 | 0.1029 | 14.7001
pf_sweep | 12 | 3 | 6 | 1 | 120 | 1 | 0.2 | 60 | 119.9892 | 0.99997 | 20.0198 | 0.024021 | -0.11756 | 0.11999 | 119.738 | 0.998 | 0.027083 | -0.25121 | -0.001971 | 0.003062 | -0.2094 | -0.1971 | 12.7476
pf_sweep | 12 | 3 | 7 | 2 | 120 | 1 | 0.2 | 60 | 119.9814 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.774 | 2.978 | 0.084781 | -0.20744 | 1.9779 | 0.060692 | -0.1729 | -0.74289 | 17.3168
pf_sweep | 12 | 3 | 8 | 2 | 120 | 1 | 0.2 | 60 | 119.9814 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.774 | 1.001 | 0.027622 | -0.20744 | 0.000904 | 0.003533 | -0.1729 | 0.0904 | 14.6669
pf_sweep | 12 | 3 | 9 | 2 | 120 | 1 | 0.2 | 60 | 119.9814 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.774 | 1 | 0.027655 | -0.20744 | -9.6e-05 | 0.003566 | -0.1729 | -0.0096 | 14.8039
pf_sweep | 12 | 3 | 10 | 2 | 120 | 1 | 0.2 | 60 | 119.9814 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.774 | 1 | 0.027683 | -0.20744 | -9.6e-05 | 0.003594 | -0.1729 | -0.0096 | 14.9201
pf_sweep | 12 | 3 | 11 | 2 | 120 | 1 | 0.2 | 60 | 119.9814 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.774 | 1 | 0.02745 | -0.20744 | -9.6e-05 | 0.003361 | -0.1729 | -0.0096 | 13.9529
pf_sweep | 12 | 3 | 12 | 2 | 120 | 1 | 0.2 | 60 | 119.9814 | 1.0001 | 20.0753 | 0.024089 | -0.11755 | 0.11999 | 119.774 | 1 | 0.027821 | -0.20744 | -9.6e-05 | 0.003732 | -0.1729 | -0.0096 | 15.493
pf_sweep | 12 | 3 | 13 | 3 | 120 | 1 | 0.2 | 60 | 119.9941 | 0.99991 | 20.0121 | 0.024011 | -0.11756 | 0.11998 | 119.776 | 1 | 0.02748 | -0.21806 | 8.5e-05 | 0.003469 | -0.1817 | 0.0085 | 14.4462
pf_sweep | 12 | 3 | 14 | 3 | 120 | 1 | 0.2 | 60 | 119.9941 | 0.99991 | 20.0121 | 0.024011 | -0.11756 | 0.11998 | 119.776 | 0.999 | 0.027251 | -0.21806 | -0.000915 | 0.00324 | -0.1817 | -0.0915 | 13.4925
pf_sweep | 12 | 3 | 15 | 3 | 120 | 1 | 0.2 | 60 | 119.9941 | 0.99991 | 20.0121 | 0.024011 | -0.11756 | 0.11998 | 119.776 | 1.001 | 0.02738 | -0.21806 | 0.001085 | 0.003369 | -0.1817 | 0.1085 | 14.0297
pf_sweep | 12 | 3 | 16 | 3 | 120 | 1 | 0.2 | 60 | 119.9941 | 0.99991 | 20.0121 | 0.024011 | -0.11756 | 0.11998 | 119.776 | 0.999 | 0.027186 | -0.21806 | -0.000915 | 0.003175 | -0.1817 | -0.0915 | 13.2218
pf_sweep | 12 | 3 | 17 | 3 | 120 | 1 | 0.2 | 60 | 119.9941 | 0.99991 | 20.0121 | 0.024011 | -0.11756 | 0.11998 | 119.776 | 1 | 0.027372 | -0.21806 | 8.5e-05 | 0.003361 | -0.1817 | 0.0085 | 13.9964
pf_sweep | 12 | 3 | 18 | 3 | 120 | 1 | 0.2 | 60 | 119.9941 | 0.99991 | 20.0121 | 0.024011 | -0.11756 | 0.11998 | 119.776 | 1 | 0.027264 | -0.21806 | 8.5e-05 | 0.003253 | -0.1817 | 0.0085 | 13.5466
pf_sweep | 13 | 1 | 1 | 1 | 120 | 1 | 0.3 | 60 | 120.0088 | 0.99997 | 30.0214 | 0.036027 | -0.11447 | 0.12 | 119.75 | 2.968 | 0.11911 | -0.25884 | 1.968 | 0.083078 | -0.2157 | -1.0633 | 10.1993
pf_sweep | 13 | 1 | 2 | 1 | 120 | 1 | 0.3 | 60 | 120.0088 | 0.99997 | 30.0214 | 0.036027 | -0.11447 | 0.12 | 119.75 | 1.001 | 0.039623 | -0.25884 | 0.001034 | 0.003596 | -0.2157 | 0.1034 | 9.981
pf_sweep | 13 | 1 | 3 | 1 | 120 | 1 | 0.3 | 60 | 120.0088 | 0.99997 | 30.0214 | 0.036027 | -0.11447 | 0.12 | 119.75 | 1.001 | 0.039569 | -0.25884 | 0.001034 | 0.003542 | -0.2157 | 0.1034 | 9.8311
pf_sweep | 13 | 1 | 4 | 1 | 120 | 1 | 0.3 | 60 | 120.0088 | 0.99997 | 30.0214 | 0.036027 | -0.11447 | 0.12 | 119.75 | 1 | 0.039559 | -0.25884 | 3.4e-05 | 0.003532 | -0.2157 | 0.0034 | 9.8033
pf_sweep | 13 | 1 | 5 | 1 | 120 | 1 | 0.3 | 60 | 120.0088 | 0.99997 | 30.0214 | 0.036027 | -0.11447 | 0.12 | 119.75 | 1.001 | 0.039563 | -0.25884 | 0.001034 | 0.003536 | -0.2157 | 0.1034 | 9.8144
pf_sweep | 13 | 1 | 6 | 1 | 120 | 1 | 0.3 | 60 | 120.0088 | 0.99997 | 30.0214 | 0.036027 | -0.11447 | 0.12 | 119.75 | 0.998 | 0.038936 | -0.25884 | -0.001966 | 0.002909 | -0.2157 | -0.1966 | 8.0741
pf_sweep | 13 | 1 | 7 | 2 | 120 | 1 | 0.3 | 60 | 119.9809 | 1.0001 | 30.0763 | 0.036088 | -0.11443 | 0.11999 | 119.775 | 2.978 | 0.12038 | -0.20595 | 1.9779 | 0.084289 | -0.1717 | -0.74018 | 11.1872
pf_sweep | 13 | 1 | 8 | 2 | 120 | 1 | 0.3 | 60 | 119.9809 | 1.0001 | 30.0763 | 0.036088 | -0.11443 | 0.11999 | 119.775 | 1.001 | 0.039554 | -0.20595 | 0.000931 | 0.003466 | -0.1717 | 0.0931 | 9.6032
pf_sweep | 13 | 1 | 9 | 2 | 120 | 1 | 0.3 | 60 | 119.9809 | 1.0001 | 30.0763 | 0.036088 | -0.11443 | 0.11999 | 119.775 | 1.001 | 0.039612 | -0.20595 | 0.000931 | 0.003524 | -0.1717 | 0.0931 | 9.7639
pf_sweep | 13 | 1 | 10 | 2 | 120 | 1 | 0.3 | 60 | 119.9809 | 1.0001 | 30.0763 | 0.036088 | -0.11443 | 0.11999 | 119.775 | 1 | 0.03965 | -0.20595 | -6.9e-05 | 0.003562 | -0.1717 | -0.0069 | 9.8692
pf_sweep | 13 | 1 | 11 | 2 | 120 | 1 | 0.3 | 60 | 119.9809 | 1.0001 | 30.0763 | 0.036088 | -0.11443 | 0.11999 | 119.775 | 0.999 | 0.039382 | -0.20595 | -0.001069 | 0.003294 | -0.1717 | -0.1069 | 9.1266
pf_sweep | 13 | 1 | 12 | 2 | 120 | 1 | 0.3 | 60 | 119.9809 | 1.0001 | 30.0763 | 0.036088 | -0.11443 | 0.11999 | 119.775 | 1 | 0.039757 | -0.20595 | -6.9e-05 | 0.003669 | -0.1717 | -0.0069 | 10.1657
pf_sweep | 13 | 1 | 13 | 3 | 120 | 1 | 0.3 | 60 | 119.9921 | 0.9999 | 30.0099 | 0.036006 | -0.11445 | 0.11998 | 119.769 | 1 | 0.039392 | -0.22307 | 0.0001 | 0.003386 | -0.1859 | 0.01 | 9.4043
pf_sweep | 13 | 1 | 14 | 3 | 120 | 1 | 0.3 | 60 | 119.9921 | 0.9999 | 30.0099 | 0.036006 | -0.11445 | 0.11998 | 119.769 | 0.999 | 0.039191 | -0.22307 | -0.0009 | 0.003185 | -0.1859 | -0.09 | 8.8461
pf_sweep | 13 | 1 | 15 | 3 | 120 | 1 | 0.3 | 60 | 119.9921 | 0.9999 | 30.0099 | 0.036006 | -0.11445 | 0.11998 | 119.769 | 1 | 0.039352 | -0.22307 | 0.0001 | 0.003346 | -0.1859 | 0.01 | 9.2932
pf_sweep | 13 | 1 | 16 | 3 | 120 | 1 | 0.3 | 60 | 119.9921 | 0.9999 | 30.0099 | 0.036006 | -0.11445 | 0.11998 | 119.769 | 0.999 | 0.039101 | -0.22307 | -0.0009 | 0.003095 | -0.1859 | -0.09 | 8.5961
pf_sweep | 13 | 1 | 17 | 3 | 120 | 1 | 0.3 | 60 | 119.9921 | 0.9999 | 30.0099 | 0.036006 | -0.11445 | 0.11998 | 119.769 | 1 | 0.039306 | -0.22307 | 0.0001 | 0.0033 | -0.1859 | 0.01 | 9.1655
pf_sweep | 13 | 1 | 18 | 3 | 120 | 1 | 0.3 | 60 | 119.9921 | 0.9999 | 30.0099 | 0.036006 | -0.11445 | 0.11998 | 119.769 | 1 | 0.03923 | -0.22307 | 0.0001 | 0.003224 | -0.1859 | 0.01 | 8.9544
pf_sweep | 13 | 2 | 1 | 1 | 120 | 1 | 0.3 | 60 | 120.0161 | 0.99998 | 30.0172 | 0.036025 | -0.11448 | 0.12001 | 119.707 | 2.968 | 0.11926 | -0.30907 | 1.968 | 0.083231 | -0.2575 | -1.0644 | 10.3469
pf_sweep | 13 | 2 | 2 | 1 | 120 | 1 | 0.3 | 60 | 120.0161 | 0.99998 | 30.0172 | 0.036025 | -0.11448 | 0.12001 | 119.707 | 1.001 | 0.03957 | -0.30907 | 0.001023 | 0.003545 | -0.2575 | 0.1023 | 9.8417
pf_sweep | 13 | 2 | 3 | 1 | 120 | 1 | 0.3 | 60 | 120.0161 | 0.99998 | 30.0172 | 0.036025 | -0.11448 | 0.12001 | 119.707 | 1.001 | 0.03962 | -0.30907 | 0.001023 | 0.003595 | -0.2575 | 0.1023 | 9.9805
pf_sweep | 13 | 2 | 4 | 1 | 120 | 1 | 0.3 | 60 | 120.0161 | 0.99998 | 30.0172 | 0.036025 | -0.11448 | 0.12001 | 119.707 | 1 | 0.039615 | -0.30907 | 2.3e-05 | 0.00359 | -0.2575 | 0.0023 | 9.9666
pf_sweep | 13 | 2 | 5 | 1 | 120 | 1 | 0.3 | 60 | 120.0161 | 0.99998 | 30.0172 | 0.036025 | -0.11448 | 0.12001 | 119.707 | 1.001 | 0.039491 | -0.30907 | 0.001023 | 0.003466 | -0.2575 | 0.1023 | 9.6224
pf_sweep | 13 | 2 | 6 | 1 | 120 | 1 | 0.3 | 60 | 120.0161 | 0.99998 | 30.0172 | 0.036025 | -0.11448 | 0.12001 | 119.707 | 0.998 | 0.038939 | -0.30907 | -0.001977 | 0.002914 | -0.2575 | -0.1977 | 8.0901
pf_sweep | 13 | 2 | 7 | 2 | 120 | 1 | 0.3 | 60 | 119.9825 | 1.0001 | 30.0762 | 0.03609 | -0.11444 | 0.11999 | 119.758 | 2.977 | 0.12025 | -0.22454 | 1.9769 | 0.084162 | -0.1871 | -0.77616 | 11.0679
pf_sweep | 13 | 2 | 8 | 2 | 120 | 1 | 0.3 | 60 | 119.9825 | 1.0001 | 30.0762 | 0.03609 | -0.11444 | 0.11999 | 119.758 | 1.001 | 0.039538 | -0.22454 | 0.000904 | 0.003448 | -0.1871 | 0.0904 | 9.555
pf_sweep | 13 | 2 | 9 | 2 | 120 | 1 | 0.3 | 60 | 119.9825 | 1.0001 | 30.0762 | 0.03609 | -0.11444 | 0.11999 | 119.758 | 1 | 0.039607 | -0.22454 | -9.6e-05 | 0.003517 | -0.1871 | -0.0096 | 9.7462
pf_sweep | 13 | 2 | 10 | 2 | 120 | 1 | 0.3 | 60 | 119.9825 | 1.0001 | 30.0762 | 0.03609 | -0.11444 | 0.11999 | 119.758 | 1 | 0.03963 | -0.22454 | -9.6e-05 | 0.00354 | -0.1871 | -0.0096 | 9.8099
pf_sweep | 13 | 2 | 11 | 2 | 120 | 1 | 0.3 | 60 | 119.9825 | 1.0001 | 30.0762 | 0.03609 | -0.11444 | 0.11999 | 119.758 | 0.999 | 0.039348 | -0.22454 | -0.001096 | 0.003258 | -0.1871 | -0.1096 | 9.0285
pf_sweep | 13 | 2 | 12 | 2 | 120 | 1 | 0.3 | 60 | 119.9825 | 1.0001 | 30.0762 | 0.03609 | -0.11444 | 0.11999 | 119.758 | 1 | 0.039761 | -0.22454 | -9.6e-05 | 0.003671 | -0.1871 | -0.0096 | 10.1729
pf_sweep | 13 | 2 | 13 | 3 | 120 | 1 | 0.3 | 60 | 119.9893 | 0.99993 | 30.0122 | 0.036009 | -0.11445 | 0.11998 | 119.772 | 1 | 0.039397 | -0.21733 | 7.1e-05 | 0.003388 | -0.1811 | 0.0071 | 9.4092
pf_sweep | 13 | 2 | 14 | 3 | 120 | 1 | 0.3 | 60 | 119.9893 | 0.99993 | 30.0122 | 0.036009 | -0.11445 | 0.11998 | 119.772 | 0.999 | 0.03921 | -0.21733 | -0.000929 | 0.003201 | -0.1811 | -0.0929 | 8.8899
pf_sweep | 13 | 2 | 15 | 3 | 120 | 1 | 0.3 | 60 | 119.9893 | 0.99993 | 30.0122 | 0.036009 | -0.11445 | 0.11998 | 119.772 | 1.001 | 0.039349 | -0.21733 | 0.001071 | 0.00334 | -0.1811 | 0.1071 | 9.2759
pf_sweep | 13 | 2 | 16 | 3 | 120 | 1 | 0.3 | 60 | 119.9893 | 0.99993 | 30.0122 | 0.036009 | -0.11445 | 0.11998 | 119.772 | 1 | 0.039106 | -0.21733 | 7.1e-05 | 0.003097 | -0.1811 | 0.0071 | 8.601
pf_sweep | 13 | 2 | 17 | 3 | 120 | 1 | 0.3 | 60 | 119.9893 | 0.99993 | 30.0122 | 0.036009 | -0.11445 | 0.11998 | 119.772 | 1 | 0.039306 | -0.21733 | 7.1e-05 | 0.003297 | -0.1811 | 0.0071 | 9.1565
pf_sweep | 13 | 2 | 18 | 3 | 120 | 1 | 0.3 | 60 | 119.9893 | 0.99993 | 30.0122 | 0.036009 | -0.11445 | 0.11998 | 119.772 | 1 | 0.039197 | -0.21733 | 7.1e-05 | 0.003188 | -0.1811 | 0.0071 | 8.8538
pf_sweep | 13 | 3 | 1 | 1 | 120 | 1 | 0.3 | 60 | 120.0077 | 0.99995 | 30.0195 | 0.036024 | -0.11447 | 0.12 | 119.73 | 2.968 | 0.11921 | -0.27768 | 1.9681 | 0.083187 | -0.2314 | -1.0614 | 10.3075
pf_sweep | 13 | 3 | 2 | 1 | 120 | 1 | 0.3 | 60 | 120.0077 | 0.99995 | 30.0195 | 0.036024 | -0.11447 | 0.12 | 119.73 | 1.001 | 0.039577 | -0.27768 | 0.001053 | 0.003553 | -0.2314 | 0.1053 | 9.8634
pf_sweep | 13 | 3 | 3 | 1 | 120 | 1 | 0.3 | 60 | 120.0077 | 0.99995 | 30.0195 | 0.036024 | -0.11447 | 0.12 | 119.73 | 1.001 | 0.03964 | -0.27768 | 0.001053 | 0.003616 | -0.2314 | 0.1053 | 10.0383
pf_sweep | 13 | 3 | 4 | 1 | 120 | 1 | 0.3 | 60 | 120.0077 | 0.99995 | 30.0195 | 0.036024 | -0.11447 | 0.12 | 119.73 | 1 | 0.039577 | -0.27768 | 5.3e-05 | 0.003553 | -0.2314 | 0.0053 | 9.8634
pf_sweep | 13 | 3 | 5 | 1 | 120 | 1 | 0.3 | 60 | 120.0077 | 0.99995 | 30.0195 | 0.036024 | -0.11447 | 0.12 | 119.73 | 1.001 | 0.039488 | -0.27768 | 0.001053 | 0.003464 | -0.2314 | 0.1053 | 9.6163
pf_sweep | 13 | 3 | 6 | 1 | 120 | 1 | 0.3 | 60 | 120.0077 | 0.99995 | 30.0195 | 0.036024 | -0.11447 | 0.12 | 119.73 | 0.998 | 0.038988 | -0.27768 | -0.001947 | 0.002964 | -0.2314 | -0.1947 | 8.2284
pf_sweep | 13 | 3 | 7 | 2 | 120 | 1 | 0.3 | 60 | 119.9834 | 1.0001 | 30.0763 | 0.03609 | -0.11444 | 0.12 | 119.776 | 2.978 | 0.12029 | -0.20738 | 1.9779 | 0.0842 | -0.1728 | -0.74362 | 11.101
pf_sweep | 13 | 3 | 8 | 2 | 120 | 1 | 0.3 | 60 | 119.9834 | 1.0001 | 30.0763 | 0.03609 | -0.11444 | 0.12 | 119.776 | 1.001 | 0.03956 | -0.20738 | 0.000896 | 0.00347 | -0.1728 | 0.0896 | 9.614
pf_sweep | 13 | 3 | 9 | 2 | 120 | 1 | 0.3 | 60 | 119.9834 | 1.0001 | 30.0763 | 0.03609 | -0.11444 | 0.12 | 119.776 | 1.001 | 0.039584 | -0.20738 | 0.000896 | 0.003494 | -0.1728 | 0.0896 | 9.6805
pf_sweep | 13 | 3 | 10 | 2 | 120 | 1 | 0.3 | 60 | 119.9834 | 1.0001 | 30.0763 | 0.03609 | -0.11444 | 0.12 | 119.776 | 1 | 0.039636 | -0.20738 | -0.000104 | 0.003546 | -0.1728 | -0.0104 | 9.8245
pf_sweep | 13 | 3 | 11 | 2 | 120 | 1 | 0.3 | 60 | 119.9834 | 1.0001 | 30.0763 | 0.03609 | -0.11444 | 0.12 | 119.776 | 1 | 0.039398 | -0.20738 | -0.000104 | 0.003308 | -0.1728 | -0.0104 | 9.1651
pf_sweep | 13 | 3 | 12 | 2 | 120 | 1 | 0.3 | 60 | 119.9834 | 1.0001 | 30.0763 | 0.03609 | -0.11444 | 0.12 | 119.776 | 1 | 0.039729 | -0.20738 | -0.000104 | 0.003639 | -0.1728 | -0.0104 | 10.0822
pf_sweep | 13 | 3 | 13 | 3 | 120 | 1 | 0.3 | 60 | 119.9892 | 0.99993 | 30.0114 | 0.036008 | -0.11445 | 0.11998 | 119.773 | 1 | 0.039431 | -0.21615 | 6.9e-05 | 0.003423 | -0.1801 | 0.0069 | 9.5065
pf_sweep | 13 | 3 | 14 | 3 | 120 | 1 | 0.3 | 60 | 119.9892 | 0.99993 | 30.0114 | 0.036008 | -0.11445 | 0.11998 | 119.773 | 0.999 | 0.039185 | -0.21615 | -0.000931 | 0.003177 | -0.1801 | -0.0931 | 8.8233
pf_sweep | 13 | 3 | 15 | 3 | 120 | 1 | 0.3 | 60 | 119.9892 | 0.99993 | 30.0114 | 0.036008 | -0.11445 | 0.11998 | 119.773 | 1.001 | 0.039376 | -0.21615 | 0.001069 | 0.003368 | -0.1801 | 0.1069 | 9.3537
pf_sweep | 13 | 3 | 16 | 3 | 120 | 1 | 0.3 | 60 | 119.9892 | 0.99993 | 30.0114 | 0.036008 | -0.11445 | 0.11998 | 119.773 | 0.999 | 0.03914 | -0.21615 | -0.000931 | 0.003132 | -0.1801 | -0.0931 | 8.6983
pf_sweep | 13 | 3 | 17 | 3 | 120 | 1 | 0.3 | 60 | 119.9892 | 0.99993 | 30.0114 | 0.036008 | -0.11445 | 0.11998 | 119.773 | 1 | 0.0393 | -0.21615 | 6.9e-05 | 0.003292 | -0.1801 | 0.0069 | 9.1427
pf_sweep | 13 | 3 | 18 | 3 | 120 | 1 | 0.3 | 60 | 119.9892 | 0.99993 | 30.0114 | 0.036008 | -0.11445 | 0.11998 | 119.773 | 1 | 0.039225 | -0.21615 | 6.9e-05 | 0.003217 | -0.1801 | 0.0069 | 8.9344
pf_sweep | 14 | 1 | 1 | 1 | 120 | 1 | 0.4 | 60 | 119.9931 | 0.99999 | 40.0152 | 0.048015 | -0.10997 | 0.11999 | 119.757 | 2.968 | 0.15443 | -0.23613 | 1.968 | 0.10641 | -0.1968 | -1.0658 | 7.2085
pf_sweep | 14 | 1 | 2 | 1 | 120 | 1 | 0.4 | 60 | 119.9931 | 0.99999 | 40.0152 | 0.048015 | -0.10997 | 0.11999 | 119.757 | 1.001 | 0.051463 | -0.23613 | 0.001008 | 0.003448 | -0.1968 | 0.1008 | 7.1808
pf_sweep | 14 | 1 | 3 | 1 | 120 | 1 | 0.4 | 60 | 119.9931 | 0.99999 | 40.0152 | 0.048015 | -0.10997 | 0.11999 | 119.757 | 1.001 | 0.051507 | -0.23613 | 0.001008 | 0.003492 | -0.1968 | 0.1008 | 7.2724
pf_sweep | 14 | 1 | 4 | 1 | 120 | 1 | 0.4 | 60 | 119.9931 | 0.99999 | 40.0152 | 0.048015 | -0.10997 | 0.11999 | 119.757 | 1 | 0.051508 | -0.23613 | 8e-06 | 0.003493 | -0.1968 | 0.0008 | 7.2745
pf_sweep | 14 | 1 | 5 | 1 | 120 | 1 | 0.4 | 60 | 119.9931 | 0.99999 | 40.0152 | 0.048015 | -0.10997 | 0.11999 | 119.757 | 1.001 | 0.051416 | -0.23613 | 0.001008 | 0.003401 | -0.1968 | 0.1008 | 7.0829
pf_sweep | 14 | 1 | 6 | 1 | 120 | 1 | 0.4 | 60 | 119.9931 | 0.99999 | 40.0152 | 0.048015 | -0.10997 | 0.11999 | 119.757 | 0.998 | 0.050849 | -0.23613 | -0.001992 | 0.002834 | -0.1968 | -0.1992 | 5.902
pf_sweep | 14 | 1 | 7 | 2 | 120 | 1 | 0.4 | 60 | 119.984 | 1.0001 | 40.0741 | 0.048087 | -0.10994 | 0.12 | 119.772 | 2.978 | 0.15569 | -0.21196 | 1.9779 | 0.10761 | -0.1767 | -0.74293 | 7.9248
pf_sweep | 14 | 1 | 8 | 2 | 120 | 1 | 0.4 | 60 | 119.984 | 1.0001 | 40.0741 | 0.048087 | -0.10994 | 0.12 | 119.772 | 1.001 | 0.051427 | -0.21196 | 0.000903 | 0.00334 | -0.1767 | 0.0903 | 6.9454
pf_sweep | 14 | 1 | 9 | 2 | 120 | 1 | 0.4 | 60 | 119.984 | 1.0001 | 40.0741 | 0.048087 | -0.10994 | 0.12 | 119.772 | 1.001 | 0.051513 | -0.21196 | 0.000903 | 0.003426 | -0.1767 | 0.0903 | 7.1242
pf_sweep | 14 | 1 | 10 | 2 | 120 | 1 | 0.4 | 60 | 119.984 | 1.0001 | 40.0741 | 0.048087 | -0.10994 | 0.12 | 119.772 | 1 | 0.051565 | -0.21196 | -9.7e-05 | 0.003478 | -0.1767 | -0.0097 | 7.2323
pf_sweep | 14 | 1 | 11 | 2 | 120 | 1 | 0.4 | 60 | 119.984 | 1.0001 | 40.0741 | 0.048087 | -0.10994 | 0.12 | 119.772 | 1 | 0.051256 | -0.21196 | -9.7e-05 | 0.003169 | -0.1767 | -0.0097 | 6.5897
pf_sweep | 14 | 1 | 12 | 2 | 120 | 1 | 0.4 | 60 | 119.984 | 1.0001 | 40.0741 | 0.048087 | -0.10994 | 0.12 | 119.772 | 1 | 0.051646 | -0.21196 | -9.7e-05 | 0.003559 | -0.1767 | -0.0097 | 7.4008
pf_sweep | 14 | 1 | 13 | 3 | 120 | 1 | 0.4 | 60 | 119.9891 | 0.99992 | 40.0115 | 0.048006 | -0.10996 | 0.11998 | 119.779 | 1 | 0.051284 | -0.21009 | 8e-05 | 0.003278 | -0.1751 | 0.008 | 6.8292
pf_sweep | 14 | 1 | 14 | 3 | 120 | 1 | 0.4 | 60 | 119.9891 | 0.99992 | 40.0115 | 0.048006 | -0.10996 | 0.11998 | 119.779 | 0.999 | 0.051115 | -0.21009 | -0.00092 | 0.003109 | -0.1751 | -0.0921 | 6.4772
pf_sweep | 14 | 1 | 15 | 3 | 120 | 1 | 0.4 | 60 | 119.9891 | 0.99992 | 40.0115 | 0.048006 | -0.10996 | 0.11998 | 119.779 | 1.001 | 0.051283 | -0.21009 | 0.00108 | 0.003277 | -0.1751 | 0.108 | 6.8271
pf_sweep | 14 | 1 | 16 | 3 | 120 | 1 | 0.4 | 60 | 119.9891 | 0.99992 | 40.0115 | 0.048006 | -0.10996 | 0.11998 | 119.779 | 0.999 | 0.050995 | -0.21009 | -0.00092 | 0.002989 | -0.1751 | -0.0921 | 6.2272
pf_sweep | 14 | 1 | 17 | 3 | 120 | 1 | 0.4 | 60 | 119.9891 | 0.99992 | 40.0115 | 0.048006 | -0.10996 | 0.11998 | 119.779 | 1 | 0.051215 | -0.21009 | 8e-05 | 0.003209 | -0.1751 | 0.008 | 6.6855
pf_sweep | 14 | 1 | 18 | 3 | 120 | 1 | 0.4 | 60 | 119.9891 | 0.99992 | 40.0115 | 0.048006 | -0.10996 | 0.11998 | 119.779 | 1 | 0.051145 | -0.21009 | 8e-05 | 0.003139 | -0.1751 | 0.008 | 6.5396
pf_sweep | 14 | 2 | 1 | 1 | 120 | 1 | 0.4 | 60 | 119.9971 | 0.99996 | 40.0166 | 0.048017 | -0.10997 | 0.11999 | 119.753 | 2.968 | 0.15447 | -0.24405 | 1.968 | 0.10646 | -0.2034 | -1.0626 | 7.2354
pf_sweep | 14 | 2 | 2 | 1 | 120 | 1 | 0.4 | 60 | 119.9971 | 0.99996 | 40.0166 | 0.048017 | -0.10997 | 0.11999 | 119.753 | 1.001 | 0.051467 | -0.24405 | 0.001041 | 0.00345 | -0.2034 | 0.1041 | 7.1855
pf_sweep | 14 | 2 | 3 | 1 | 120 | 1 | 0.4 | 60 | 119.9971 | 0.99996 | 40.0166 | 0.048017 | -0.10997 | 0.11999 | 119.753 | 1.001 | 0.051512 | -0.24405 | 0.001041 | 0.003495 | -0.2034 | 0.1041 | 7.2792
pf_sweep | 14 | 2 | 4 | 1 | 120 | 1 | 0.4 | 60 | 119.9971 | 0.99996 | 40.0166 | 0.048017 | -0.10997 | 0.11999 | 119.753 | 1 | 0.051515 | -0.24405 | 4.1e-05 | 0.003498 | -0.2034 | 0.0041 | 7.2854
pf_sweep | 14 | 2 | 5 | 1 | 120 | 1 | 0.4 | 60 | 119.9971 | 0.99996 | 40.0166 | 0.048017 | -0.10997 | 0.11999 | 119.753 | 1.001 | 0.051401 | -0.24405 | 0.001041 | 0.003384 | -0.2034 | 0.1041 | 7.048
pf_sweep | 14 | 2 | 6 | 1 | 120 | 1 | 0.4 | 60 | 119.9971 | 0.99996 | 40.0166 | 0.048017 | -0.10997 | 0.11999 | 119.753 | 0.999 | 0.050894 | -0.24405 | -0.000959 | 0.002877 | -0.2034 | -0.0959 | 5.9921
pf_sweep | 14 | 2 | 7 | 2 | 120 | 1 | 0.4 | 60 | 119.9823 | 1.0001 | 40.0762 | 0.048089 | -0.10994 | 0.11999 | 119.765 | 2.977 | 0.15554 | -0.21732 | 1.9769 | 0.10745 | -0.1811 | -0.77616 | 7.8127
pf_sweep | 14 | 2 | 8 | 2 | 120 | 1 | 0.4 | 60 | 119.9823 | 1.0001 | 40.0762 | 0.048089 | -0.10994 | 0.11999 | 119.765 | 1.001 | 0.05147 | -0.21732 | 0.000904 | 0.003381 | -0.1811 | 0.0904 | 7.0308
pf_sweep | 14 | 2 | 9 | 2 | 120 | 1 | 0.4 | 60 | 119.9823 | 1.0001 | 40.0762 | 0.048089 | -0.10994 | 0.11999 | 119.765 | 1.001 | 0.05151 | -0.21732 | 0.000904 | 0.003421 | -0.1811 | 0.0904 | 7.114
pf_sweep | 14 | 2 | 10 | 2 | 120 | 1 | 0.4 | 60 | 119.9823 | 1.0001 | 40.0762 | 0.048089 | -0.10994 | 0.11999 | 119.765 | 1 | 0.051523 | -0.21732 | -9.6e-05 | 0.003434 | -0.1811 | -0.0096 | 7.141
pf_sweep | 14 | 2 | 11 | 2 | 120 | 1 | 0.4 | 60 | 119.9823 | 1.0001 | 40.0762 | 0.048089 | -0.10994 | 0.11999 | 119.765 | 1 | 0.051289 | -0.21732 | -9.6e-05 | 0.0032 | -0.1811 | -0.0096 | 6.6544
pf_sweep | 14 | 2 | 12 | 2 | 120 | 1 | 0.4 | 60 | 119.9823 | 1.0001 | 40.0762 | 0.048089 | -0.10994 | 0.11999 | 119.765 | 1 | 0.051643 | -0.21732 | -9.6e-05 | 0.003554 | -0.1811 | -0.0096 | 7.3906
pf_sweep | 14 | 2 | 13 | 3 | 120 | 1 | 0.4 | 60 | 119.9895 | 0.99992 | 40.012 | 0.048006 | -0.10996 | 0.11998 | 119.767 | 1 | 0.051298 | -0.22249 | 7.9e-05 | 0.003292 | -0.1854 | 0.0079 | 6.8567
pf_sweep | 14 | 2 | 14 | 3 | 120 | 1 | 0.4 | 60 | 119.9895 | 0.99992 | 40.012 | 0.048006 | -0.10996 | 0.11998 | 119.767 | 0.999 | 0.051088 | -0.22249 | -0.000921 | 0.003082 | -0.1854 | -0.0921 | 6.4193
pf_sweep | 14 | 2 | 15 | 3 | 120 | 1 | 0.4 | 60 | 119.9895 | 0.99992 | 40.012 | 0.048006 | -0.10996 | 0.11998 | 119.767 | 1.001 | 0.051238 | -0.22249 | 0.001079 | 0.003232 | -0.1854 | 0.108 | 6.7317
pf_sweep | 14 | 2 | 16 | 3 | 120 | 1 | 0.4 | 60 | 119.9895 | 0.99992 | 40.012 | 0.048006 | -0.10996 | 0.11998 | 119.767 | 0.999 | 0.051028 | -0.22249 | -0.000921 | 0.003022 | -0.1854 | -0.0921 | 6.2943
pf_sweep | 14 | 2 | 17 | 3 | 120 | 1 | 0.4 | 60 | 119.9895 | 0.99992 | 40.012 | 0.048006 | -0.10996 | 0.11998 | 119.767 | 1 | 0.051197 | -0.22249 | 7.9e-05 | 0.003191 | -0.1854 | 0.0079 | 6.6463
pf_sweep | 14 | 2 | 18 | 3 | 120 | 1 | 0.4 | 60 | 119.9895 | 0.99992 | 40.012 | 0.048006 | -0.10996 | 0.11998 | 119.767 | 1 | 0.05111 | -0.22249 | 7.9e-05 | 0.003104 | -0.1854 | 0.0079 | 6.4651
pf_sweep | 14 | 3 | 1 | 1 | 120 | 1 | 0.4 | 60 | 120.0024 | 0.99998 | 40.018 | 0.048022 | -0.10997 | 0.12 | 119.782 | 2.968 | 0.15453 | -0.22039 | 1.968 | 0.10651 | -0.1837 | -1.0646 | 7.2671
pf_sweep | 14 | 3 | 2 | 1 | 120 | 1 | 0.4 | 60 | 120.0024 | 0.99998 | 40.018 | 0.048022 | -0.10997 | 0.12 | 119.782 | 1.001 | 0.051519 | -0.22039 | 0.00102 | 0.003497 | -0.1837 | 0.102 | 7.2831
pf_sweep | 14 | 3 | 3 | 1 | 120 | 1 | 0.4 | 60 | 120.0024 | 0.99998 | 40.018 | 0.048022 | -0.10997 | 0.12 | 119.782 | 1.001 | 0.05154 | -0.22039 | 0.00102 | 0.003518 | -0.1837 | 0.102 | 7.3268
pf_sweep | 14 | 3 | 4 | 1 | 120 | 1 | 0.4 | 60 | 120.0024 | 0.99998 | 40.018 | 0.048022 | -0.10997 | 0.12 | 119.782 | 1 | 0.051538 | -0.22039 | 2e-05 | 0.003516 | -0.1837 | 0.002 | 7.3227
pf_sweep | 14 | 3 | 5 | 1 | 120 | 1 | 0.4 | 60 | 120.0024 | 0.99998 | 40.018 | 0.048022 | -0.10997 | 0.12 | 119.782 | 1.001 | 0.051442 | -0.22039 | 0.00102 | 0.00342 | -0.1837 | 0.102 | 7.1228
pf_sweep | 14 | 3 | 6 | 1 | 120 | 1 | 0.4 | 60 | 120.0024 | 0.99998 | 40.018 | 0.048022 | -0.10997 | 0.12 | 119.782 | 0.998 | 0.050902 | -0.22039 | -0.00198 | 0.00288 | -0.1837 | -0.198 | 5.9983
pf_sweep | 14 | 3 | 7 | 2 | 120 | 1 | 0.4 | 60 | 119.9828 | 1.0001 | 40.0742 | 0.048088 | -0.10994 | 0.12 | 119.762 | 2.978 | 0.15558 | -0.22081 | 1.9779 | 0.10749 | -0.184 | -0.7446 | 7.8469
pf_sweep | 14 | 3 | 8 | 2 | 120 | 1 | 0.4 | 60 | 119.9828 | 1.0001 | 40.0742 | 0.048088 | -0.10994 | 0.12 | 119.762 | 1.001 | 0.051452 | -0.22081 | 0.000887 | 0.003364 | -0.184 | 0.0886 | 6.9964
pf_sweep | 14 | 3 | 9 | 2 | 120 | 1 | 0.4 | 60 | 119.9828 | 1.0001 | 40.0742 | 0.048088 | -0.10994 | 0.12 | 119.762 | 1.001 | 0.051499 | -0.22081 | 0.000887 | 0.003411 | -0.184 | 0.0886 | 7.0941
pf_sweep | 14 | 3 | 10 | 2 | 120 | 1 | 0.4 | 60 | 119.9828 | 1.0001 | 40.0742 | 0.048088 | -0.10994 | 0.12 | 119.762 | 1 | 0.051512 | -0.22081 | -0.000113 | 0.003424 | -0.184 | -0.0113 | 7.1211
pf_sweep | 14 | 3 | 11 | 2 | 120 | 1 | 0.4 | 60 | 119.9828 | 1.0001 | 40.0742 | 0.048088 | -0.10994 | 0.12 | 119.762 | 0.999 | 0.051271 | -0.22081 | -0.001113 | 0.003183 | -0.184 | -0.1113 | 6.62
pf_sweep | 14 | 3 | 12 | 2 | 120 | 1 | 0.4 | 60 | 119.9828 | 1.0001 | 40.0742 | 0.048088 | -0.10994 | 0.12 | 119.762 | 1 | 0.051633 | -0.22081 | -0.000113 | 0.003545 | -0.184 | -0.0113 | 7.3728
pf_sweep | 14 | 3 | 13 | 3 | 120 | 1 | 0.4 | 60 | 119.9922 | 0.99992 | 40.0116 | 0.048007 | -0.10996 | 0.11998 | 119.771 | 1 | 0.0513 | -0.22117 | 7.7e-05 | 0.003293 | -0.1843 | 0.0077 | 6.8593
pf_sweep | 14 | 3 | 14 | 3 | 120 | 1 | 0.4 | 60 | 119.9922 | 0.99992 | 40.0116 | 0.048007 | -0.10996 | 0.11998 | 119.771 | 0.999 | 0.051081 | -0.22117 | -0.000923 | 0.003074 | -0.1843 | -0.0923 | 6.4031
pf_sweep | 14 | 3 | 15 | 3 | 120 | 1 | 0.4 | 60 | 119.9922 | 0.99992 | 40.0116 | 0.048007 | -0.10996 | 0.11998 | 119.771 | 1.001 | 0.051247 | -0.22117 | 0.001077 | 0.00324 | -0.1843 | 0.1078 | 6.7489
pf_sweep | 14 | 3 | 16 | 3 | 120 | 1 | 0.4 | 60 | 119.9922 | 0.99992 | 40.0116 | 0.048007 | -0.10996 | 0.11998 | 119.771 | 0.999 | 0.05102 | -0.22117 | -0.000923 | 0.003013 | -0.1843 | -0.0923 | 6.2761
pf_sweep | 14 | 3 | 17 | 3 | 120 | 1 | 0.4 | 60 | 119.9922 | 0.99992 | 40.0116 | 0.048007 | -0.10996 | 0.11998 | 119.771 | 1 | 0.05118 | -0.22117 | 7.7e-05 | 0.003173 | -0.1843 | 0.0077 | 6.6093
pf_sweep | 14 | 3 | 18 | 3 | 120 | 1 | 0.4 | 60 | 119.9922 | 0.99992 | 40.0116 | 0.048007 | -0.10996 | 0.11998 | 119.771 | 1 | 0.051109 | -0.22117 | 7.7e-05 | 0.003102 | -0.1843 | 0.0077 | 6.4614
pf_sweep | 15 | 1 | 1 | 1 | 120 | 1 | 0.499 | 60 | 120.0003 | 0.99997 | 49.9158 | 0.059897 | -0.10398 | 0.12 | 119.754 | 2.968 | 0.18893 | -0.24627 | 1.968 | 0.12903 | -0.2052 | -1.0639 | 5.1402
pf_sweep | 15 | 1 | 2 | 1 | 120 | 1 | 0.499 | 60 | 120.0003 | 0.99997 | 49.9158 | 0.059897 | -0.10398 | 0.12 | 119.754 | 1.001 | 0.063223 | -0.24627 | 0.001028 | 0.003326 | -0.2052 | 0.1028 | 5.552
pf_sweep | 15 | 1 | 3 | 1 | 120 | 1 | 0.499 | 60 | 120.0003 | 0.99997 | 49.9158 | 0.059897 | -0.10398 | 0.12 | 119.754 | 1.001 | 0.063273 | -0.24627 | 0.001028 | 0.003376 | -0.2052 | 0.1028 | 5.6355
pf_sweep | 15 | 1 | 4 | 1 | 120 | 1 | 0.499 | 60 | 120.0003 | 0.99997 | 49.9158 | 0.059897 | -0.10398 | 0.12 | 119.754 | 1 | 0.063149 | -0.24627 | 2.8e-05 | 0.003252 | -0.2052 | 0.0028 | 5.4285
pf_sweep | 15 | 1 | 5 | 1 | 120 | 1 | 0.499 | 60 | 120.0003 | 0.99997 | 49.9158 | 0.059897 | -0.10398 | 0.12 | 119.754 | 1.001 | 0.063151 | -0.24627 | 0.001028 | 0.003254 | -0.2052 | 0.1028 | 5.4318
pf_sweep | 15 | 1 | 6 | 1 | 120 | 1 | 0.499 | 60 | 120.0003 | 0.99997 | 49.9158 | 0.059897 | -0.10398 | 0.12 | 119.754 | 0.999 | 0.062648 | -0.24627 | -0.000972 | 0.002751 | -0.2052 | -0.0972 | 4.592
pf_sweep | 15 | 1 | 7 | 2 | 120 | 1 | 0.499 | 60 | 119.9841 | 1.0001 | 49.968 | 0.05996 | -0.10394 | 0.12 | 119.774 | 2.978 | 0.19038 | -0.21014 | 1.9779 | 0.13042 | -0.1751 | -0.74306 | 5.8381
pf_sweep | 15 | 1 | 8 | 2 | 120 | 1 | 0.499 | 60 | 119.9841 | 1.0001 | 49.968 | 0.05996 | -0.10394 | 0.12 | 119.774 | 1.001 | 0.063226 | -0.21014 | 0.000902 | 0.003266 | -0.1751 | 0.0902 | 5.4478
pf_sweep | 15 | 1 | 9 | 2 | 120 | 1 | 0.499 | 60 | 119.9841 | 1.0001 | 49.968 | 0.05996 | -0.10394 | 0.12 | 119.774 | 1 | 0.063249 | -0.21014 | -9.8e-05 | 0.003289 | -0.1751 | -0.0098 | 5.4862
pf_sweep | 15 | 1 | 10 | 2 | 120 | 1 | 0.499 | 60 | 119.9841 | 1.0001 | 49.968 | 0.05996 | -0.10394 | 0.12 | 119.774 | 1 | 0.063271 | -0.21014 | -9.8e-05 | 0.003311 | -0.1751 | -0.0098 | 5.5229
pf_sweep | 15 | 1 | 11 | 2 | 120 | 1 | 0.499 | 60 | 119.9841 | 1.0001 | 49.968 | 0.05996 | -0.10394 | 0.12 | 119.774 | 1 | 0.063044 | -0.21014 | -9.8e-05 | 0.003084 | -0.1751 | -0.0098 | 5.1443
pf_sweep | 15 | 1 | 12 | 2 | 120 | 1 | 0.499 | 60 | 119.9841 | 1.0001 | 49.968 | 0.05996 | -0.10394 | 0.12 | 119.774 | 1 | 0.06337 | -0.21014 | -9.8e-05 | 0.00341 | -0.1751 | -0.0098 | 5.688
pf_sweep | 15 | 1 | 13 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99993 | 49.9137 | 0.05989 | -0.10397 | 0.11999 | 119.77 | 1 | 0.063043 | -0.22534 | 7e-05 | 0.003153 | -0.1878 | 0.007 | 5.2648
pf_sweep | 15 | 1 | 14 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99993 | 49.9137 | 0.05989 | -0.10397 | 0.11999 | 119.77 | 0.999 | 0.062829 | -0.22534 | -0.00093 | 0.002939 | -0.1878 | -0.093 | 4.9075
pf_sweep | 15 | 1 | 15 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99993 | 49.9137 | 0.05989 | -0.10397 | 0.11999 | 119.77 | 1 | 0.062995 | -0.22534 | 7e-05 | 0.003105 | -0.1878 | 0.007 | 5.1846
pf_sweep | 15 | 1 | 16 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99993 | 49.9137 | 0.05989 | -0.10397 | 0.11999 | 119.77 | 0.999 | 0.062784 | -0.22534 | -0.00093 | 0.002894 | -0.1878 | -0.093 | 4.8323
pf_sweep | 15 | 1 | 17 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99993 | 49.9137 | 0.05989 | -0.10397 | 0.11999 | 119.77 | 1 | 0.06293 | -0.22534 | 7e-05 | 0.00304 | -0.1878 | 0.007 | 5.0761
pf_sweep | 15 | 1 | 18 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99993 | 49.9137 | 0.05989 | -0.10397 | 0.11999 | 119.77 | 1 | 0.06286 | -0.22534 | 7e-05 | 0.00297 | -0.1878 | 0.007 | 4.9592
pf_sweep | 15 | 2 | 1 | 1 | 120 | 1 | 0.499 | 60 | 120.0108 | 0.99996 | 49.9192 | 0.059906 | -0.10398 | 0.12001 | 119.741 | 2.968 | 0.189 | -0.26979 | 1.968 | 0.1291 | -0.2248 | -1.0626 | 5.167
pf_sweep | 15 | 2 | 2 | 1 | 120 | 1 | 0.499 | 60 | 120.0108 | 0.99996 | 49.9192 | 0.059906 | -0.10398 | 0.12001 | 119.741 | 1.001 | 0.06322 | -0.26979 | 0.001041 | 0.003314 | -0.2248 | 0.1041 | 5.5321
pf_sweep | 15 | 2 | 3 | 1 | 120 | 1 | 0.499 | 60 | 120.0108 | 0.99996 | 49.9192 | 0.059906 | -0.10398 | 0.12001 | 119.741 | 1.001 | 0.063256 | -0.26979 | 0.001041 | 0.00335 | -0.2248 | 0.1041 | 5.5921
pf_sweep | 15 | 2 | 4 | 1 | 120 | 1 | 0.499 | 60 | 120.0108 | 0.99996 | 49.9192 | 0.059906 | -0.10398 | 0.12001 | 119.741 | 1 | 0.063189 | -0.26979 | 4.1e-05 | 0.003283 | -0.2248 | 0.0041 | 5.4803
pf_sweep | 15 | 2 | 5 | 1 | 120 | 1 | 0.499 | 60 | 120.0108 | 0.99996 | 49.9192 | 0.059906 | -0.10398 | 0.12001 | 119.741 | 1.001 | 0.063167 | -0.26979 | 0.001041 | 0.003261 | -0.2248 | 0.1041 | 5.4436
pf_sweep | 15 | 2 | 6 | 1 | 120 | 1 | 0.499 | 60 | 120.0108 | 0.99996 | 49.9192 | 0.059906 | -0.10398 | 0.12001 | 119.741 | 0.998 | 0.062584 | -0.26979 | -0.001959 | 0.002678 | -0.2248 | -0.1959 | 4.4704
pf_sweep | 15 | 2 | 7 | 2 | 120 | 1 | 0.499 | 60 | 119.9824 | 1.0001 | 49.9691 | 0.059961 | -0.10394 | 0.12 | 119.757 | 2.978 | 0.19045 | -0.22541 | 1.9779 | 0.13049 | -0.1879 | -0.74423 | 5.8753
pf_sweep | 15 | 2 | 8 | 2 | 120 | 1 | 0.499 | 60 | 119.9824 | 1.0001 | 49.9691 | 0.059961 | -0.10394 | 0.12 | 119.757 | 1.001 | 0.063185 | -0.22541 | 0.00089 | 0.003224 | -0.1879 | 0.089 | 5.3772
pf_sweep | 15 | 2 | 9 | 2 | 120 | 1 | 0.499 | 60 | 119.9824 | 1.0001 | 49.9691 | 0.059961 | -0.10394 | 0.12 | 119.757 | 1.001 | 0.063231 | -0.22541 | 0.00089 | 0.00327 | -0.1879 | 0.089 | 5.454
pf_sweep | 15 | 2 | 10 | 2 | 120 | 1 | 0.499 | 60 | 119.9824 | 1.0001 | 49.9691 | 0.059961 | -0.10394 | 0.12 | 119.757 | 1 | 0.063278 | -0.22541 | -0.00011 | 0.003317 | -0.1879 | -0.011 | 5.5323
pf_sweep | 15 | 2 | 11 | 2 | 120 | 1 | 0.499 | 60 | 119.9824 | 1.0001 | 49.9691 | 0.059961 | -0.10394 | 0.12 | 119.757 | 0.999 | 0.063013 | -0.22541 | -0.00111 | 0.003052 | -0.1879 | -0.111 | 5.0904
pf_sweep | 15 | 2 | 12 | 2 | 120 | 1 | 0.499 | 60 | 119.9824 | 1.0001 | 49.9691 | 0.059961 | -0.10394 | 0.12 | 119.757 | 1 | 0.063352 | -0.22541 | -0.00011 | 0.003391 | -0.1879 | -0.011 | 5.6558
pf_sweep | 15 | 2 | 13 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99994 | 49.9141 | 0.059891 | -0.10397 | 0.11999 | 119.777 | 1 | 0.06301 | -0.21832 | 6.1e-05 | 0.003119 | -0.1819 | 0.0061 | 5.208
pf_sweep | 15 | 2 | 14 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99994 | 49.9141 | 0.059891 | -0.10397 | 0.11999 | 119.777 | 0.999 | 0.062802 | -0.21832 | -0.000939 | 0.002911 | -0.1819 | -0.0939 | 4.8607
pf_sweep | 15 | 2 | 15 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99994 | 49.9141 | 0.059891 | -0.10397 | 0.11999 | 119.777 | 1.001 | 0.063028 | -0.21832 | 0.001061 | 0.003137 | -0.1819 | 0.1061 | 5.238
pf_sweep | 15 | 2 | 16 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99994 | 49.9141 | 0.059891 | -0.10397 | 0.11999 | 119.777 | 0.999 | 0.06275 | -0.21832 | -0.000939 | 0.002859 | -0.1819 | -0.0939 | 4.7738
pf_sweep | 15 | 2 | 17 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99994 | 49.9141 | 0.059891 | -0.10397 | 0.11999 | 119.777 | 1 | 0.062933 | -0.21832 | 6.1e-05 | 0.003042 | -0.1819 | 0.0061 | 5.0794
pf_sweep | 15 | 2 | 18 | 3 | 120 | 1 | 0.499 | 60 | 119.9953 | 0.99994 | 49.9141 | 0.059891 | -0.10397 | 0.11999 | 119.777 | 1 | 0.062863 | -0.21832 | 6.1e-05 | 0.002972 | -0.1819 | 0.0061 | 4.9625
pf_sweep | 15 | 3 | 1 | 1 | 120 | 1 | 0.499 | 60 | 120.0053 | 0.99998 | 49.9179 | 0.059903 | -0.10398 | 0.12 | 119.726 | 2.968 | 0.18913 | -0.27932 | 1.968 | 0.12923 | -0.2328 | -1.0651 | 5.2425
pf_sweep | 15 | 3 | 2 | 1 | 120 | 1 | 0.499 | 60 | 120.0053 | 0.99998 | 49.9179 | 0.059903 | -0.10398 | 0.12 | 119.726 | 1.001 | 0.063241 | -0.27932 | 0.001015 | 0.003338 | -0.2328 | 0.1015 | 5.572
pf_sweep | 15 | 3 | 3 | 1 | 120 | 1 | 0.499 | 60 | 120.0053 | 0.99998 | 49.9179 | 0.059903 | -0.10398 | 0.12 | 119.726 | 1.001 | 0.063239 | -0.27932 | 0.001015 | 0.003336 | -0.2328 | 0.1015 | 5.5686
pf_sweep | 15 | 3 | 4 | 1 | 120 | 1 | 0.499 | 60 | 120.0053 | 0.99998 | 49.9179 | 0.059903 | -0.10398 | 0.12 | 119.726 | 1 | 0.063226 | -0.27932 | 1.5e-05 | 0.003323 | -0.2328 | 0.0015 | 5.5469
pf_sweep | 15 | 3 | 5 | 1 | 120 | 1 | 0.499 | 60 | 120.0053 | 0.99998 | 49.9179 | 0.059903 | -0.10398 | 0.12 | 119.726 | 1.001 | 0.063179 | -0.27932 | 0.001015 | 0.003276 | -0.2328 | 0.1015 | 5.4685
pf_sweep | 15 | 3 | 6 | 1 | 120 | 1 | 0.499 | 60 | 120.0053 | 0.99998 | 49.9179 | 0.059903 | -0.10398 | 0.12 | 119.726 | 0.998 | 0.062576 | -0.27932 | -0.001985 | 0.002673 | -0.2328 | -0.1985 | 4.4618
pf_sweep | 15 | 3 | 7 | 2 | 120 | 1 | 0.499 | 60 | 119.9837 | 1.0001 | 49.9687 | 0.059961 | -0.10394 | 0.12 | 119.76 | 2.978 | 0.1904 | -0.22372 | 1.9779 | 0.13044 | -0.1865 | -0.7441 | 5.8481
pf_sweep | 15 | 3 | 8 | 2 | 120 | 1 | 0.499 | 60 | 119.9837 | 1.0001 | 49.9687 | 0.059961 | -0.10394 | 0.12 | 119.76 | 1.001 | 0.063171 | -0.22372 | 0.000892 | 0.00321 | -0.1865 | 0.0891 | 5.3539
pf_sweep | 15 | 3 | 9 | 2 | 120 | 1 | 0.499 | 60 | 119.9837 | 1.0001 | 49.9687 | 0.059961 | -0.10394 | 0.12 | 119.76 | 1.001 | 0.063252 | -0.22372 | 0.000892 | 0.003291 | -0.1865 | 0.0891 | 5.489
pf_sweep | 15 | 3 | 10 | 2 | 120 | 1 | 0.499 | 60 | 119.9837 | 1.0001 | 49.9687 | 0.059961 | -0.10394 | 0.12 | 119.76 | 1 | 0.063268 | -0.22372 | -0.000108 | 0.003307 | -0.1865 | -0.0108 | 5.5157
pf_sweep | 15 | 3 | 11 | 2 | 120 | 1 | 0.499 | 60 | 119.9837 | 1.0001 | 49.9687 | 0.059961 | -0.10394 | 0.12 | 119.76 | 1 | 0.062989 | -0.22372 | -0.000108 | 0.003028 | -0.1865 | -0.0108 | 5.0504
pf_sweep | 15 | 3 | 12 | 2 | 120 | 1 | 0.499 | 60 | 119.9837 | 1.0001 | 49.9687 | 0.059961 | -0.10394 | 0.12 | 119.76 | 1 | 0.063363 | -0.22372 | -0.000108 | 0.003402 | -0.1865 | -0.0108 | 5.6741
pf_sweep | 15 | 3 | 13 | 3 | 120 | 1 | 0.499 | 60 | 119.9932 | 0.99994 | 49.91 | 0.059885 | -0.10397 | 0.11999 | 119.782 | 1 | 0.063015 | -0.21124 | 6.4e-05 | 0.00313 | -0.176 | 0.0064 | 5.2271
pf_sweep | 15 | 3 | 14 | 3 | 120 | 1 | 0.499 | 60 | 119.9932 | 0.99994 | 49.91 | 0.059885 | -0.10397 | 0.11999 | 119.782 | 0.999 | 0.062835 | -0.21124 | -0.000936 | 0.00295 | -0.176 | -0.0936 | 4.9265
pf_sweep | 15 | 3 | 15 | 3 | 120 | 1 | 0.499 | 60 | 119.9932 | 0.99994 | 49.91 | 0.059885 | -0.10397 | 0.11999 | 119.782 | 1.001 | 0.06304 | -0.21124 | 0.001064 | 0.003155 | -0.176 | 0.1064 | 5.2688
pf_sweep | 15 | 3 | 16 | 3 | 120 | 1 | 0.499 | 60 | 119.9932 | 0.99994 | 49.91 | 0.059885 | -0.10397 | 0.11999 | 119.782 | 0.999 | 0.062736 | -0.21124 | -0.000936 | 0.002851 | -0.176 | -0.0936 | 4.7612
pf_sweep | 15 | 3 | 17 | 3 | 120 | 1 | 0.499 | 60 | 119.9932 | 0.99994 | 49.91 | 0.059885 | -0.10397 | 0.11999 | 119.782 | 1 | 0.062956 | -0.21124 | 6.4e-05 | 0.003071 | -0.176 | 0.0064 | 5.1286
pf_sweep | 15 | 3 | 18 | 3 | 120 | 1 | 0.499 | 60 | 119.9932 | 0.99994 | 49.91 | 0.059885 | -0.10397 | 0.11999 | 119.782 | 1 | 0.062876 | -0.21124 | 6.4e-05 | 0.002991 | -0.176 | 0.0064 | 4.995
pf_sweep | 16 | 1 | 1 | 1 | 120 | 1 | 0.599 | 60 | 120.0023 | 0.99996 | 59.9169 | 0.071899 | -0.096073 | 0.12 | 119.704 | 2.968 | 0.22379 | -0.29832 | 1.968 | 0.15189 | -0.2486 | -1.0629 | 3.7512
pf_sweep | 16 | 1 | 2 | 1 | 120 | 1 | 0.599 | 60 | 120.0023 | 0.99996 | 59.9169 | 0.071899 | -0.096073 | 0.12 | 119.704 | 1.001 | 0.074966 | -0.29832 | 0.001038 | 0.003067 | -0.2486 | 0.1038 | 4.2658
pf_sweep | 16 | 1 | 3 | 1 | 120 | 1 | 0.599 | 60 | 120.0023 | 0.99996 | 59.9169 | 0.071899 | -0.096073 | 0.12 | 119.704 | 1.001 | 0.074997 | -0.29832 | 0.001038 | 0.003098 | -0.2486 | 0.1038 | 4.3089
pf_sweep | 16 | 1 | 4 | 1 | 120 | 1 | 0.599 | 60 | 120.0023 | 0.99996 | 59.9169 | 0.071899 | -0.096073 | 0.12 | 119.704 | 1 | 0.074959 | -0.29832 | 3.8e-05 | 0.00306 | -0.2486 | 0.0038 | 4.2561
pf_sweep | 16 | 1 | 5 | 1 | 120 | 1 | 0.599 | 60 | 120.0023 | 0.99996 | 59.9169 | 0.071899 | -0.096073 | 0.12 | 119.704 | 1.001 | 0.074917 | -0.29832 | 0.001038 | 0.003018 | -0.2486 | 0.1038 | 4.1977
pf_sweep | 16 | 1 | 6 | 1 | 120 | 1 | 0.599 | 60 | 120.0023 | 0.99996 | 59.9169 | 0.071899 | -0.096073 | 0.12 | 119.704 | 0.999 | 0.074368 | -0.29832 | -0.000962 | 0.002469 | -0.2486 | -0.0962 | 3.4341
pf_sweep | 16 | 1 | 7 | 2 | 120 | 1 | 0.599 | 60 | 119.9812 | 1.0001 | 59.9638 | 0.071954 | -0.096029 | 0.12 | 119.772 | 2.978 | 0.22538 | -0.2092 | 1.9779 | 0.15342 | -0.1744 | -0.7455 | 4.4074
pf_sweep | 16 | 1 | 8 | 2 | 120 | 1 | 0.599 | 60 | 119.9812 | 1.0001 | 59.9638 | 0.071954 | -0.096029 | 0.12 | 119.772 | 1.001 | 0.075003 | -0.2092 | 0.000877 | 0.003049 | -0.1744 | 0.0877 | 4.2373
pf_sweep | 16 | 1 | 9 | 2 | 120 | 1 | 0.599 | 60 | 119.9812 | 1.0001 | 59.9638 | 0.071954 | -0.096029 | 0.12 | 119.772 | 1.001 | 0.075066 | -0.2092 | 0.000877 | 0.003112 | -0.1744 | 0.0877 | 4.3249
pf_sweep | 16 | 1 | 10 | 2 | 120 | 1 | 0.599 | 60 | 119.9812 | 1.0001 | 59.9638 | 0.071954 | -0.096029 | 0.12 | 119.772 | 1 | 0.075077 | -0.2092 | -0.000123 | 0.003123 | -0.1744 | -0.0123 | 4.3402
pf_sweep | 16 | 1 | 11 | 2 | 120 | 1 | 0.599 | 60 | 119.9812 | 1.0001 | 59.9638 | 0.071954 | -0.096029 | 0.12 | 119.772 | 1 | 0.074831 | -0.2092 | -0.000123 | 0.002877 | -0.1744 | -0.0123 | 3.9983
pf_sweep | 16 | 1 | 12 | 2 | 120 | 1 | 0.599 | 60 | 119.9812 | 1.0001 | 59.9638 | 0.071954 | -0.096029 | 0.12 | 119.772 | 1 | 0.075175 | -0.2092 | -0.000123 | 0.003221 | -0.1744 | -0.0123 | 4.4764
pf_sweep | 16 | 1 | 13 | 3 | 120 | 1 | 0.599 | 60 | 119.9925 | 0.99993 | 59.9093 | 0.071882 | -0.096069 | 0.11998 | 119.765 | 1 | 0.074799 | -0.22753 | 6.9e-05 | 0.002917 | -0.1896 | 0.0069 | 4.0585
pf_sweep | 16 | 1 | 14 | 3 | 120 | 1 | 0.599 | 60 | 119.9925 | 0.99993 | 59.9093 | 0.071882 | -0.096069 | 0.11998 | 119.765 | 0.999 | 0.074648 | -0.22753 | -0.000931 | 0.002766 | -0.1896 | -0.0931 | 3.8485
pf_sweep | 16 | 1 | 15 | 3 | 120 | 1 | 0.599 | 60 | 119.9925 | 0.99993 | 59.9093 | 0.071882 | -0.096069 | 0.11998 | 119.765 | 1 | 0.074813 | -0.22753 | 6.9e-05 | 0.002931 | -0.1896 | 0.0069 | 4.078
pf_sweep | 16 | 1 | 16 | 3 | 120 | 1 | 0.599 | 60 | 119.9925 | 0.99993 | 59.9093 | 0.071882 | -0.096069 | 0.11998 | 119.765 | 0.999 | 0.074561 | -0.22753 | -0.000931 | 0.002679 | -0.1896 | -0.0931 | 3.7274
pf_sweep | 16 | 1 | 17 | 3 | 120 | 1 | 0.599 | 60 | 119.9925 | 0.99993 | 59.9093 | 0.071882 | -0.096069 | 0.11998 | 119.765 | 1 | 0.074763 | -0.22753 | 6.9e-05 | 0.002881 | -0.1896 | 0.0069 | 4.0085
pf_sweep | 16 | 1 | 18 | 3 | 120 | 1 | 0.599 | 60 | 119.9925 | 0.99993 | 59.9093 | 0.071882 | -0.096069 | 0.11998 | 119.765 | 1 | 0.074662 | -0.22753 | 6.9e-05 | 0.00278 | -0.1896 | 0.0069 | 3.868
pf_sweep | 16 | 2 | 1 | 1 | 120 | 1 | 0.599 | 60 | 120.0125 | 0.99998 | 59.9163 | 0.071905 | -0.096083 | 0.12001 | 119.731 | 2.968 | 0.22385 | -0.28151 | 1.968 | 0.15194 | -0.2346 | -1.0644 | 3.7692
pf_sweep | 16 | 2 | 2 | 1 | 120 | 1 | 0.599 | 60 | 120.0125 | 0.99998 | 59.9163 | 0.071905 | -0.096083 | 0.12001 | 119.731 | 1.001 | 0.074976 | -0.28151 | 0.001023 | 0.003071 | -0.2346 | 0.1023 | 4.2703
pf_sweep | 16 | 2 | 3 | 1 | 120 | 1 | 0.599 | 60 | 120.0125 | 0.99998 | 59.9163 | 0.071905 | -0.096083 | 0.12001 | 119.731 | 1.001 | 0.07505 | -0.28151 | 0.001023 | 0.003145 | -0.2346 | 0.1023 | 4.3732
pf_sweep | 16 | 2 | 4 | 1 | 120 | 1 | 0.599 | 60 | 120.0125 | 0.99998 | 59.9163 | 0.071905 | -0.096083 | 0.12001 | 119.731 | 1 | 0.074969 | -0.28151 | 2.3e-05 | 0.003064 | -0.2346 | 0.0023 | 4.2606
pf_sweep | 16 | 2 | 5 | 1 | 120 | 1 | 0.599 | 60 | 120.0125 | 0.99998 | 59.9163 | 0.071905 | -0.096083 | 0.12001 | 119.731 | 1.001 | 0.074927 | -0.28151 | 0.001023 | 0.003022 | -0.2346 | 0.1023 | 4.2022
pf_sweep | 16 | 2 | 6 | 1 | 120 | 1 | 0.599 | 60 | 120.0125 | 0.99998 | 59.9163 | 0.071905 | -0.096083 | 0.12001 | 119.731 | 0.998 | 0.074382 | -0.28151 | -0.001977 | 0.002477 | -0.2346 | -0.1977 | 3.4442
pf_sweep | 16 | 2 | 7 | 2 | 120 | 1 | 0.599 | 60 | 119.9818 | 1.0001 | 59.9669 | 0.071957 | -0.096025 | 0.11999 | 119.756 | 2.978 | 0.22531 | -0.2258 | 1.9779 | 0.15335 | -0.1882 | -0.74375 | 4.3703
pf_sweep | 16 | 2 | 8 | 2 | 120 | 1 | 0.599 | 60 | 119.9818 | 1.0001 | 59.9669 | 0.071957 | -0.096025 | 0.11999 | 119.756 | 1.001 | 0.074981 | -0.2258 | 0.000895 | 0.003024 | -0.1882 | 0.0895 | 4.2026
pf_sweep | 16 | 2 | 9 | 2 | 120 | 1 | 0.599 | 60 | 119.9818 | 1.0001 | 59.9669 | 0.071957 | -0.096025 | 0.11999 | 119.756 | 1 | 0.075018 | -0.2258 | -0.000105 | 0.003061 | -0.1882 | -0.0105 | 4.254
pf_sweep | 16 | 2 | 10 | 2 | 120 | 1 | 0.599 | 60 | 119.9818 | 1.0001 | 59.9669 | 0.071957 | -0.096025 | 0.11999 | 119.756 | 1 | 0.07505 | -0.2258 | -0.000105 | 0.003093 | -0.1882 | -0.0105 | 4.2985
pf_sweep | 16 | 2 | 11 | 2 | 120 | 1 | 0.599 | 60 | 119.9818 | 1.0001 | 59.9669 | 0.071957 | -0.096025 | 0.11999 | 119.756 | 1 | 0.074819 | -0.2258 | -0.000105 | 0.002862 | -0.1882 | -0.0105 | 3.9775
pf_sweep | 16 | 2 | 12 | 2 | 120 | 1 | 0.599 | 60 | 119.9818 | 1.0001 | 59.9669 | 0.071957 | -0.096025 | 0.11999 | 119.756 | 1 | 0.075117 | -0.2258 | -0.000105 | 0.00316 | -0.1882 | -0.0105 | 4.3916
pf_sweep | 16 | 2 | 13 | 3 | 120 | 1 | 0.599 | 60 | 119.991 | 0.99992 | 59.9099 | 0.071881 | -0.096066 | 0.11998 | 119.784 | 1 | 0.074852 | -0.20699 | 8.2e-05 | 0.002971 | -0.1725 | 0.0082 | 4.1338
pf_sweep | 16 | 2 | 14 | 3 | 120 | 1 | 0.599 | 60 | 119.991 | 0.99992 | 59.9099 | 0.071881 | -0.096066 | 0.11998 | 119.784 | 0.999 | 0.074649 | -0.20699 | -0.000918 | 0.002768 | -0.1725 | -0.0918 | 3.8514
pf_sweep | 16 | 2 | 15 | 3 | 120 | 1 | 0.599 | 60 | 119.991 | 0.99992 | 59.9099 | 0.071881 | -0.096066 | 0.11998 | 119.784 | 1.001 | 0.07483 | -0.20699 | 0.001082 | 0.002949 | -0.1725 | 0.1083 | 4.1032
pf_sweep | 16 | 2 | 16 | 3 | 120 | 1 | 0.599 | 60 | 119.991 | 0.99992 | 59.9099 | 0.071881 | -0.096066 | 0.11998 | 119.784 | 0.999 | 0.074575 | -0.20699 | -0.000918 | 0.002694 | -0.1725 | -0.0918 | 3.7484
pf_sweep | 16 | 2 | 17 | 3 | 120 | 1 | 0.599 | 60 | 119.991 | 0.99992 | 59.9099 | 0.071881 | -0.096066 | 0.11998 | 119.784 | 1 | 0.074753 | -0.20699 | 8.2e-05 | 0.002872 | -0.1725 | 0.0082 | 3.9961
pf_sweep | 16 | 2 | 18 | 3 | 120 | 1 | 0.599 | 60 | 119.991 | 0.99992 | 59.9099 | 0.071881 | -0.096066 | 0.11998 | 119.784 | 1 | 0.074688 | -0.20699 | 8.2e-05 | 0.002807 | -0.1725 | 0.0082 | 3.9057
pf_sweep | 16 | 3 | 1 | 1 | 120 | 1 | 0.599 | 60 | 120.0076 | 0.99996 | 59.9136 | 0.071898 | -0.09608 | 0.12 | 119.748 | 2.968 | 0.22379 | -0.25958 | 1.968 | 0.1519 | -0.2163 | -1.0629 | 3.7547
pf_sweep | 16 | 3 | 2 | 1 | 120 | 1 | 0.599 | 60 | 120.0076 | 0.99996 | 59.9136 | 0.071898 | -0.09608 | 0.12 | 119.748 | 1.001 | 0.075025 | -0.25958 | 0.001038 | 0.003127 | -0.2163 | 0.1038 | 4.3491
pf_sweep | 16 | 3 | 3 | 1 | 120 | 1 | 0.599 | 60 | 120.0076 | 0.99996 | 59.9136 | 0.071898 | -0.09608 | 0.12 | 119.748 | 1.001 | 0.075012 | -0.25958 | 0.001038 | 0.003114 | -0.2163 | 0.1038 | 4.331
pf_sweep | 16 | 3 | 4 | 1 | 120 | 1 | 0.599 | 60 | 120.0076 | 0.99996 | 59.9136 | 0.071898 | -0.09608 | 0.12 | 119.748 | 1 | 0.07495 | -0.25958 | 3.8e-05 | 0.003052 | -0.2163 | 0.0038 | 4.2448
pf_sweep | 16 | 3 | 5 | 1 | 120 | 1 | 0.599 | 60 | 120.0076 | 0.99996 | 59.9136 | 0.071898 | -0.09608 | 0.12 | 119.748 | 1.001 | 0.074986 | -0.25958 | 0.001038 | 0.003088 | -0.2163 | 0.1038 | 4.2948
pf_sweep | 16 | 3 | 6 | 1 | 120 | 1 | 0.599 | 60 | 120.0076 | 0.99996 | 59.9136 | 0.071898 | -0.09608 | 0.12 | 119.748 | 0.998 | 0.074344 | -0.25958 | -0.001962 | 0.002446 | -0.2163 | -0.1962 | 3.4019
pf_sweep | 16 | 3 | 7 | 2 | 120 | 1 | 0.599 | 60 | 119.9843 | 1.0001 | 59.9652 | 0.071956 | -0.096028 | 0.12 | 119.776 | 2.978 | 0.2254 | -0.20831 | 1.9779 | 0.15344 | -0.1736 | -0.74265 | 4.414
pf_sweep | 16 | 3 | 8 | 2 | 120 | 1 | 0.599 | 60 | 119.9843 | 1.0001 | 59.9652 | 0.071956 | -0.096028 | 0.12 | 119.776 | 1.001 | 0.075013 | -0.20831 | 0.000906 | 0.003057 | -0.1736 | 0.0906 | 4.2491
pf_sweep | 16 | 3 | 9 | 2 | 120 | 1 | 0.599 | 60 | 119.9843 | 1.0001 | 59.9652 | 0.071956 | -0.096028 | 0.12 | 119.776 | 1.001 | 0.075026 | -0.20831 | 0.000906 | 0.00307 | -0.1736 | 0.0906 | 4.2671
pf_sweep | 16 | 3 | 10 | 2 | 120 | 1 | 0.599 | 60 | 119.9843 | 1.0001 | 59.9652 | 0.071956 | -0.096028 | 0.12 | 119.776 | 1 | 0.075065 | -0.20831 | -9.4e-05 | 0.003109 | -0.1736 | -0.0094 | 4.3213
pf_sweep | 16 | 3 | 11 | 2 | 120 | 1 | 0.599 | 60 | 119.9843 | 1.0001 | 59.9652 | 0.071956 | -0.096028 | 0.12 | 119.776 | 1 | 0.074851 | -0.20831 | -9.4e-05 | 0.002895 | -0.1736 | -0.0094 | 4.0239
pf_sweep | 16 | 3 | 12 | 2 | 120 | 1 | 0.599 | 60 | 119.9843 | 1.0001 | 59.9652 | 0.071956 | -0.096028 | 0.12 | 119.776 | 1 | 0.075135 | -0.20831 | -9.4e-05 | 0.003179 | -0.1736 | -0.0094 | 4.4186
pf_sweep | 16 | 3 | 13 | 3 | 120 | 1 | 0.599 | 60 | 119.9897 | 0.99991 | 59.9116 | 0.071882 | -0.096063 | 0.11998 | 119.766 | 1 | 0.074818 | -0.22372 | 8.6e-05 | 0.002936 | -0.1864 | 0.0086 | 4.0851
pf_sweep | 16 | 3 | 14 | 3 | 120 | 1 | 0.599 | 60 | 119.9897 | 0.99991 | 59.9116 | 0.071882 | -0.096063 | 0.11998 | 119.766 | 0.999 | 0.074609 | -0.22372 | -0.000914 | 0.002727 | -0.1864 | -0.0914 | 3.7944
pf_sweep | 16 | 3 | 15 | 3 | 120 | 1 | 0.599 | 60 | 119.9897 | 0.99991 | 59.9116 | 0.071882 | -0.096063 | 0.11998 | 119.766 | 1.001 | 0.074836 | -0.22372 | 0.001086 | 0.002954 | -0.1864 | 0.1086 | 4.1102
pf_sweep | 16 | 3 | 16 | 3 | 120 | 1 | 0.599 | 60 | 119.9897 | 0.99991 | 59.9116 | 0.071882 | -0.096063 | 0.11998 | 119.766 | 0.999 | 0.074551 | -0.22372 | -0.000914 | 0.002669 | -0.1864 | -0.0914 | 3.7137
pf_sweep | 16 | 3 | 17 | 3 | 120 | 1 | 0.599 | 60 | 119.9897 | 0.99991 | 59.9116 | 0.071882 | -0.096063 | 0.11998 | 119.766 | 1 | 0.074723 | -0.22372 | 8.6e-05 | 0.002841 | -0.1864 | 0.0086 | 3.953
pf_sweep | 16 | 3 | 18 | 3 | 120 | 1 | 0.599 | 60 | 119.9897 | 0.99991 | 59.9116 | 0.071882 | -0.096063 | 0.11998 | 119.766 | 1 | 0.074704 | -0.22372 | 8.6e-05 | 0.002822 | -0.1864 | 0.0086 | 3.9265
pf_sweep | 17 | 1 | 1 | 1 | 120 | 1 | 0.699 | 60 | 120.0138 | 1 | 69.9151 | 0.083907 | -0.085806 | 0.12001 | 119.731 | 2.968 | 0.25848 | -0.28279 | 1.968 | 0.17457 | -0.2356 | -1.0662 | 2.6835
pf_sweep | 17 | 1 | 2 | 1 | 120 | 1 | 0.699 | 60 | 120.0138 | 1 | 69.9151 | 0.083907 | -0.085806 | 0.12001 | 119.731 | 1.001 | 0.086743 | -0.28279 | 0.001005 | 0.002836 | -0.2356 | 0.1005 | 3.3795
pf_sweep | 17 | 1 | 3 | 1 | 120 | 1 | 0.699 | 60 | 120.0138 | 1 | 69.9151 | 0.083907 | -0.085806 | 0.12001 | 119.731 | 1.001 | 0.086745 | -0.28279 | 0.001005 | 0.002838 | -0.2356 | 0.1005 | 3.3819
pf_sweep | 17 | 1 | 4 | 1 | 120 | 1 | 0.699 | 60 | 120.0138 | 1 | 69.9151 | 0.083907 | -0.085806 | 0.12001 | 119.731 | 1 | 0.086706 | -0.28279 | 5e-06 | 0.002799 | -0.2356 | 0.0005 | 3.3354
pf_sweep | 17 | 1 | 5 | 1 | 120 | 1 | 0.699 | 60 | 120.0138 | 1 | 69.9151 | 0.083907 | -0.085806 | 0.12001 | 119.731 | 1.001 | 0.086678 | -0.28279 | 0.001005 | 0.002771 | -0.2356 | 0.1005 | 3.3021
pf_sweep | 17 | 1 | 6 | 1 | 120 | 1 | 0.699 | 60 | 120.0138 | 1 | 69.9151 | 0.083907 | -0.085806 | 0.12001 | 119.731 | 0.998 | 0.086101 | -0.28279 | -0.001995 | 0.002194 | -0.2356 | -0.1995 | 2.6144
pf_sweep | 17 | 1 | 7 | 2 | 120 | 1 | 0.699 | 60 | 119.985 | 1.0001 | 69.96 | 0.083953 | -0.085745 | 0.12 | 119.763 | 2.978 | 0.26001 | -0.22198 | 1.9779 | 0.17605 | -0.185 | -0.74716 | 3.2352
pf_sweep | 17 | 1 | 8 | 2 | 120 | 1 | 0.699 | 60 | 119.985 | 1.0001 | 69.96 | 0.083953 | -0.085745 | 0.12 | 119.763 | 1.001 | 0.086725 | -0.22198 | 0.000861 | 0.002772 | -0.185 | 0.0861 | 3.3016
pf_sweep | 17 | 1 | 9 | 2 | 120 | 1 | 0.699 | 60 | 119.985 | 1.0001 | 69.96 | 0.083953 | -0.085745 | 0.12 | 119.763 | 1.001 | 0.086778 | -0.22198 | 0.000861 | 0.002825 | -0.185 | 0.0861 | 3.3647
pf_sweep | 17 | 1 | 10 | 2 | 120 | 1 | 0.699 | 60 | 119.985 | 1.0001 | 69.96 | 0.083953 | -0.085745 | 0.12 | 119.763 | 1 | 0.086788 | -0.22198 | -0.000139 | 0.002835 | -0.185 | -0.0139 | 3.3766
pf_sweep | 17 | 1 | 11 | 2 | 120 | 1 | 0.699 | 60 | 119.985 | 1.0001 | 69.96 | 0.083953 | -0.085745 | 0.12 | 119.763 | 1 | 0.086554 | -0.22198 | -0.000139 | 0.002601 | -0.185 | -0.0139 | 3.0979
pf_sweep | 17 | 1 | 12 | 2 | 120 | 1 | 0.699 | 60 | 119.985 | 1.0001 | 69.96 | 0.083953 | -0.085745 | 0.12 | 119.763 | 1 | 0.086855 | -0.22198 | -0.000139 | 0.002902 | -0.185 | -0.0139 | 3.4564
pf_sweep | 17 | 1 | 13 | 3 | 120 | 1 | 0.699 | 60 | 119.9912 | 0.99992 | 69.9086 | 0.083878 | -0.085791 | 0.11998 | 119.784 | 1 | 0.086527 | -0.20724 | 8e-05 | 0.002649 | -0.1727 | 0.008 | 3.1587
pf_sweep | 17 | 1 | 14 | 3 | 120 | 1 | 0.699 | 60 | 119.9912 | 0.99992 | 69.9086 | 0.083878 | -0.085791 | 0.11998 | 119.784 | 0.999 | 0.086384 | -0.20724 | -0.00092 | 0.002506 | -0.1727 | -0.092 | 2.9883
pf_sweep | 17 | 1 | 15 | 3 | 120 | 1 | 0.699 | 60 | 119.9912 | 0.99992 | 69.9086 | 0.083878 | -0.085791 | 0.11998 | 119.784 | 1.001 | 0.086585 | -0.20724 | 0.00108 | 0.002707 | -0.1727 | 0.108 | 3.2279
pf_sweep | 17 | 1 | 16 | 3 | 120 | 1 | 0.699 | 60 | 119.9912 | 0.99992 | 69.9086 | 0.083878 | -0.085791 | 0.11998 | 119.784 | 0.999 | 0.086309 | -0.20724 | -0.00092 | 0.002431 | -0.1727 | -0.092 | 2.8988
pf_sweep | 17 | 1 | 17 | 3 | 120 | 1 | 0.699 | 60 | 119.9912 | 0.99992 | 69.9086 | 0.083878 | -0.085791 | 0.11998 | 119.784 | 1 | 0.086481 | -0.20724 | 8e-05 | 0.002603 | -0.1727 | 0.008 | 3.1039
pf_sweep | 17 | 1 | 18 | 3 | 120 | 1 | 0.699 | 60 | 119.9912 | 0.99992 | 69.9086 | 0.083878 | -0.085791 | 0.11998 | 119.784 | 1 | 0.086436 | -0.20724 | 8e-05 | 0.002558 | -0.1727 | 0.008 | 3.0502
pf_sweep | 17 | 2 | 1 | 1 | 120 | 1 | 0.699 | 60 | 120.0242 | 0.99996 | 69.9135 | 0.08391 | -0.085812 | 0.12002 | 119.752 | 2.968 | 0.25852 | -0.27225 | 1.968 | 0.17461 | -0.2268 | -1.0627 | 2.6972
pf_sweep | 17 | 2 | 2 | 1 | 120 | 1 | 0.699 | 60 | 120.0242 | 0.99996 | 69.9135 | 0.08391 | -0.085812 | 0.12002 | 119.752 | 1.001 | 0.086692 | -0.27225 | 0.00104 | 0.002782 | -0.2268 | 0.104 | 3.3157
pf_sweep | 17 | 2 | 3 | 1 | 120 | 1 | 0.699 | 60 | 120.0242 | 0.99996 | 69.9135 | 0.08391 | -0.085812 | 0.12002 | 119.752 | 1.001 | 0.086754 | -0.27225 | 0.00104 | 0.002844 | -0.2268 | 0.104 | 3.3896
pf_sweep | 17 | 2 | 4 | 1 | 120 | 1 | 0.699 | 60 | 120.0242 | 0.99996 | 69.9135 | 0.08391 | -0.085812 | 0.12002 | 119.752 | 1 | 0.086721 | -0.27225 | 4e-05 | 0.002811 | -0.2268 | 0.004 | 3.3502
pf_sweep | 17 | 2 | 5 | 1 | 120 | 1 | 0.699 | 60 | 120.0242 | 0.99996 | 69.9135 | 0.08391 | -0.085812 | 0.12002 | 119.752 | 1.001 | 0.086638 | -0.27225 | 0.00104 | 0.002728 | -0.2268 | 0.104 | 3.2513
pf_sweep | 17 | 2 | 6 | 1 | 120 | 1 | 0.699 | 60 | 120.0242 | 0.99996 | 69.9135 | 0.08391 | -0.085812 | 0.12002 | 119.752 | 0.999 | 0.086139 | -0.27225 | -0.00096 | 0.002229 | -0.2268 | -0.096 | 2.6566
pf_sweep | 17 | 2 | 7 | 2 | 120 | 1 | 0.699 | 60 | 119.9843 | 1.0001 | 69.9614 | 0.083952 | -0.08574 | 0.12 | 119.768 | 2.978 | 0.26003 | -0.21634 | 1.9779 | 0.17607 | -0.1803 | -0.74396 | 3.2442
pf_sweep | 17 | 2 | 8 | 2 | 120 | 1 | 0.699 | 60 | 119.9843 | 1.0001 | 69.9614 | 0.083952 | -0.08574 | 0.12 | 119.768 | 1.001 | 0.086739 | -0.21634 | 0.000893 | 0.002787 | -0.1803 | 0.0893 | 3.3201
pf_sweep | 17 | 2 | 9 | 2 | 120 | 1 | 0.699 | 60 | 119.9843 | 1.0001 | 69.9614 | 0.083952 | -0.08574 | 0.12 | 119.768 | 1.001 | 0.086805 | -0.21634 | 0.000893 | 0.002853 | -0.1803 | 0.0893 | 3.3987
pf_sweep | 17 | 2 | 10 | 2 | 120 | 1 | 0.699 | 60 | 119.9843 | 1.0001 | 69.9614 | 0.083952 | -0.08574 | 0.12 | 119.768 | 1 | 0.086796 | -0.21634 | -0.000107 | 0.002844 | -0.1803 | -0.0107 | 3.388
pf_sweep | 17 | 2 | 11 | 2 | 120 | 1 | 0.699 | 60 | 119.9843 | 1.0001 | 69.9614 | 0.083952 | -0.08574 | 0.12 | 119.768 | 1 | 0.086577 | -0.21634 | -0.000107 | 0.002625 | -0.1803 | -0.0107 | 3.1271
pf_sweep | 17 | 2 | 12 | 2 | 120 | 1 | 0.699 | 60 | 119.9843 | 1.0001 | 69.9614 | 0.083952 | -0.08574 | 0.12 | 119.768 | 1 | 0.086892 | -0.21634 | -0.000107 | 0.00294 | -0.1803 | -0.0107 | 3.5023
pf_sweep | 17 | 2 | 13 | 3 | 120 | 1 | 0.699 | 60 | 119.9898 | 0.99993 | 69.9088 | 0.083877 | -0.085791 | 0.11998 | 119.766 | 1 | 0.086546 | -0.22375 | 7.3e-05 | 0.002669 | -0.1865 | 0.0073 | 3.1816
pf_sweep | 17 | 2 | 14 | 3 | 120 | 1 | 0.699 | 60 | 119.9898 | 0.99993 | 69.9088 | 0.083877 | -0.085791 | 0.11998 | 119.766 | 0.999 | 0.086389 | -0.22375 | -0.000927 | 0.002512 | -0.1865 | -0.0927 | 2.9944
pf_sweep | 17 | 2 | 15 | 3 | 120 | 1 | 0.699 | 60 | 119.9898 | 0.99993 | 69.9088 | 0.083877 | -0.085791 | 0.11998 | 119.766 | 1.001 | 0.086567 | -0.22375 | 0.001073 | 0.00269 | -0.1865 | 0.1073 | 3.2066
pf_sweep | 17 | 2 | 16 | 3 | 120 | 1 | 0.699 | 60 | 119.9898 | 0.99993 | 69.9088 | 0.083877 | -0.085791 | 0.11998 | 119.766 | 0.999 | 0.086309 | -0.22375 | -0.000927 | 0.002432 | -0.1865 | -0.0927 | 2.8991
pf_sweep | 17 | 2 | 17 | 3 | 120 | 1 | 0.699 | 60 | 119.9898 | 0.99993 | 69.9088 | 0.083877 | -0.085791 | 0.11998 | 119.766 | 1 | 0.086486 | -0.22375 | 7.3e-05 | 0.002609 | -0.1865 | 0.0073 | 3.1101
pf_sweep | 17 | 2 | 18 | 3 | 120 | 1 | 0.699 | 60 | 119.9898 | 0.99993 | 69.9088 | 0.083877 | -0.085791 | 0.11998 | 119.766 | 1 | 0.086399 | -0.22375 | 7.3e-05 | 0.002522 | -0.1865 | 0.0073 | 3.0064
pf_sweep | 17 | 3 | 1 | 1 | 120 | 1 | 0.699 | 60 | 120.0079 | 0.99998 | 69.9128 | 0.083899 | -0.085803 | 0.12001 | 119.737 | 2.968 | 0.25839 | -0.27095 | 1.968 | 0.17449 | -0.2258 | -1.0645 | 2.6586
pf_sweep | 17 | 3 | 2 | 1 | 120 | 1 | 0.699 | 60 | 120.0079 | 0.99998 | 69.9128 | 0.083899 | -0.085803 | 0.12001 | 119.737 | 1.001 | 0.08681 | -0.27095 | 0.001022 | 0.002911 | -0.2258 | 0.1022 | 3.4695
pf_sweep | 17 | 3 | 3 | 1 | 120 | 1 | 0.699 | 60 | 120.0079 | 0.99998 | 69.9128 | 0.083899 | -0.085803 | 0.12001 | 119.737 | 1.001 | 0.086786 | -0.27095 | 0.001022 | 0.002887 | -0.2258 | 0.1022 | 3.4409
pf_sweep | 17 | 3 | 4 | 1 | 120 | 1 | 0.699 | 60 | 120.0079 | 0.99998 | 69.9128 | 0.083899 | -0.085803 | 0.12001 | 119.737 | 1 | 0.086682 | -0.27095 | 2.2e-05 | 0.002783 | -0.2258 | 0.0022 | 3.3169
pf_sweep | 17 | 3 | 5 | 1 | 120 | 1 | 0.699 | 60 | 120.0079 | 0.99998 | 69.9128 | 0.083899 | -0.085803 | 0.12001 | 119.737 | 1.001 | 0.086774 | -0.27095 | 0.001022 | 0.002875 | -0.2258 | 0.1022 | 3.4265
pf_sweep | 17 | 3 | 6 | 1 | 120 | 1 | 0.699 | 60 | 120.0079 | 0.99998 | 69.9128 | 0.083899 | -0.085803 | 0.12001 | 119.737 | 0.998 | 0.086122 | -0.27095 | -0.001978 | 0.002223 | -0.2258 | -0.1978 | 2.6494
pf_sweep | 17 | 3 | 7 | 2 | 120 | 1 | 0.699 | 60 | 119.9849 | 1.0001 | 69.9576 | 0.083947 | -0.085745 | 0.12 | 119.761 | 2.978 | 0.26001 | -0.22392 | 1.9779 | 0.17606 | -0.1866 | -0.74371 | 3.2436
pf_sweep | 17 | 3 | 8 | 2 | 120 | 1 | 0.699 | 60 | 119.9849 | 1.0001 | 69.9576 | 0.083947 | -0.085745 | 0.12 | 119.761 | 1 | 0.086732 | -0.22392 | -0.000105 | 0.002785 | -0.1866 | -0.0105 | 3.3171
pf_sweep | 17 | 3 | 9 | 2 | 120 | 1 | 0.699 | 60 | 119.9849 | 1.0001 | 69.9576 | 0.083947 | -0.085745 | 0.12 | 119.761 | 1 | 0.086754 | -0.22392 | -0.000105 | 0.002807 | -0.1866 | -0.0105 | 3.3433
pf_sweep | 17 | 3 | 10 | 2 | 120 | 1 | 0.699 | 60 | 119.9849 | 1.0001 | 69.9576 | 0.083947 | -0.085745 | 0.12 | 119.761 | 1 | 0.086781 | -0.22392 | -0.000105 | 0.002834 | -0.1866 | -0.0105 | 3.3754
pf_sweep | 17 | 3 | 11 | 2 | 120 | 1 | 0.699 | 60 | 119.9849 | 1.0001 | 69.9576 | 0.083947 | -0.085745 | 0.12 | 119.761 | 1 | 0.08657 | -0.22392 | -0.000105 | 0.002623 | -0.1866 | -0.0105 | 3.1241
pf_sweep | 17 | 3 | 12 | 2 | 120 | 1 | 0.699 | 60 | 119.9849 | 1.0001 | 69.9576 | 0.083947 | -0.085745 | 0.12 | 119.761 | 1 | 0.086821 | -0.22392 | -0.000105 | 0.002874 | -0.1866 | -0.0105 | 3.4231
pf_sweep | 17 | 3 | 13 | 3 | 120 | 1 | 0.699 | 60 | 119.9917 | 0.99993 | 69.9114 | 0.083882 | -0.085789 | 0.11998 | 119.784 | 1 | 0.086588 | -0.20767 | 7e-05 | 0.002706 | -0.1731 | 0.007 | 3.226
pf_sweep | 17 | 3 | 14 | 3 | 120 | 1 | 0.699 | 60 | 119.9917 | 0.99993 | 69.9114 | 0.083882 | -0.085789 | 0.11998 | 119.784 | 0.999 | 0.086366 | -0.20767 | -0.00093 | 0.002484 | -0.1731 | -0.093 | 2.9614
pf_sweep | 17 | 3 | 15 | 3 | 120 | 1 | 0.699 | 60 | 119.9917 | 0.99993 | 69.9114 | 0.083882 | -0.085789 | 0.11998 | 119.784 | 1.001 | 0.086577 | -0.20767 | 0.00107 | 0.002695 | -0.1731 | 0.107 | 3.2129
pf_sweep | 17 | 3 | 16 | 3 | 120 | 1 | 0.699 | 60 | 119.9917 | 0.99993 | 69.9114 | 0.083882 | -0.085789 | 0.11998 | 119.784 | 0.999 | 0.086341 | -0.20767 | -0.00093 | 0.002459 | -0.1731 | -0.093 | 2.9315
pf_sweep | 17 | 3 | 17 | 3 | 120 | 1 | 0.699 | 60 | 119.9917 | 0.99993 | 69.9114 | 0.083882 | -0.085789 | 0.11998 | 119.784 | 1 | 0.086453 | -0.20767 | 7e-05 | 0.002571 | -0.1731 | 0.007 | 3.0651
pf_sweep | 17 | 3 | 18 | 3 | 120 | 1 | 0.699 | 60 | 119.9917 | 0.99993 | 69.9114 | 0.083882 | -0.085789 | 0.11998 | 119.784 | 1 | 0.086429 | -0.20767 | 7e-05 | 0.002547 | -0.1731 | 0.007 | 3.0365
pf_sweep | 18 | 1 | 1 | 1 | 120 | 1 | 0.799 | 60 | 120.0105 | 0.99998 | 79.9127 | 0.095902 | -0.072144 | 0.12001 | 119.709 | 2.968 | 0.29252 | -0.30153 | 1.968 | 0.19662 | -0.2513 | -1.0649 | 1.6747
pf_sweep | 18 | 1 | 2 | 1 | 120 | 1 | 0.799 | 60 | 120.0105 | 0.99998 | 79.9127 | 0.095902 | -0.072144 | 0.12001 | 119.709 | 1.001 | 0.098307 | -0.30153 | 0.001018 | 0.002405 | -0.2513 | 0.1018 | 2.5078
pf_sweep | 18 | 1 | 3 | 1 | 120 | 1 | 0.799 | 60 | 120.0105 | 0.99998 | 79.9127 | 0.095902 | -0.072144 | 0.12001 | 119.709 | 1.001 | 0.098371 | -0.30153 | 0.001018 | 0.002469 | -0.2513 | 0.1018 | 2.5745
pf_sweep | 18 | 1 | 4 | 1 | 120 | 1 | 0.799 | 60 | 120.0105 | 0.99998 | 79.9127 | 0.095902 | -0.072144 | 0.12001 | 119.709 | 1 | 0.098264 | -0.30153 | 1.8e-05 | 0.002362 | -0.2513 | 0.0018 | 2.463
pf_sweep | 18 | 1 | 5 | 1 | 120 | 1 | 0.799 | 60 | 120.0105 | 0.99998 | 79.9127 | 0.095902 | -0.072144 | 0.12001 | 119.709 | 1.001 | 0.098277 | -0.30153 | 0.001018 | 0.002375 | -0.2513 | 0.1018 | 2.4765
pf_sweep | 18 | 1 | 6 | 1 | 120 | 1 | 0.799 | 60 | 120.0105 | 0.99998 | 79.9127 | 0.095902 | -0.072144 | 0.12001 | 119.709 | 0.998 | 0.09777 | -0.30153 | -0.001982 | 0.001868 | -0.2513 | -0.1982 | 1.9478
pf_sweep | 18 | 1 | 7 | 2 | 120 | 1 | 0.799 | 60 | 119.986 | 1.0001 | 79.9496 | 0.09594 | -0.072081 | 0.12 | 119.777 | 2.978 | 0.29439 | -0.20899 | 1.9779 | 0.19845 | -0.1742 | -0.74519 | 2.2822
pf_sweep | 18 | 1 | 8 | 2 | 120 | 1 | 0.799 | 60 | 119.986 | 1.0001 | 79.9496 | 0.09594 | -0.072081 | 0.12 | 119.777 | 1 | 0.098377 | -0.20899 | -0.000119 | 0.002437 | -0.1742 | -0.0119 | 2.5404
pf_sweep | 18 | 1 | 9 | 2 | 120 | 1 | 0.799 | 60 | 119.986 | 1.0001 | 79.9496 | 0.09594 | -0.072081 | 0.12 | 119.777 | 1.001 | 0.098383 | -0.20899 | 0.000881 | 0.002443 | -0.1742 | 0.088 | 2.5466
pf_sweep | 18 | 1 | 10 | 2 | 120 | 1 | 0.799 | 60 | 119.986 | 1.0001 | 79.9496 | 0.09594 | -0.072081 | 0.12 | 119.777 | 1 | 0.09843 | -0.20899 | -0.000119 | 0.00249 | -0.1742 | -0.0119 | 2.5956
pf_sweep | 18 | 1 | 11 | 2 | 120 | 1 | 0.799 | 60 | 119.986 | 1.0001 | 79.9496 | 0.09594 | -0.072081 | 0.12 | 119.777 | 1 | 0.098235 | -0.20899 | -0.000119 | 0.002295 | -0.1742 | -0.0119 | 2.3923
pf_sweep | 18 | 1 | 12 | 2 | 120 | 1 | 0.799 | 60 | 119.986 | 1.0001 | 79.9496 | 0.09594 | -0.072081 | 0.12 | 119.777 | 1 | 0.098438 | -0.20899 | -0.000119 | 0.002498 | -0.1742 | -0.0119 | 2.6039
pf_sweep | 18 | 1 | 13 | 3 | 120 | 1 | 0.799 | 60 | 119.9911 | 0.99992 | 79.9062 | 0.095873 | -0.072139 | 0.11998 | 119.769 | 1 | 0.098179 | -0.22212 | 7.6e-05 | 0.002306 | -0.1851 | 0.0076 | 2.4051
pf_sweep | 18 | 1 | 14 | 3 | 120 | 1 | 0.799 | 60 | 119.9911 | 0.99992 | 79.9062 | 0.095873 | -0.072139 | 0.11998 | 119.769 | 0.999 | 0.097993 | -0.22212 | -0.000924 | 0.00212 | -0.1851 | -0.0924 | 2.2111
pf_sweep | 18 | 1 | 15 | 3 | 120 | 1 | 0.799 | 60 | 119.9911 | 0.99992 | 79.9062 | 0.095873 | -0.072139 | 0.11998 | 119.769 | 1.001 | 0.098231 | -0.22212 | 0.001076 | 0.002358 | -0.1851 | 0.1076 | 2.4594
pf_sweep | 18 | 1 | 16 | 3 | 120 | 1 | 0.799 | 60 | 119.9911 | 0.99992 | 79.9062 | 0.095873 | -0.072139 | 0.11998 | 119.769 | 0.999 | 0.097991 | -0.22212 | -0.000924 | 0.002118 | -0.1851 | -0.0924 | 2.2091
pf_sweep | 18 | 1 | 17 | 3 | 120 | 1 | 0.799 | 60 | 119.9911 | 0.99992 | 79.9062 | 0.095873 | -0.072139 | 0.11998 | 119.769 | 1 | 0.098102 | -0.22212 | 7.6e-05 | 0.002229 | -0.1851 | 0.0076 | 2.3248
pf_sweep | 18 | 1 | 18 | 3 | 120 | 1 | 0.799 | 60 | 119.9911 | 0.99992 | 79.9062 | 0.095873 | -0.072139 | 0.11998 | 119.769 | 1 | 0.098105 | -0.22212 | 7.6e-05 | 0.002232 | -0.1851 | 0.0076 | 2.328
pf_sweep | 18 | 2 | 1 | 1 | 120 | 1 | 0.799 | 60 | 120.008 | 0.99998 | 79.9089 | 0.095896 | -0.072149 | 0.12001 | 119.733 | 2.968 | 0.29273 | -0.27498 | 1.968 | 0.19683 | -0.2291 | -1.0651 | 1.7517
pf_sweep | 18 | 2 | 2 | 1 | 120 | 1 | 0.799 | 60 | 120.008 | 0.99998 | 79.9089 | 0.095896 | -0.072149 | 0.12001 | 119.733 | 1.001 | 0.098241 | -0.27498 | 0.001015 | 0.002345 | -0.2291 | 0.1016 | 2.4458
pf_sweep | 18 | 2 | 3 | 1 | 120 | 1 | 0.799 | 60 | 120.008 | 0.99998 | 79.9089 | 0.095896 | -0.072149 | 0.12001 | 119.733 | 1.001 | 0.098359 | -0.27498 | 0.001015 | 0.002463 | -0.2291 | 0.1016 | 2.5689
pf_sweep | 18 | 2 | 4 | 1 | 120 | 1 | 0.799 | 60 | 120.008 | 0.99998 | 79.9089 | 0.095896 | -0.072149 | 0.12001 | 119.733 | 1 | 0.098349 | -0.27498 | 1.5e-05 | 0.002453 | -0.2291 | 0.0015 | 2.5585
pf_sweep | 18 | 2 | 5 | 1 | 120 | 1 | 0.799 | 60 | 120.008 | 0.99998 | 79.9089 | 0.095896 | -0.072149 | 0.12001 | 119.733 | 1.001 | 0.09821 | -0.27498 | 0.001015 | 0.002314 | -0.2291 | 0.1016 | 2.4135
pf_sweep | 18 | 2 | 6 | 1 | 120 | 1 | 0.799 | 60 | 120.008 | 0.99998 | 79.9089 | 0.095896 | -0.072149 | 0.12001 | 119.733 | 0.998 | 0.097767 | -0.27498 | -0.001985 | 0.001871 | -0.2291 | -0.1985 | 1.9516
pf_sweep | 18 | 2 | 7 | 2 | 120 | 1 | 0.799 | 60 | 119.9849 | 1.0001 | 79.9474 | 0.095937 | -0.072084 | 0.12 | 119.765 | 2.978 | 0.29425 | -0.21989 | 1.9779 | 0.19831 | -0.1833 | -0.74611 | 2.2375
pf_sweep | 18 | 2 | 8 | 2 | 120 | 1 | 0.799 | 60 | 119.9849 | 1.0001 | 79.9474 | 0.095937 | -0.072084 | 0.12 | 119.765 | 1 | 0.098356 | -0.21989 | -0.000129 | 0.002419 | -0.1833 | -0.0129 | 2.5213
pf_sweep | 18 | 2 | 9 | 2 | 120 | 1 | 0.799 | 60 | 119.9849 | 1.0001 | 79.9474 | 0.095937 | -0.072084 | 0.12 | 119.765 | 1.001 | 0.098412 | -0.21989 | 0.000871 | 0.002475 | -0.1833 | 0.0871 | 2.5797
pf_sweep | 18 | 2 | 10 | 2 | 120 | 1 | 0.799 | 60 | 119.9849 | 1.0001 | 79.9474 | 0.095937 | -0.072084 | 0.12 | 119.765 | 1 | 0.098377 | -0.21989 | -0.000129 | 0.00244 | -0.1833 | -0.0129 | 2.5432
pf_sweep | 18 | 2 | 11 | 2 | 120 | 1 | 0.799 | 60 | 119.9849 | 1.0001 | 79.9474 | 0.095937 | -0.072084 | 0.12 | 119.765 | 1 | 0.098195 | -0.21989 | -0.000129 | 0.002258 | -0.1833 | -0.0129 | 2.3535
pf_sweep | 18 | 2 | 12 | 2 | 120 | 1 | 0.799 | 60 | 119.9849 | 1.0001 | 79.9474 | 0.095937 | -0.072084 | 0.12 | 119.765 | 1 | 0.098467 | -0.21989 | -0.000129 | 0.00253 | -0.1833 | -0.0129 | 2.637
pf_sweep | 18 | 2 | 13 | 3 | 120 | 1 | 0.799 | 60 | 119.9912 | 0.99993 | 79.9097 | 0.095878 | -0.072134 | 0.11998 | 119.784 | 1 | 0.098178 | -0.20719 | 7.3e-05 | 0.0023 | -0.1727 | 0.0073 | 2.3994
pf_sweep | 18 | 2 | 14 | 3 | 120 | 1 | 0.799 | 60 | 119.9912 | 0.99993 | 79.9097 | 0.095878 | -0.072134 | 0.11998 | 119.784 | 0.999 | 0.098051 | -0.20719 | -0.000927 | 0.002173 | -0.1727 | -0.0927 | 2.2669
pf_sweep | 18 | 2 | 15 | 3 | 120 | 1 | 0.799 | 60 | 119.9912 | 0.99993 | 79.9097 | 0.095878 | -0.072134 | 0.11998 | 119.784 | 1.001 | 0.098235 | -0.20719 | 0.001073 | 0.002357 | -0.1727 | 0.1074 | 2.4588
pf_sweep | 18 | 2 | 16 | 3 | 120 | 1 | 0.799 | 60 | 119.9912 | 0.99993 | 79.9097 | 0.095878 | -0.072134 | 0.11998 | 119.784 | 0.999 | 0.097971 | -0.20719 | -0.000927 | 0.002093 | -0.1727 | -0.0927 | 2.1835
pf_sweep | 18 | 2 | 17 | 3 | 120 | 1 | 0.799 | 60 | 119.9912 | 0.99993 | 79.9097 | 0.095878 | -0.072134 | 0.11998 | 119.784 | 1 | 0.09815 | -0.20719 | 7.3e-05 | 0.002272 | -0.1727 | 0.0073 | 2.3702
pf_sweep | 18 | 2 | 18 | 3 | 120 | 1 | 0.799 | 60 | 119.9912 | 0.99993 | 79.9097 | 0.095878 | -0.072134 | 0.11998 | 119.784 | 1 | 0.09807 | -0.20719 | 7.3e-05 | 0.002192 | -0.1727 | 0.0073 | 2.2867
pf_sweep | 18 | 3 | 1 | 1 | 120 | 1 | 0.799 | 60 | 120.0055 | 0.99997 | 79.9121 | 0.095896 | -0.072141 | 0.12 | 119.755 | 2.968 | 0.29282 | -0.25053 | 1.968 | 0.19693 | -0.2088 | -1.0636 | 1.7846
pf_sweep | 18 | 3 | 2 | 1 | 120 | 1 | 0.799 | 60 | 120.0055 | 0.99997 | 79.9121 | 0.095896 | -0.072141 | 0.12 | 119.755 | 1.001 | 0.098332 | -0.25053 | 0.001031 | 0.002436 | -0.2088 | 0.1032 | 2.5403
pf_sweep | 18 | 3 | 3 | 1 | 120 | 1 | 0.799 | 60 | 120.0055 | 0.99997 | 79.9121 | 0.095896 | -0.072141 | 0.12 | 119.755 | 1.001 | 0.098406 | -0.25053 | 0.001031 | 0.00251 | -0.2088 | 0.1032 | 2.6174
pf_sweep | 18 | 3 | 4 | 1 | 120 | 1 | 0.799 | 60 | 120.0055 | 0.99997 | 79.9121 | 0.095896 | -0.072141 | 0.12 | 119.755 | 1 | 0.098383 | -0.25053 | 3.1e-05 | 0.002487 | -0.2088 | 0.0031 | 2.5935
pf_sweep | 18 | 3 | 5 | 1 | 120 | 1 | 0.799 | 60 | 120.0055 | 0.99997 | 79.9121 | 0.095896 | -0.072141 | 0.12 | 119.755 | 1.001 | 0.09834 | -0.25053 | 0.001031 | 0.002444 | -0.2088 | 0.1032 | 2.5486
pf_sweep | 18 | 3 | 6 | 1 | 120 | 1 | 0.799 | 60 | 120.0055 | 0.99997 | 79.9121 | 0.095896 | -0.072141 | 0.12 | 119.755 | 0.999 | 0.097834 | -0.25053 | -0.000969 | 0.001938 | -0.2088 | -0.0969 | 2.021
pf_sweep | 18 | 3 | 7 | 2 | 120 | 1 | 0.799 | 60 | 119.9882 | 1.0001 | 79.9494 | 0.095941 | -0.072082 | 0.12 | 119.764 | 2.978 | 0.29437 | -0.2242 | 1.9779 | 0.19843 | -0.1868 | -0.7445 | 2.2753
pf_sweep | 18 | 3 | 8 | 2 | 120 | 1 | 0.799 | 60 | 119.9882 | 1.0001 | 79.9494 | 0.095941 | -0.072082 | 0.12 | 119.764 | 1 | 0.098395 | -0.2242 | -0.000113 | 0.002454 | -0.1868 | -0.0113 | 2.5581
pf_sweep | 18 | 3 | 9 | 2 | 120 | 1 | 0.799 | 60 | 119.9882 | 1.0001 | 79.9494 | 0.095941 | -0.072082 | 0.12 | 119.764 | 1 | 0.098382 | -0.2242 | -0.000113 | 0.002441 | -0.1868 | -0.0113 | 2.5446
pf_sweep | 18 | 3 | 10 | 2 | 120 | 1 | 0.799 | 60 | 119.9882 | 1.0001 | 79.9494 | 0.095941 | -0.072082 | 0.12 | 119.764 | 1 | 0.098413 | -0.2242 | -0.000113 | 0.002472 | -0.1868 | -0.0113 | 2.5769
pf_sweep | 18 | 3 | 11 | 2 | 120 | 1 | 0.799 | 60 | 119.9882 | 1.0001 | 79.9494 | 0.095941 | -0.072082 | 0.12 | 119.764 | 1 | 0.098234 | -0.2242 | -0.000113 | 0.002293 | -0.1868 | -0.0113 | 2.3903
pf_sweep | 18 | 3 | 12 | 2 | 120 | 1 | 0.799 | 60 | 119.9882 | 1.0001 | 79.9494 | 0.095941 | -0.072082 | 0.12 | 119.764 | 1 | 0.098427 | -0.2242 | -0.000113 | 0.002486 | -0.1868 | -0.0113 | 2.5915
pf_sweep | 18 | 3 | 13 | 3 | 120 | 1 | 0.799 | 60 | 119.9942 | 0.99991 | 79.9091 | 0.095878 | -0.072135 | 0.11998 | 119.768 | 1 | 0.09819 | -0.22616 | 9e-05 | 0.002312 | -0.1885 | 0.009 | 2.4118
pf_sweep | 18 | 3 | 14 | 3 | 120 | 1 | 0.799 | 60 | 119.9942 | 0.99991 | 79.9091 | 0.095878 | -0.072135 | 0.11998 | 119.768 | 0.999 | 0.098004 | -0.22616 | -0.00091 | 0.002126 | -0.1885 | -0.091 | 2.2178
pf_sweep | 18 | 3 | 15 | 3 | 120 | 1 | 0.799 | 60 | 119.9942 | 0.99991 | 79.9091 | 0.095878 | -0.072135 | 0.11998 | 119.768 | 1.001 | 0.098212 | -0.22616 | 0.00109 | 0.002334 | -0.1885 | 0.109 | 2.4347
pf_sweep | 18 | 3 | 16 | 3 | 120 | 1 | 0.799 | 60 | 119.9942 | 0.99991 | 79.9091 | 0.095878 | -0.072135 | 0.11998 | 119.768 | 0.999 | 0.097973 | -0.22616 | -0.00091 | 0.002095 | -0.1885 | -0.091 | 2.1854
pf_sweep | 18 | 3 | 17 | 3 | 120 | 1 | 0.799 | 60 | 119.9942 | 0.99991 | 79.9091 | 0.095878 | -0.072135 | 0.11998 | 119.768 | 1 | 0.098112 | -0.22616 | 9e-05 | 0.002234 | -0.1885 | 0.009 | 2.3304
pf_sweep | 18 | 3 | 18 | 3 | 120 | 1 | 0.799 | 60 | 119.9942 | 0.99991 | 79.9091 | 0.095878 | -0.072135 | 0.11998 | 119.768 | 1 | 0.098067 | -0.22616 | 9e-05 | 0.002189 | -0.1885 | 0.009 | 2.2835
pf_sweep | 19 | 1 | 1 | 1 | 120 | 1 | 0.899 | 60 | 119.9641 | 0.99998 | 89.9037 | 0.10785 | -0.052527 | 0.11996 | 119.743 | 2.968 | 0.32617 | -0.22106 | 1.968 | 0.21832 | -0.1843 | -1.0644 | 0.80883
pf_sweep | 19 | 1 | 2 | 1 | 120 | 1 | 0.899 | 60 | 119.9641 | 0.99998 | 89.9037 | 0.10785 | -0.052527 | 0.11996 | 119.743 | 1.001 | 0.10986 | -0.22106 | 0.001023 | 0.002013 | -0.1843 | 0.1023 | 1.8668
pf_sweep | 19 | 1 | 3 | 1 | 120 | 1 | 0.899 | 60 | 119.9641 | 0.99998 | 89.9037 | 0.10785 | -0.052527 | 0.11996 | 119.743 | 1.001 | 0.10984 | -0.22106 | 0.001023 | 0.001986 | -0.1843 | 0.1023 | 1.8417
pf_sweep | 19 | 1 | 4 | 1 | 120 | 1 | 0.899 | 60 | 119.9641 | 0.99998 | 89.9037 | 0.10785 | -0.052527 | 0.11996 | 119.743 | 1 | 0.10972 | -0.22106 | 2.3e-05 | 0.001874 | -0.1843 | 0.0023 | 1.7379
pf_sweep | 19 | 1 | 5 | 1 | 120 | 1 | 0.899 | 60 | 119.9641 | 0.99998 | 89.9037 | 0.10785 | -0.052527 | 0.11996 | 119.743 | 1.001 | 0.10985 | -0.22106 | 0.001023 | 0.001997 | -0.1843 | 0.1023 | 1.8519
pf_sweep | 19 | 1 | 6 | 1 | 120 | 1 | 0.899 | 60 | 119.9641 | 0.99998 | 89.9037 | 0.10785 | -0.052527 | 0.11996 | 119.743 | 0.998 | 0.1093 | -0.22106 | -0.001977 | 0.001447 | -0.1843 | -0.1977 | 1.342
pf_sweep | 19 | 1 | 7 | 2 | 120 | 1 | 0.899 | 60 | 119.9906 | 1.0001 | 89.9393 | 0.10793 | -0.052459 | 0.12001 | 119.775 | 2.978 | 0.32804 | -0.21559 | 1.9779 | 0.22011 | -0.1797 | -0.74534 | 1.3116
pf_sweep | 19 | 1 | 8 | 2 | 120 | 1 | 0.899 | 60 | 119.9906 | 1.0001 | 89.9393 | 0.10793 | -0.052459 | 0.12001 | 119.775 | 1 | 0.10981 | -0.21559 | -0.000121 | 0.001877 | -0.1797 | -0.0121 | 1.7393
pf_sweep | 19 | 1 | 9 | 2 | 120 | 1 | 0.899 | 60 | 119.9906 | 1.0001 | 89.9393 | 0.10793 | -0.052459 | 0.12001 | 119.775 | 1.001 | 0.10986 | -0.21559 | 0.000879 | 0.001924 | -0.1797 | 0.0879 | 1.7828
pf_sweep | 19 | 1 | 10 | 2 | 120 | 1 | 0.899 | 60 | 119.9906 | 1.0001 | 89.9393 | 0.10793 | -0.052459 | 0.12001 | 119.775 | 1 | 0.10987 | -0.21559 | -0.000121 | 0.001937 | -0.1797 | -0.0121 | 1.7949
pf_sweep | 19 | 1 | 11 | 2 | 120 | 1 | 0.899 | 60 | 119.9906 | 1.0001 | 89.9393 | 0.10793 | -0.052459 | 0.12001 | 119.775 | 1 | 0.10967 | -0.21559 | -0.000121 | 0.001735 | -0.1797 | -0.0121 | 1.6077
pf_sweep | 19 | 1 | 12 | 2 | 120 | 1 | 0.899 | 60 | 119.9906 | 1.0001 | 89.9393 | 0.10793 | -0.052459 | 0.12001 | 119.775 | 1 | 0.10984 | -0.21559 | -0.000121 | 0.001908 | -0.1797 | -0.0121 | 1.768
pf_sweep | 19 | 1 | 13 | 3 | 120 | 1 | 0.899 | 60 | 119.99 | 0.99993 | 89.9053 | 0.10787 | -0.052533 | 0.11998 | 119.784 | 1 | 0.10966 | -0.20602 | 6.9e-05 | 0.00179 | -0.1717 | 0.0069 | 1.6595
pf_sweep | 19 | 1 | 14 | 3 | 120 | 1 | 0.899 | 60 | 119.99 | 0.99993 | 89.9053 | 0.10787 | -0.052533 | 0.11998 | 119.784 | 0.999 | 0.10949 | -0.20602 | -0.000931 | 0.001622 | -0.1717 | -0.0931 | 1.5038
pf_sweep | 19 | 1 | 15 | 3 | 120 | 1 | 0.899 | 60 | 119.99 | 0.99993 | 89.9053 | 0.10787 | -0.052533 | 0.11998 | 119.784 | 1.001 | 0.10972 | -0.20602 | 0.001069 | 0.001847 | -0.1717 | 0.1069 | 1.7124
pf_sweep | 19 | 1 | 16 | 3 | 120 | 1 | 0.899 | 60 | 119.99 | 0.99993 | 89.9053 | 0.10787 | -0.052533 | 0.11998 | 119.784 | 0.999 | 0.10951 | -0.20602 | -0.000931 | 0.00164 | -0.1717 | -0.0931 | 1.5205
pf_sweep | 19 | 1 | 17 | 3 | 120 | 1 | 0.899 | 60 | 119.99 | 0.99993 | 89.9053 | 0.10787 | -0.052533 | 0.11998 | 119.784 | 1 | 0.10958 | -0.20602 | 6.9e-05 | 0.001712 | -0.1717 | 0.0069 | 1.5872
pf_sweep | 19 | 1 | 18 | 3 | 120 | 1 | 0.899 | 60 | 119.99 | 0.99993 | 89.9053 | 0.10787 | -0.052533 | 0.11998 | 119.784 | 1 | 0.10959 | -0.20602 | 6.9e-05 | 0.001724 | -0.1717 | 0.0069 | 1.5983
pf_sweep | 19 | 2 | 1 | 1 | 120 | 1 | 0.899 | 60 | 120.0184 | 0.99998 | 89.9093 | 0.10791 | -0.052538 | 0.12002 | 119.746 | 2.968 | 0.3263 | -0.27238 | 1.968 | 0.21839 | -0.2269 | -1.0652 | 0.79724
pf_sweep | 19 | 2 | 2 | 1 | 120 | 1 | 0.899 | 60 | 120.0184 | 0.99998 | 89.9093 | 0.10791 | -0.052538 | 0.12002 | 119.746 | 1.001 | 0.1098 | -0.27238 | 0.001015 | 0.00189 | -0.2269 | 0.1015 | 1.7515
pf_sweep | 19 | 2 | 3 | 1 | 120 | 1 | 0.899 | 60 | 120.0184 | 0.99998 | 89.9093 | 0.10791 | -0.052538 | 0.12002 | 119.746 | 1.001 | 0.10983 | -0.27238 | 0.001015 | 0.001919 | -0.2269 | 0.1015 | 1.7783
pf_sweep | 19 | 2 | 4 | 1 | 120 | 1 | 0.899 | 60 | 120.0184 | 0.99998 | 89.9093 | 0.10791 | -0.052538 | 0.12002 | 119.746 | 1 | 0.10977 | -0.27238 | 1.5e-05 | 0.001859 | -0.2269 | 0.0015 | 1.7227
pf_sweep | 19 | 2 | 5 | 1 | 120 | 1 | 0.899 | 60 | 120.0184 | 0.99998 | 89.9093 | 0.10791 | -0.052538 | 0.12002 | 119.746 | 1.001 | 0.10978 | -0.27238 | 0.001015 | 0.001874 | -0.2269 | 0.1015 | 1.7366
pf_sweep | 19 | 2 | 6 | 1 | 120 | 1 | 0.899 | 60 | 120.0184 | 0.99998 | 89.9093 | 0.10791 | -0.052538 | 0.12002 | 119.746 | 0.998 | 0.10923 | -0.27238 | -0.001985 | 0.001322 | -0.2269 | -0.1985 | 1.2251
pf_sweep | 19 | 2 | 7 | 2 | 120 | 1 | 0.899 | 60 | 119.988 | 1.0001 | 89.9345 | 0.10792 | -0.05247 | 0.12 | 119.762 | 2.978 | 0.32795 | -0.22599 | 1.9779 | 0.22003 | -0.1883 | -0.74585 | 1.2918
pf_sweep | 19 | 2 | 8 | 2 | 120 | 1 | 0.899 | 60 | 119.988 | 1.0001 | 89.9345 | 0.10792 | -0.05247 | 0.12 | 119.762 | 1 | 0.1098 | -0.22599 | -0.000126 | 0.001874 | -0.1883 | -0.0126 | 1.7362
pf_sweep | 19 | 2 | 9 | 2 | 120 | 1 | 0.899 | 60 | 119.988 | 1.0001 | 89.9345 | 0.10792 | -0.05247 | 0.12 | 119.762 | 1.001 | 0.10987 | -0.22599 | 0.000874 | 0.001947 | -0.1883 | 0.0874 | 1.8039
pf_sweep | 19 | 2 | 10 | 2 | 120 | 1 | 0.899 | 60 | 119.988 | 1.0001 | 89.9345 | 0.10792 | -0.05247 | 0.12 | 119.762 | 1 | 0.10983 | -0.22599 | -0.000126 | 0.001911 | -0.1883 | -0.0126 | 1.7705
pf_sweep | 19 | 2 | 11 | 2 | 120 | 1 | 0.899 | 60 | 119.988 | 1.0001 | 89.9345 | 0.10792 | -0.05247 | 0.12 | 119.762 | 1 | 0.10968 | -0.22599 | -0.000126 | 0.001752 | -0.1883 | -0.0126 | 1.6232
pf_sweep | 19 | 2 | 12 | 2 | 120 | 1 | 0.899 | 60 | 119.988 | 1.0001 | 89.9345 | 0.10792 | -0.05247 | 0.12 | 119.762 | 1 | 0.10989 | -0.22599 | -0.000126 | 0.001969 | -0.1883 | -0.0126 | 1.8242
pf_sweep | 19 | 2 | 13 | 3 | 120 | 1 | 0.899 | 60 | 119.9904 | 0.99992 | 89.9069 | 0.10787 | -0.052529 | 0.11998 | 119.776 | 1 | 0.10962 | -0.21436 | 7.7e-05 | 0.001746 | -0.1786 | 0.0077 | 1.6183
pf_sweep | 19 | 2 | 14 | 3 | 120 | 1 | 0.899 | 60 | 119.9904 | 0.99992 | 89.9069 | 0.10787 | -0.052529 | 0.11998 | 119.776 | 0.999 | 0.10951 | -0.21436 | -0.000923 | 0.001637 | -0.1786 | -0.0924 | 1.5173
pf_sweep | 19 | 2 | 15 | 3 | 120 | 1 | 0.899 | 60 | 119.9904 | 0.99992 | 89.9069 | 0.10787 | -0.052529 | 0.11998 | 119.776 | 1 | 0.10968 | -0.21436 | 7.7e-05 | 0.001811 | -0.1786 | 0.0077 | 1.6786
pf_sweep | 19 | 2 | 16 | 3 | 120 | 1 | 0.899 | 60 | 119.9904 | 0.99992 | 89.9069 | 0.10787 | -0.052529 | 0.11998 | 119.776 | 0.999 | 0.10946 | -0.21436 | -0.000923 | 0.001587 | -0.1786 | -0.0924 | 1.4709
pf_sweep | 19 | 2 | 17 | 3 | 120 | 1 | 0.899 | 60 | 119.9904 | 0.99992 | 89.9069 | 0.10787 | -0.052529 | 0.11998 | 119.776 | 1 | 0.1096 | -0.21436 | 7.7e-05 | 0.001727 | -0.1786 | 0.0077 | 1.6007
pf_sweep | 19 | 2 | 18 | 3 | 120 | 1 | 0.899 | 60 | 119.9904 | 0.99992 | 89.9069 | 0.10787 | -0.052529 | 0.11998 | 119.776 | 1 | 0.10956 | -0.21436 | 7.7e-05 | 0.001688 | -0.1786 | 0.0077 | 1.5645
pf_sweep | 19 | 3 | 1 | 1 | 120 | 1 | 0.899 | 60 | 120.0255 | 0.99997 | 89.91 | 0.10791 | -0.052538 | 0.12002 | 119.749 | 2.968 | 0.32622 | -0.27647 | 1.968 | 0.21831 | -0.2303 | -1.0638 | 0.76691
pf_sweep | 19 | 3 | 2 | 1 | 120 | 1 | 0.899 | 60 | 120.0255 | 0.99997 | 89.91 | 0.10791 | -0.052538 | 0.12002 | 119.749 | 1.001 | 0.1098 | -0.27647 | 0.001029 | 0.001886 | -0.2303 | 0.1029 | 1.748
pf_sweep | 19 | 3 | 3 | 1 | 120 | 1 | 0.899 | 60 | 120.0255 | 0.99997 | 89.91 | 0.10791 | -0.052538 | 0.12002 | 119.749 | 1.001 | 0.1098 | -0.27647 | 0.001029 | 0.001884 | -0.2303 | 0.1029 | 1.7461
pf_sweep | 19 | 3 | 4 | 1 | 120 | 1 | 0.899 | 60 | 120.0255 | 0.99997 | 89.91 | 0.10791 | -0.052538 | 0.12002 | 119.749 | 1 | 0.10971 | -0.27647 | 2.9e-05 | 0.001803 | -0.2303 | 0.0029 | 1.671
pf_sweep | 19 | 3 | 5 | 1 | 120 | 1 | 0.899 | 60 | 120.0255 | 0.99997 | 89.91 | 0.10791 | -0.052538 | 0.12002 | 119.749 | 1.001 | 0.10977 | -0.27647 | 0.001029 | 0.001861 | -0.2303 | 0.1029 | 1.7248
pf_sweep | 19 | 3 | 6 | 1 | 120 | 1 | 0.899 | 60 | 120.0255 | 0.99997 | 89.91 | 0.10791 | -0.052538 | 0.12002 | 119.749 | 0.998 | 0.10926 | -0.27647 | -0.001971 | 0.001345 | -0.2303 | -0.1971 | 1.2466
pf_sweep | 19 | 3 | 7 | 2 | 120 | 1 | 0.899 | 60 | 119.9873 | 1.0001 | 89.9342 | 0.10792 | -0.052471 | 0.12 | 119.779 | 2.978 | 0.32799 | -0.20832 | 1.9779 | 0.22007 | -0.1736 | -0.74662 | 1.3027
pf_sweep | 19 | 3 | 8 | 2 | 120 | 1 | 0.899 | 60 | 119.9873 | 1.0001 | 89.9342 | 0.10792 | -0.052471 | 0.12 | 119.779 | 1 | 0.10983 | -0.20832 | -0.000134 | 0.001909 | -0.1736 | -0.0134 | 1.7688
pf_sweep | 19 | 3 | 9 | 2 | 120 | 1 | 0.899 | 60 | 119.9873 | 1.0001 | 89.9342 | 0.10792 | -0.052471 | 0.12 | 119.779 | 1.001 | 0.10983 | -0.20832 | 0.000866 | 0.001908 | -0.1736 | 0.0866 | 1.7678
pf_sweep | 19 | 3 | 10 | 2 | 120 | 1 | 0.899 | 60 | 119.9873 | 1.0001 | 89.9342 | 0.10792 | -0.052471 | 0.12 | 119.779 | 1 | 0.10985 | -0.20832 | -0.000134 | 0.001926 | -0.1736 | -0.0134 | 1.7845
pf_sweep | 19 | 3 | 11 | 2 | 120 | 1 | 0.899 | 60 | 119.9873 | 1.0001 | 89.9342 | 0.10792 | -0.052471 | 0.12 | 119.779 | 1 | 0.10972 | -0.20832 | -0.000134 | 0.001797 | -0.1736 | -0.0134 | 1.665
pf_sweep | 19 | 3 | 12 | 2 | 120 | 1 | 0.899 | 60 | 119.9873 | 1.0001 | 89.9342 | 0.10792 | -0.052471 | 0.12 | 119.779 | 1 | 0.10984 | -0.20832 | -0.000134 | 0.001921 | -0.1736 | -0.0134 | 1.7799
pf_sweep | 19 | 3 | 13 | 3 | 120 | 1 | 0.899 | 60 | 119.9934 | 0.99992 | 89.9055 | 0.10787 | -0.052533 | 0.11998 | 119.77 | 1 | 0.10966 | -0.22342 | 7.8e-05 | 0.001786 | -0.1862 | 0.0078 | 1.6554
pf_sweep | 19 | 3 | 14 | 3 | 120 | 1 | 0.899 | 60 | 119.9934 | 0.99992 | 89.9055 | 0.10787 | -0.052533 | 0.11998 | 119.77 | 0.999 | 0.10949 | -0.22342 | -0.000922 | 0.001621 | -0.1862 | -0.0922 | 1.5025
pf_sweep | 19 | 3 | 15 | 3 | 120 | 1 | 0.899 | 60 | 119.9934 | 0.99992 | 89.9055 | 0.10787 | -0.052533 | 0.11998 | 119.77 | 1.001 | 0.10971 | -0.22342 | 0.001078 | 0.001843 | -0.1862 | 0.1078 | 1.7083
pf_sweep | 19 | 3 | 16 | 3 | 120 | 1 | 0.899 | 60 | 119.9934 | 0.99992 | 89.9055 | 0.10787 | -0.052533 | 0.11998 | 119.77 | 0.999 | 0.10951 | -0.22342 | -0.000922 | 0.001636 | -0.1862 | -0.0922 | 1.5164
pf_sweep | 19 | 3 | 17 | 3 | 120 | 1 | 0.899 | 60 | 119.9934 | 0.99992 | 89.9055 | 0.10787 | -0.052533 | 0.11998 | 119.77 | 1 | 0.10958 | -0.22342 | 7.8e-05 | 0.00171 | -0.1862 | 0.0078 | 1.585
pf_sweep | 19 | 3 | 18 | 3 | 120 | 1 | 0.899 | 60 | 119.9934 | 0.99992 | 89.9055 | 0.10787 | -0.052533 | 0.11998 | 119.77 | 1 | 0.10955 | -0.22342 | 7.8e-05 | 0.001681 | -0.1862 | 0.0078 | 1.5581
pf_sweep | 20 | 1 | 1 | 1 | 120 | 1 | 0.999 | 60 | 120.0173 | 0.99998 | 99.9006 | 0.1199 | -0.0053444 | 0.12001 | 119.73 | 2.968 | 0.35543 | -0.28735 | 1.968 | 0.23553 | -0.2394 | -1.0643 | -1.1837
pf_sweep | 20 | 1 | 2 | 1 | 120 | 1 | 0.999 | 60 | 120.0173 | 0.99998 | 99.9006 | 0.1199 | -0.0053444 | 0.12001 | 119.73 | 1.001 | 0.11991 | -0.28735 | 0.001024 | 1.8e-05 | -0.2394 | 0.1024 | 0.0149
pf_sweep | 20 | 1 | 3 | 1 | 120 | 1 | 0.999 | 60 | 120.0173 | 0.99998 | 99.9006 | 0.1199 | -0.0053444 | 0.12001 | 119.73 | 1.001 | 0.11994 | -0.28735 | 0.001024 | 5e-05 | -0.2394 | 0.1024 | 0.0416
pf_sweep | 20 | 1 | 4 | 1 | 120 | 1 | 0.999 | 60 | 120.0173 | 0.99998 | 99.9006 | 0.1199 | -0.0053444 | 0.12001 | 119.73 | 1 | 0.11984 | -0.28735 | 2.4e-05 | -5.4e-05 | -0.2394 | 0.0024 | -0.0452
pf_sweep | 20 | 1 | 5 | 1 | 120 | 1 | 0.999 | 60 | 120.0173 | 0.99998 | 99.9006 | 0.1199 | -0.0053444 | 0.12001 | 119.73 | 1.001 | 0.11991 | -0.28735 | 0.001024 | 1.6e-05 | -0.2394 | 0.1024 | 0.0132
pf_sweep | 20 | 1 | 6 | 1 | 120 | 1 | 0.999 | 60 | 120.0173 | 0.99998 | 99.9006 | 0.1199 | -0.0053444 | 0.12001 | 119.73 | 0.998 | 0.11961 | -0.28735 | -0.001976 | -0.000288 | -0.2394 | -0.1976 | -0.2404
pf_sweep | 20 | 1 | 7 | 2 | 120 | 1 | 0.999 | 60 | 119.987 | 1.0002 | 99.9035 | 0.11989 | -0.0052707 | 0.12001 | 119.763 | 2.979 | 0.35682 | -0.22404 | 1.9788 | 0.23693 | -0.1867 | -0.71692 | -0.79351
pf_sweep | 20 | 1 | 8 | 2 | 120 | 1 | 0.999 | 60 | 119.987 | 1.0002 | 99.9035 | 0.11989 | -0.0052707 | 0.12001 | 119.763 | 1 | 0.11983 | -0.22404 | -0.00017 | -6.1e-05 | -0.1867 | -0.017 | -0.0506
pf_sweep | 20 | 1 | 9 | 2 | 120 | 1 | 0.999 | 60 | 119.987 | 1.0002 | 99.9035 | 0.11989 | -0.0052707 | 0.12001 | 119.763 | 1.001 | 0.11989 | -0.22404 | 0.00083 | -4e-06 | -0.1867 | 0.0829 | -0.0031
pf_sweep | 20 | 1 | 10 | 2 | 120 | 1 | 0.999 | 60 | 119.987 | 1.0002 | 99.9035 | 0.11989 | -0.0052707 | 0.12001 | 119.763 | 1 | 0.11986 | -0.22404 | -0.00017 | -3e-05 | -0.1867 | -0.017 | -0.0248
pf_sweep | 20 | 1 | 11 | 2 | 120 | 1 | 0.999 | 60 | 119.987 | 1.0002 | 99.9035 | 0.11989 | -0.0052707 | 0.12001 | 119.763 | 1 | 0.1198 | -0.22404 | -0.00017 | -8.9e-05 | -0.1867 | -0.017 | -0.074
pf_sweep | 20 | 1 | 12 | 2 | 120 | 1 | 0.999 | 60 | 119.987 | 1.0002 | 99.9035 | 0.11989 | -0.0052707 | 0.12001 | 119.763 | 1 | 0.11981 | -0.22404 | -0.00017 | -8.4e-05 | -0.1867 | -0.017 | -0.0698
pf_sweep | 20 | 1 | 13 | 3 | 120 | 1 | 0.999 | 60 | 119.9886 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.779 | 1 | 0.11979 | -0.20965 | 8.2e-05 | -7.4e-05 | -0.1747 | 0.0082 | -0.0614
pf_sweep | 20 | 1 | 14 | 3 | 120 | 1 | 0.999 | 60 | 119.9886 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.779 | 0.999 | 0.11972 | -0.20965 | -0.000918 | -0.000136 | -0.1747 | -0.0918 | -0.1131
pf_sweep | 20 | 1 | 15 | 3 | 120 | 1 | 0.999 | 60 | 119.9886 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.779 | 1.001 | 0.1199 | -0.20965 | 0.001082 | 4.1e-05 | -0.1747 | 0.1082 | 0.0345
pf_sweep | 20 | 1 | 16 | 3 | 120 | 1 | 0.999 | 60 | 119.9886 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.779 | 0.999 | 0.11975 | -0.20965 | -0.000918 | -0.000115 | -0.1747 | -0.0918 | -0.0956
pf_sweep | 20 | 1 | 17 | 3 | 120 | 1 | 0.999 | 60 | 119.9886 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.779 | 1 | 0.11978 | -0.20965 | 8.2e-05 | -8e-05 | -0.1747 | 0.0082 | -0.0664
pf_sweep | 20 | 1 | 18 | 3 | 120 | 1 | 0.999 | 60 | 119.9886 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.779 | 1 | 0.1198 | -0.20965 | 8.2e-05 | -6.2e-05 | -0.1747 | 0.0082 | -0.0514
pf_sweep | 20 | 2 | 1 | 1 | 120 | 1 | 0.999 | 60 | 120.007 | 0.99999 | 99.9006 | 0.11989 | -0.0053452 | 0.12001 | 119.769 | 2.968 | 0.35557 | -0.23803 | 1.968 | 0.23568 | -0.1983 | -1.0654 | -1.1368
pf_sweep | 20 | 2 | 2 | 1 | 120 | 1 | 0.999 | 60 | 120.007 | 0.99999 | 99.9006 | 0.11989 | -0.0053452 | 0.12001 | 119.769 | 1.001 | 0.11991 | -0.23803 | 0.001013 | 2.8e-05 | -0.1983 | 0.1013 | 0.0232
pf_sweep | 20 | 2 | 3 | 1 | 120 | 1 | 0.999 | 60 | 120.007 | 0.99999 | 99.9006 | 0.11989 | -0.0053452 | 0.12001 | 119.769 | 1.001 | 0.11999 | -0.23803 | 0.001013 | 0.000103 | -0.1983 | 0.1013 | 0.0858
pf_sweep | 20 | 2 | 4 | 1 | 120 | 1 | 0.999 | 60 | 120.007 | 0.99999 | 99.9006 | 0.11989 | -0.0053452 | 0.12001 | 119.769 | 1.001 | 0.11989 | -0.23803 | 0.001013 | 4e-06 | -0.1983 | 0.1013 | 0.0032
pf_sweep | 20 | 2 | 5 | 1 | 120 | 1 | 0.999 | 60 | 120.007 | 0.99999 | 99.9006 | 0.11989 | -0.0053452 | 0.12001 | 119.769 | 1.001 | 0.11991 | -0.23803 | 0.001013 | 2.3e-05 | -0.1983 | 0.1013 | 0.019
pf_sweep | 20 | 2 | 6 | 1 | 120 | 1 | 0.999 | 60 | 120.007 | 0.99999 | 99.9006 | 0.11989 | -0.0053452 | 0.12001 | 119.769 | 0.998 | 0.11962 | -0.23803 | -0.001987 | -0.000266 | -0.1983 | -0.1987 | -0.222
pf_sweep | 20 | 2 | 7 | 2 | 120 | 1 | 0.999 | 60 | 119.9888 | 1.0001 | 99.9037 | 0.11989 | -0.0052656 | 0.12001 | 119.772 | 2.978 | 0.35677 | -0.21682 | 1.9779 | 0.23688 | -0.1807 | -0.74746 | -0.80597
pf_sweep | 20 | 2 | 8 | 2 | 120 | 1 | 0.999 | 60 | 119.9888 | 1.0001 | 99.9037 | 0.11989 | -0.0052656 | 0.12001 | 119.772 | 1 | 0.11982 | -0.21682 | -0.000142 | -7.3e-05 | -0.1807 | -0.0142 | -0.0611
pf_sweep | 20 | 2 | 9 | 2 | 120 | 1 | 0.999 | 60 | 119.9888 | 1.0001 | 99.9037 | 0.11989 | -0.0052656 | 0.12001 | 119.772 | 1.001 | 0.1199 | -0.21682 | 0.000858 | 1.1e-05 | -0.1807 | 0.0858 | 0.0089
pf_sweep | 20 | 2 | 10 | 2 | 120 | 1 | 0.999 | 60 | 119.9888 | 1.0001 | 99.9037 | 0.11989 | -0.0052656 | 0.12001 | 119.772 | 1 | 0.11984 | -0.21682 | -0.000142 | -4.6e-05 | -0.1807 | -0.0142 | -0.0386
pf_sweep | 20 | 2 | 11 | 2 | 120 | 1 | 0.999 | 60 | 119.9888 | 1.0001 | 99.9037 | 0.11989 | -0.0052656 | 0.12001 | 119.772 | 1 | 0.11979 | -0.21682 | -0.000142 | -0.000101 | -0.1807 | -0.0142 | -0.0845
pf_sweep | 20 | 2 | 12 | 2 | 120 | 1 | 0.999 | 60 | 119.9888 | 1.0001 | 99.9037 | 0.11989 | -0.0052656 | 0.12001 | 119.772 | 1 | 0.11983 | -0.21682 | -0.000142 | -5.8e-05 | -0.1807 | -0.0142 | -0.0486
pf_sweep | 20 | 2 | 13 | 3 | 120 | 1 | 0.999 | 60 | 119.9921 | 0.99993 | 99.9007 | 0.11987 | -0.0053447 | 0.11998 | 119.767 | 1 | 0.11978 | -0.22514 | 6.7e-05 | -8.9e-05 | -0.1876 | 0.0067 | -0.0743
pf_sweep | 20 | 2 | 14 | 3 | 120 | 1 | 0.999 | 60 | 119.9921 | 0.99993 | 99.9007 | 0.11987 | -0.0053447 | 0.11998 | 119.767 | 0.999 | 0.11975 | -0.22514 | -0.000933 | -0.000111 | -0.1876 | -0.0934 | -0.0926
pf_sweep | 20 | 2 | 15 | 3 | 120 | 1 | 0.999 | 60 | 119.9921 | 0.99993 | 99.9007 | 0.11987 | -0.0053447 | 0.11998 | 119.767 | 1.001 | 0.11989 | -0.22514 | 0.001067 | 2.7e-05 | -0.1876 | 0.1067 | 0.0225
pf_sweep | 20 | 2 | 16 | 3 | 120 | 1 | 0.999 | 60 | 119.9921 | 0.99993 | 99.9007 | 0.11987 | -0.0053447 | 0.11998 | 119.767 | 0.999 | 0.11974 | -0.22514 | -0.000933 | -0.000128 | -0.1876 | -0.0934 | -0.1068
pf_sweep | 20 | 2 | 17 | 3 | 120 | 1 | 0.999 | 60 | 119.9921 | 0.99993 | 99.9007 | 0.11987 | -0.0053447 | 0.11998 | 119.767 | 1 | 0.11983 | -0.22514 | 6.7e-05 | -3.5e-05 | -0.1876 | 0.0067 | -0.0292
pf_sweep | 20 | 2 | 18 | 3 | 120 | 1 | 0.999 | 60 | 119.9921 | 0.99993 | 99.9007 | 0.11987 | -0.0053447 | 0.11998 | 119.767 | 1 | 0.11978 | -0.22514 | 6.7e-05 | -8.6e-05 | -0.1876 | 0.0067 | -0.0718
pf_sweep | 20 | 3 | 1 | 1 | 120 | 1 | 0.999 | 60 | 120.001 | 0.99997 | 99.9005 | 0.11988 | -0.0053421 | 0.12 | 119.712 | 2.968 | 0.35537 | -0.28903 | 1.968 | 0.23549 | -0.2409 | -1.0639 | -1.1855
pf_sweep | 20 | 3 | 2 | 1 | 120 | 1 | 0.999 | 60 | 120.001 | 0.99997 | 99.9005 | 0.11988 | -0.0053421 | 0.12 | 119.712 | 1.001 | 0.11993 | -0.28903 | 0.001028 | 5.6e-05 | -0.2409 | 0.1028 | 0.0465
pf_sweep | 20 | 3 | 3 | 1 | 120 | 1 | 0.999 | 60 | 120.001 | 0.99997 | 99.9005 | 0.11988 | -0.0053421 | 0.12 | 119.712 | 1.001 | 0.11997 | -0.28903 | 0.001028 | 9.4e-05 | -0.2409 | 0.1028 | 0.0782
pf_sweep | 20 | 3 | 4 | 1 | 120 | 1 | 0.999 | 60 | 120.001 | 0.99997 | 99.9005 | 0.11988 | -0.0053421 | 0.12 | 119.712 | 1 | 0.11982 | -0.28903 | 2.8e-05 | -5.8e-05 | -0.2409 | 0.0028 | -0.0486
pf_sweep | 20 | 3 | 5 | 1 | 120 | 1 | 0.999 | 60 | 120.001 | 0.99997 | 99.9005 | 0.11988 | -0.0053421 | 0.12 | 119.712 | 1.001 | 0.11994 | -0.28903 | 0.001028 | 6e-05 | -0.2409 | 0.1028 | 0.0499
pf_sweep | 20 | 3 | 6 | 1 | 120 | 1 | 0.999 | 60 | 120.001 | 0.99997 | 99.9005 | 0.11988 | -0.0053421 | 0.12 | 119.712 | 0.998 | 0.11963 | -0.28903 | -0.001972 | -0.000249 | -0.2409 | -0.1972 | -0.2079
pf_sweep | 20 | 3 | 7 | 2 | 120 | 1 | 0.999 | 60 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052648 | 0.12 | 119.762 | 2.978 | 0.35673 | -0.22171 | 1.9778 | 0.23684 | -0.1848 | -0.7492 | -0.81662
pf_sweep | 20 | 3 | 8 | 2 | 120 | 1 | 0.999 | 60 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052648 | 0.12 | 119.762 | 1 | 0.11983 | -0.22171 | -0.00016 | -5.3e-05 | -0.1848 | -0.016 | -0.0445
pf_sweep | 20 | 3 | 9 | 2 | 120 | 1 | 0.999 | 60 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052648 | 0.12 | 119.762 | 1 | 0.11988 | -0.22171 | -0.00016 | -6e-06 | -0.1848 | -0.016 | -0.0053
pf_sweep | 20 | 3 | 10 | 2 | 120 | 1 | 0.999 | 60 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052648 | 0.12 | 119.762 | 1 | 0.11983 | -0.22171 | -0.00016 | -5.3e-05 | -0.1848 | -0.016 | -0.0445
pf_sweep | 20 | 3 | 11 | 2 | 120 | 1 | 0.999 | 60 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052648 | 0.12 | 119.762 | 1 | 0.11981 | -0.22171 | -0.00016 | -7.8e-05 | -0.1848 | -0.016 | -0.0654
pf_sweep | 20 | 3 | 12 | 2 | 120 | 1 | 0.999 | 60 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052648 | 0.12 | 119.762 | 1 | 0.11982 | -0.22171 | -0.00016 | -6.7e-05 | -0.1848 | -0.016 | -0.0562
pf_sweep | 20 | 3 | 13 | 3 | 120 | 1 | 0.999 | 60 | 119.9927 | 0.9999 | 99.9006 | 0.11986 | -0.0053469 | 0.11998 | 119.781 | 1 | 0.11981 | -0.21167 | 9.6e-05 | -5.6e-05 | -0.1764 | 0.0096 | -0.0466
pf_sweep | 20 | 3 | 14 | 3 | 120 | 1 | 0.999 | 60 | 119.9927 | 0.9999 | 99.9006 | 0.11986 | -0.0053469 | 0.11998 | 119.781 | 0.999 | 0.11973 | -0.21167 | -0.000904 | -0.000131 | -0.1764 | -0.0904 | -0.1092
pf_sweep | 20 | 3 | 15 | 3 | 120 | 1 | 0.999 | 60 | 119.9927 | 0.9999 | 99.9006 | 0.11986 | -0.0053469 | 0.11998 | 119.781 | 1 | 0.11988 | -0.21167 | 9.6e-05 | 1.6e-05 | -0.1764 | 0.0096 | 0.0134
pf_sweep | 20 | 3 | 16 | 3 | 120 | 1 | 0.999 | 60 | 119.9927 | 0.9999 | 99.9006 | 0.11986 | -0.0053469 | 0.11998 | 119.781 | 0.999 | 0.11976 | -0.21167 | -0.000904 | -0.0001 | -0.1764 | -0.0904 | -0.0833
pf_sweep | 20 | 3 | 17 | 3 | 120 | 1 | 0.999 | 60 | 119.9927 | 0.9999 | 99.9006 | 0.11986 | -0.0053469 | 0.11998 | 119.781 | 1 | 0.1198 | -0.21167 | 9.6e-05 | -6.5e-05 | -0.1764 | 0.0096 | -0.0541
pf_sweep | 20 | 3 | 18 | 3 | 120 | 1 | 0.999 | 60 | 119.9927 | 0.9999 | 99.9006 | 0.11986 | -0.0053469 | 0.11998 | 119.781 | 1 | 0.11978 | -0.21167 | 9.6e-05 | -7.8e-05 | -0.1764 | 0.0096 | -0.065

## Freq_Sweep
sweep | point | sample | channel | phase | source_vln | source_i | source_pf | source_freq | ref_vln | ref_i | ref_pf | ref_kW | ref_kVAR | ref_kVA | dut_vln | dut_i | dut_kW | err_vln | err_i | err_kW | pct_err_vln | pct_err_i | pct_err_kW
freq_sweep | 1 | 1 | 1 | 1 | 120 | 1 | 0.999 | 50 | 120.0104 | 1 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.721 | 2.966 | 0.35521 | -0.28941 | 1.966 | 0.23532 | -0.2412 | -1.1329 | -1.2405
freq_sweep | 1 | 1 | 2 | 1 | 120 | 1 | 0.999 | 50 | 120.0104 | 1 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.721 | 1.001 | 0.11993 | -0.28941 | 0.001005 | 3.5e-05 | -0.2412 | 0.1005 | 0.0296
freq_sweep | 1 | 1 | 3 | 1 | 120 | 1 | 0.999 | 50 | 120.0104 | 1 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.721 | 1.001 | 0.11996 | -0.28941 | 0.001005 | 6.5e-05 | -0.2412 | 0.1005 | 0.0546
freq_sweep | 1 | 1 | 4 | 1 | 120 | 1 | 0.999 | 50 | 120.0104 | 1 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.721 | 1 | 0.11981 | -0.28941 | 5e-06 | -7.9e-05 | -0.2412 | 0.0005 | -0.0655
freq_sweep | 1 | 1 | 5 | 1 | 120 | 1 | 0.999 | 50 | 120.0104 | 1 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.721 | 1.001 | 0.11992 | -0.28941 | 0.001005 | 2.7e-05 | -0.2412 | 0.1005 | 0.0229
freq_sweep | 1 | 1 | 6 | 1 | 120 | 1 | 0.999 | 50 | 120.0104 | 1 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.721 | 0.998 | 0.11961 | -0.28941 | -0.001995 | -0.000286 | -0.2412 | -0.1995 | -0.2382
freq_sweep | 1 | 1 | 7 | 2 | 120 | 1 | 0.999 | 50 | 119.9827 | 1.0002 | 99.9037 | 0.11989 | -0.0052654 | 0.12 | 119.757 | 2.977 | 0.35663 | -0.22566 | 1.9768 | 0.23675 | -0.1881 | -0.78358 | -0.84232
freq_sweep | 1 | 1 | 8 | 2 | 120 | 1 | 0.999 | 50 | 119.9827 | 1.0002 | 99.9037 | 0.11989 | -0.0052654 | 0.12 | 119.757 | 1 | 0.11983 | -0.22566 | -0.00017 | -6.3e-05 | -0.1881 | -0.017 | -0.0521
freq_sweep | 1 | 1 | 9 | 2 | 120 | 1 | 0.999 | 50 | 119.9827 | 1.0002 | 99.9037 | 0.11989 | -0.0052654 | 0.12 | 119.757 | 1 | 0.11987 | -0.22566 | -0.00017 | -1.4e-05 | -0.1881 | -0.017 | -0.0113
freq_sweep | 1 | 1 | 10 | 2 | 120 | 1 | 0.999 | 50 | 119.9827 | 1.0002 | 99.9037 | 0.11989 | -0.0052654 | 0.12 | 119.757 | 1 | 0.11983 | -0.22566 | -0.00017 | -6.1e-05 | -0.1881 | -0.017 | -0.0505
freq_sweep | 1 | 1 | 11 | 2 | 120 | 1 | 0.999 | 50 | 119.9827 | 1.0002 | 99.9037 | 0.11989 | -0.0052654 | 0.12 | 119.757 | 1 | 0.11979 | -0.22566 | -0.00017 | -0.000101 | -0.1881 | -0.017 | -0.0838
freq_sweep | 1 | 1 | 12 | 2 | 120 | 1 | 0.999 | 50 | 119.9827 | 1.0002 | 99.9037 | 0.11989 | -0.0052654 | 0.12 | 119.757 | 1 | 0.1198 | -0.22566 | -0.00017 | -9e-05 | -0.1881 | -0.017 | -0.0747
freq_sweep | 1 | 1 | 13 | 3 | 120 | 1 | 0.999 | 50 | 119.9877 | 0.99993 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.781 | 1 | 0.11979 | -0.20672 | 6.7e-05 | -7e-05 | -0.1723 | 0.0067 | -0.0586
freq_sweep | 1 | 1 | 14 | 3 | 120 | 1 | 0.999 | 50 | 119.9877 | 0.99993 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.781 | 0.999 | 0.11972 | -0.20672 | -0.000933 | -0.000141 | -0.1723 | -0.0933 | -0.1178
freq_sweep | 1 | 1 | 15 | 3 | 120 | 1 | 0.999 | 50 | 119.9877 | 0.99993 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.781 | 1 | 0.11988 | -0.20672 | 6.7e-05 | 1.9e-05 | -0.1723 | 0.0067 | 0.0157
freq_sweep | 1 | 1 | 16 | 3 | 120 | 1 | 0.999 | 50 | 119.9877 | 0.99993 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.781 | 0.999 | 0.11977 | -0.20672 | -0.000933 | -9.3e-05 | -0.1723 | -0.0933 | -0.0778
freq_sweep | 1 | 1 | 17 | 3 | 120 | 1 | 0.999 | 50 | 119.9877 | 0.99993 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.781 | 1 | 0.11979 | -0.20672 | 6.7e-05 | -7.4e-05 | -0.1723 | 0.0067 | -0.0619
freq_sweep | 1 | 1 | 18 | 3 | 120 | 1 | 0.999 | 50 | 119.9877 | 0.99993 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.781 | 1 | 0.11977 | -0.20672 | 6.7e-05 | -8.7e-05 | -0.1723 | 0.0067 | -0.0728
freq_sweep | 1 | 2 | 1 | 1 | 120 | 1 | 0.999 | 50 | 119.9933 | 0.99999 | 99.9002 | 0.11987 | -0.0053476 | 0.11999 | 119.708 | 2.967 | 0.35518 | -0.2853 | 1.967 | 0.23531 | -0.2378 | -1.099 | -1.2338
freq_sweep | 1 | 2 | 2 | 1 | 120 | 1 | 0.999 | 50 | 119.9933 | 0.99999 | 99.9002 | 0.11987 | -0.0053476 | 0.11999 | 119.708 | 1.001 | 0.11993 | -0.2853 | 0.00101 | 5.6e-05 | -0.2378 | 0.101 | 0.0464
freq_sweep | 1 | 2 | 3 | 1 | 120 | 1 | 0.999 | 50 | 119.9933 | 0.99999 | 99.9002 | 0.11987 | -0.0053476 | 0.11999 | 119.708 | 1.001 | 0.11992 | -0.2853 | 0.00101 | 5e-05 | -0.2378 | 0.101 | 0.0414
freq_sweep | 1 | 2 | 4 | 1 | 120 | 1 | 0.999 | 50 | 119.9933 | 0.99999 | 99.9002 | 0.11987 | -0.0053476 | 0.11999 | 119.708 | 1 | 0.11981 | -0.2853 | 1e-05 | -6.6e-05 | -0.2378 | 0.001 | -0.0554
freq_sweep | 1 | 2 | 5 | 1 | 120 | 1 | 0.999 | 50 | 119.9933 | 0.99999 | 99.9002 | 0.11987 | -0.0053476 | 0.11999 | 119.708 | 1.001 | 0.11995 | -0.2853 | 0.00101 | 7.5e-05 | -0.2378 | 0.101 | 0.0623
freq_sweep | 1 | 2 | 6 | 1 | 120 | 1 | 0.999 | 50 | 119.9933 | 0.99999 | 99.9002 | 0.11987 | -0.0053476 | 0.11999 | 119.708 | 0.998 | 0.11953 | -0.2853 | -0.00199 | -0.000338 | -0.2378 | -0.199 | -0.2823
freq_sweep | 1 | 2 | 7 | 2 | 120 | 1 | 0.999 | 50 | 119.9834 | 1.0002 | 99.9037 | 0.11989 | -0.0052645 | 0.12 | 119.763 | 2.978 | 0.3567 | -0.22039 | 1.9778 | 0.23681 | -0.1837 | -0.75074 | -0.82427
freq_sweep | 1 | 2 | 8 | 2 | 120 | 1 | 0.999 | 50 | 119.9834 | 1.0002 | 99.9037 | 0.11989 | -0.0052645 | 0.12 | 119.763 | 1 | 0.11982 | -0.22039 | -0.000175 | -7e-05 | -0.1837 | -0.0175 | -0.0583
freq_sweep | 1 | 2 | 9 | 2 | 120 | 1 | 0.999 | 50 | 119.9834 | 1.0002 | 99.9037 | 0.11989 | -0.0052645 | 0.12 | 119.763 | 1.001 | 0.11988 | -0.22039 | 0.000825 | -1.2e-05 | -0.1837 | 0.0824 | -0.0099
freq_sweep | 1 | 2 | 10 | 2 | 120 | 1 | 0.999 | 50 | 119.9834 | 1.0002 | 99.9037 | 0.11989 | -0.0052645 | 0.12 | 119.763 | 1 | 0.11983 | -0.22039 | -0.000175 | -5.8e-05 | -0.1837 | -0.0175 | -0.0483
freq_sweep | 1 | 2 | 11 | 2 | 120 | 1 | 0.999 | 50 | 119.9834 | 1.0002 | 99.9037 | 0.11989 | -0.0052645 | 0.12 | 119.763 | 1 | 0.11979 | -0.22039 | -0.000175 | -0.000104 | -0.1837 | -0.0175 | -0.0866
freq_sweep | 1 | 2 | 12 | 2 | 120 | 1 | 0.999 | 50 | 119.9834 | 1.0002 | 99.9037 | 0.11989 | -0.0052645 | 0.12 | 119.763 | 1 | 0.1198 | -0.22039 | -0.000175 | -9.3e-05 | -0.1837 | -0.0175 | -0.0775
freq_sweep | 1 | 2 | 13 | 3 | 120 | 1 | 0.999 | 50 | 119.9898 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.784 | 1 | 0.1198 | -0.20577 | 7.8e-05 | -6e-05 | -0.1715 | 0.0078 | -0.0502
freq_sweep | 1 | 2 | 14 | 3 | 120 | 1 | 0.999 | 50 | 119.9898 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.784 | 0.999 | 0.11971 | -0.20577 | -0.000922 | -0.000147 | -0.1715 | -0.0922 | -0.1228
freq_sweep | 1 | 2 | 15 | 3 | 120 | 1 | 0.999 | 50 | 119.9898 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.784 | 1 | 0.11989 | -0.20577 | 7.8e-05 | 2.6e-05 | -0.1715 | 0.0078 | 0.0216
freq_sweep | 1 | 2 | 16 | 3 | 120 | 1 | 0.999 | 50 | 119.9898 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.784 | 0.999 | 0.11975 | -0.20577 | -0.000922 | -0.00011 | -0.1715 | -0.0922 | -0.0919
freq_sweep | 1 | 2 | 17 | 3 | 120 | 1 | 0.999 | 50 | 119.9898 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.784 | 1 | 0.11977 | -0.20577 | 7.8e-05 | -9.2e-05 | -0.1715 | 0.0078 | -0.0769
freq_sweep | 1 | 2 | 18 | 3 | 120 | 1 | 0.999 | 50 | 119.9898 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.784 | 1 | 0.11979 | -0.20577 | 7.8e-05 | -7.5e-05 | -0.1715 | 0.0078 | -0.0627
freq_sweep | 1 | 3 | 1 | 1 | 120 | 1 | 0.999 | 50 | 119.998 | 0.99998 | 99.9005 | 0.11988 | -0.0053432 | 0.12 | 119.74 | 2.967 | 0.35531 | -0.25799 | 1.967 | 0.23543 | -0.215 | -1.0984 | -1.2012
freq_sweep | 1 | 3 | 2 | 1 | 120 | 1 | 0.999 | 50 | 119.998 | 0.99998 | 99.9005 | 0.11988 | -0.0053432 | 0.12 | 119.74 | 1.001 | 0.11994 | -0.25799 | 0.001016 | 6.3e-05 | -0.215 | 0.1016 | 0.0528
freq_sweep | 1 | 3 | 3 | 1 | 120 | 1 | 0.999 | 50 | 119.998 | 0.99998 | 99.9005 | 0.11988 | -0.0053432 | 0.12 | 119.74 | 1.001 | 0.11994 | -0.25799 | 0.001016 | 6.1e-05 | -0.215 | 0.1016 | 0.0512
freq_sweep | 1 | 3 | 4 | 1 | 120 | 1 | 0.999 | 50 | 119.998 | 0.99998 | 99.9005 | 0.11988 | -0.0053432 | 0.12 | 119.74 | 1 | 0.11982 | -0.25799 | 1.6e-05 | -5.7e-05 | -0.215 | 0.0016 | -0.0473
freq_sweep | 1 | 3 | 5 | 1 | 120 | 1 | 0.999 | 50 | 119.998 | 0.99998 | 99.9005 | 0.11988 | -0.0053432 | 0.12 | 119.74 | 1.001 | 0.11993 | -0.25799 | 0.001016 | 5.7e-05 | -0.215 | 0.1016 | 0.0478
freq_sweep | 1 | 3 | 6 | 1 | 120 | 1 | 0.999 | 50 | 119.998 | 0.99998 | 99.9005 | 0.11988 | -0.0053432 | 0.12 | 119.74 | 0.998 | 0.11958 | -0.25799 | -0.001984 | -0.000295 | -0.215 | -0.1984 | -0.2458
freq_sweep | 1 | 3 | 7 | 2 | 120 | 1 | 0.999 | 50 | 119.9811 | 1.0002 | 99.9036 | 0.11988 | -0.0052665 | 0.12 | 119.768 | 2.978 | 0.35676 | -0.21308 | 1.9778 | 0.23687 | -0.1776 | -0.74914 | -0.80481
freq_sweep | 1 | 3 | 8 | 2 | 120 | 1 | 0.999 | 50 | 119.9811 | 1.0002 | 99.9036 | 0.11988 | -0.0052665 | 0.12 | 119.768 | 1 | 0.11982 | -0.21308 | -0.000159 | -6.5e-05 | -0.1776 | -0.0159 | -0.0538
freq_sweep | 1 | 3 | 9 | 2 | 120 | 1 | 0.999 | 50 | 119.9811 | 1.0002 | 99.9036 | 0.11988 | -0.0052665 | 0.12 | 119.768 | 1.001 | 0.11989 | -0.21308 | 0.000841 | 5e-06 | -0.1776 | 0.0841 | 0.0046
freq_sweep | 1 | 3 | 10 | 2 | 120 | 1 | 0.999 | 50 | 119.9811 | 1.0002 | 99.9036 | 0.11988 | -0.0052665 | 0.12 | 119.768 | 1 | 0.11986 | -0.21308 | -0.000159 | -3e-05 | -0.1776 | -0.0159 | -0.0246
freq_sweep | 1 | 3 | 11 | 2 | 120 | 1 | 0.999 | 50 | 119.9811 | 1.0002 | 99.9036 | 0.11988 | -0.0052665 | 0.12 | 119.768 | 1 | 0.11979 | -0.21308 | -0.000159 | -9.5e-05 | -0.1776 | -0.0159 | -0.0788
freq_sweep | 1 | 3 | 12 | 2 | 120 | 1 | 0.999 | 50 | 119.9811 | 1.0002 | 99.9036 | 0.11988 | -0.0052665 | 0.12 | 119.768 | 1 | 0.11982 | -0.21308 | -0.000159 | -6.5e-05 | -0.1776 | -0.0159 | -0.0538
freq_sweep | 1 | 3 | 13 | 3 | 120 | 1 | 0.999 | 50 | 119.9943 | 0.99992 | 99.9006 | 0.11987 | -0.0053469 | 0.11998 | 119.767 | 0.999 | 0.11976 | -0.2273 | -0.000922 | -0.000103 | -0.1894 | -0.0922 | -0.0857
freq_sweep | 1 | 3 | 14 | 3 | 120 | 1 | 0.999 | 50 | 119.9943 | 0.99992 | 99.9006 | 0.11987 | -0.0053469 | 0.11998 | 119.767 | 0.999 | 0.1197 | -0.2273 | -0.000922 | -0.000168 | -0.1894 | -0.0922 | -0.1399
freq_sweep | 1 | 3 | 15 | 3 | 120 | 1 | 0.999 | 50 | 119.9943 | 0.99992 | 99.9006 | 0.11987 | -0.0053469 | 0.11998 | 119.767 | 1 | 0.11989 | -0.2273 | 7.8e-05 | 2.3e-05 | -0.1894 | 0.0078 | 0.0194
freq_sweep | 1 | 3 | 16 | 3 | 120 | 1 | 0.999 | 50 | 119.9943 | 0.99992 | 99.9006 | 0.11987 | -0.0053469 | 0.11998 | 119.767 | 0.999 | 0.11973 | -0.2273 | -0.000922 | -0.000136 | -0.1894 | -0.0922 | -0.1132
freq_sweep | 1 | 3 | 17 | 3 | 120 | 1 | 0.999 | 50 | 119.9943 | 0.99992 | 99.9006 | 0.11987 | -0.0053469 | 0.11998 | 119.767 | 1 | 0.11977 | -0.2273 | 7.8e-05 | -9.3e-05 | -0.1894 | 0.0078 | -0.0774
freq_sweep | 1 | 3 | 18 | 3 | 120 | 1 | 0.999 | 50 | 119.9943 | 0.99992 | 99.9006 | 0.11987 | -0.0053469 | 0.11998 | 119.767 | 1 | 0.1198 | -0.2273 | 7.8e-05 | -6.5e-05 | -0.1894 | 0.0078 | -0.054
freq_sweep | 2 | 1 | 1 | 1 | 120 | 1 | 0.999 | 51 | 119.9977 | 0.99999 | 99.9004 | 0.11988 | -0.0053429 | 0.12 | 119.74 | 2.967 | 0.35533 | -0.25769 | 1.967 | 0.23545 | -0.2147 | -1.099 | -1.1956
freq_sweep | 2 | 1 | 2 | 1 | 120 | 1 | 0.999 | 51 | 119.9977 | 0.99999 | 99.9004 | 0.11988 | -0.0053429 | 0.12 | 119.74 | 1.001 | 0.11998 | -0.25769 | 0.00101 | 0.000102 | -0.2147 | 0.101 | 0.0851
freq_sweep | 2 | 1 | 3 | 1 | 120 | 1 | 0.999 | 51 | 119.9977 | 0.99999 | 99.9004 | 0.11988 | -0.0053429 | 0.12 | 119.74 | 1.001 | 0.11997 | -0.25769 | 0.00101 | 8.8e-05 | -0.2147 | 0.101 | 0.0734
freq_sweep | 2 | 1 | 4 | 1 | 120 | 1 | 0.999 | 51 | 119.9977 | 0.99999 | 99.9004 | 0.11988 | -0.0053429 | 0.12 | 119.74 | 1 | 0.11983 | -0.25769 | 1e-05 | -4.8e-05 | -0.2147 | 0.001 | -0.04
freq_sweep | 2 | 1 | 5 | 1 | 120 | 1 | 0.999 | 51 | 119.9977 | 0.99999 | 99.9004 | 0.11988 | -0.0053429 | 0.12 | 119.74 | 1.001 | 0.11997 | -0.25769 | 0.00101 | 8.9e-05 | -0.2147 | 0.101 | 0.0743
freq_sweep | 2 | 1 | 6 | 1 | 120 | 1 | 0.999 | 51 | 119.9977 | 0.99999 | 99.9004 | 0.11988 | -0.0053429 | 0.12 | 119.74 | 0.998 | 0.1196 | -0.25769 | -0.00199 | -0.000281 | -0.2147 | -0.199 | -0.2344
freq_sweep | 2 | 1 | 7 | 2 | 120 | 1 | 0.999 | 51 | 119.9798 | 1.0002 | 99.9037 | 0.11989 | -0.0052629 | 0.12 | 119.756 | 2.978 | 0.35678 | -0.22383 | 1.9778 | 0.2369 | -0.1866 | -0.7507 | -0.79908
freq_sweep | 2 | 1 | 8 | 2 | 120 | 1 | 0.999 | 51 | 119.9798 | 1.0002 | 99.9037 | 0.11989 | -0.0052629 | 0.12 | 119.756 | 1 | 0.11982 | -0.22383 | -0.000175 | -6.7e-05 | -0.1866 | -0.0175 | -0.0562
freq_sweep | 2 | 1 | 9 | 2 | 120 | 1 | 0.999 | 51 | 119.9798 | 1.0002 | 99.9037 | 0.11989 | -0.0052629 | 0.12 | 119.756 | 1 | 0.11987 | -0.22383 | -0.000175 | -1.1e-05 | -0.1866 | -0.0175 | -0.0094
freq_sweep | 2 | 1 | 10 | 2 | 120 | 1 | 0.999 | 51 | 119.9798 | 1.0002 | 99.9037 | 0.11989 | -0.0052629 | 0.12 | 119.756 | 1 | 0.11985 | -0.22383 | -0.000175 | -3.4e-05 | -0.1866 | -0.0175 | -0.0286
freq_sweep | 2 | 1 | 11 | 2 | 120 | 1 | 0.999 | 51 | 119.9798 | 1.0002 | 99.9037 | 0.11989 | -0.0052629 | 0.12 | 119.756 | 1 | 0.11979 | -0.22383 | -0.000175 | -9.8e-05 | -0.1866 | -0.0175 | -0.082
freq_sweep | 2 | 1 | 12 | 2 | 120 | 1 | 0.999 | 51 | 119.9798 | 1.0002 | 99.9037 | 0.11989 | -0.0052629 | 0.12 | 119.756 | 1 | 0.1198 | -0.22383 | -0.000175 | -8.2e-05 | -0.1866 | -0.0175 | -0.0687
freq_sweep | 2 | 1 | 13 | 3 | 120 | 1 | 0.999 | 51 | 119.9917 | 0.99993 | 99.9004 | 0.11986 | -0.0053524 | 0.11998 | 119.777 | 1 | 0.11978 | -0.21468 | 7.5e-05 | -8.7e-05 | -0.1789 | 0.0075 | -0.0728
freq_sweep | 2 | 1 | 14 | 3 | 120 | 1 | 0.999 | 51 | 119.9917 | 0.99993 | 99.9004 | 0.11986 | -0.0053524 | 0.11998 | 119.777 | 0.999 | 0.11972 | -0.21468 | -0.000925 | -0.000146 | -0.1789 | -0.0925 | -0.122
freq_sweep | 2 | 1 | 15 | 3 | 120 | 1 | 0.999 | 51 | 119.9917 | 0.99993 | 99.9004 | 0.11986 | -0.0053524 | 0.11998 | 119.777 | 1 | 0.1199 | -0.21468 | 7.5e-05 | 3.3e-05 | -0.1789 | 0.0075 | 0.0273
freq_sweep | 2 | 1 | 16 | 3 | 120 | 1 | 0.999 | 51 | 119.9917 | 0.99993 | 99.9004 | 0.11986 | -0.0053524 | 0.11998 | 119.777 | 0.999 | 0.11973 | -0.21468 | -0.000925 | -0.00013 | -0.1789 | -0.0925 | -0.1086
freq_sweep | 2 | 1 | 17 | 3 | 120 | 1 | 0.999 | 51 | 119.9917 | 0.99993 | 99.9004 | 0.11986 | -0.0053524 | 0.11998 | 119.777 | 1 | 0.1198 | -0.21468 | 7.5e-05 | -6.6e-05 | -0.1789 | 0.0075 | -0.0553
freq_sweep | 2 | 1 | 18 | 3 | 120 | 1 | 0.999 | 51 | 119.9917 | 0.99993 | 99.9004 | 0.11986 | -0.0053524 | 0.11998 | 119.777 | 1 | 0.11979 | -0.21468 | 7.5e-05 | -7.4e-05 | -0.1789 | 0.0075 | -0.0619
freq_sweep | 2 | 2 | 1 | 1 | 120 | 1 | 0.999 | 51 | 120.0122 | 0.99995 | 99.9006 | 0.11989 | -0.005339 | 0.12001 | 119.806 | 2.968 | 0.35539 | -0.20616 | 1.968 | 0.2355 | -0.1718 | -1.0619 | -1.1871
freq_sweep | 2 | 2 | 2 | 1 | 120 | 1 | 0.999 | 51 | 120.0122 | 0.99995 | 99.9006 | 0.11989 | -0.005339 | 0.12001 | 119.806 | 1.001 | 0.11995 | -0.20616 | 0.001048 | 6.3e-05 | -0.1718 | 0.1048 | 0.0524
freq_sweep | 2 | 2 | 3 | 1 | 120 | 1 | 0.999 | 51 | 120.0122 | 0.99995 | 99.9006 | 0.11989 | -0.005339 | 0.12001 | 119.806 | 1.001 | 0.12001 | -0.20616 | 0.001048 | 0.00012 | -0.1718 | 0.1048 | 0.1
freq_sweep | 2 | 2 | 4 | 1 | 120 | 1 | 0.999 | 51 | 120.0122 | 0.99995 | 99.9006 | 0.11989 | -0.005339 | 0.12001 | 119.806 | 1 | 0.11983 | -0.20616 | 4.8e-05 | -5.5e-05 | -0.1718 | 0.0048 | -0.046
freq_sweep | 2 | 2 | 5 | 1 | 120 | 1 | 0.999 | 51 | 120.0122 | 0.99995 | 99.9006 | 0.11989 | -0.005339 | 0.12001 | 119.806 | 1.001 | 0.11994 | -0.20616 | 0.001048 | 5e-05 | -0.1718 | 0.1048 | 0.0416
freq_sweep | 2 | 2 | 6 | 1 | 120 | 1 | 0.999 | 51 | 120.0122 | 0.99995 | 99.9006 | 0.11989 | -0.005339 | 0.12001 | 119.806 | 0.998 | 0.11967 | -0.20616 | -0.001952 | -0.000215 | -0.1718 | -0.1952 | -0.1795
freq_sweep | 2 | 2 | 7 | 2 | 120 | 1 | 0.999 | 51 | 119.9805 | 1.0002 | 99.9036 | 0.11988 | -0.0052666 | 0.12 | 119.776 | 2.978 | 0.35668 | -0.20454 | 1.9778 | 0.23679 | -0.1705 | -0.74984 | -0.82758
freq_sweep | 2 | 2 | 8 | 2 | 120 | 1 | 0.999 | 51 | 119.9805 | 1.0002 | 99.9036 | 0.11988 | -0.0052666 | 0.12 | 119.776 | 1 | 0.11981 | -0.20454 | -0.000166 | -7e-05 | -0.1705 | -0.0166 | -0.0582
freq_sweep | 2 | 2 | 9 | 2 | 120 | 1 | 0.999 | 51 | 119.9805 | 1.0002 | 99.9036 | 0.11988 | -0.0052666 | 0.12 | 119.776 | 1.001 | 0.1199 | -0.20454 | 0.000834 | 1.9e-05 | -0.1705 | 0.0834 | 0.016
freq_sweep | 2 | 2 | 10 | 2 | 120 | 1 | 0.999 | 51 | 119.9805 | 1.0002 | 99.9036 | 0.11988 | -0.0052666 | 0.12 | 119.776 | 1 | 0.11983 | -0.20454 | -0.000166 | -5.8e-05 | -0.1705 | -0.0166 | -0.0482
freq_sweep | 2 | 2 | 11 | 2 | 120 | 1 | 0.999 | 51 | 119.9805 | 1.0002 | 99.9036 | 0.11988 | -0.0052666 | 0.12 | 119.776 | 1 | 0.11978 | -0.20454 | -0.000166 | -0.000107 | -0.1705 | -0.0166 | -0.0891
freq_sweep | 2 | 2 | 12 | 2 | 120 | 1 | 0.999 | 51 | 119.9805 | 1.0002 | 99.9036 | 0.11988 | -0.0052666 | 0.12 | 119.776 | 1 | 0.11983 | -0.20454 | -0.000166 | -5.2e-05 | -0.1705 | -0.0166 | -0.0432
freq_sweep | 2 | 2 | 13 | 3 | 120 | 1 | 0.999 | 51 | 119.9923 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.775 | 1 | 0.11979 | -0.21729 | 8.7e-05 | -6.8e-05 | -0.1811 | 0.0087 | -0.0563
freq_sweep | 2 | 2 | 14 | 3 | 120 | 1 | 0.999 | 51 | 119.9923 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.775 | 0.999 | 0.11974 | -0.21729 | -0.000913 | -0.000122 | -0.1811 | -0.0913 | -0.1014
freq_sweep | 2 | 2 | 15 | 3 | 120 | 1 | 0.999 | 51 | 119.9923 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.775 | 1 | 0.11989 | -0.21729 | 8.7e-05 | 2.9e-05 | -0.1811 | 0.0087 | 0.0246
freq_sweep | 2 | 2 | 16 | 3 | 120 | 1 | 0.999 | 51 | 119.9923 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.775 | 0.999 | 0.11974 | -0.21729 | -0.000913 | -0.00012 | -0.1811 | -0.0913 | -0.0997
freq_sweep | 2 | 2 | 17 | 3 | 120 | 1 | 0.999 | 51 | 119.9923 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.775 | 1 | 0.11981 | -0.21729 | 8.7e-05 | -4.8e-05 | -0.1811 | 0.0087 | -0.0396
freq_sweep | 2 | 2 | 18 | 3 | 120 | 1 | 0.999 | 51 | 119.9923 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.775 | 0.999 | 0.11976 | -0.21729 | -0.000913 | -9.8e-05 | -0.1811 | -0.0913 | -0.0813
freq_sweep | 2 | 3 | 1 | 1 | 120 | 1 | 0.999 | 51 | 120.0066 | 0.99998 | 99.9006 | 0.11988 | -0.0053409 | 0.12 | 119.754 | 2.967 | 0.35541 | -0.25261 | 1.967 | 0.23553 | -0.2105 | -1.0979 | -1.1801
freq_sweep | 2 | 3 | 2 | 1 | 120 | 1 | 0.999 | 51 | 120.0066 | 0.99998 | 99.9006 | 0.11988 | -0.0053409 | 0.12 | 119.754 | 1.001 | 0.1199 | -0.25261 | 0.001021 | 1.6e-05 | -0.2105 | 0.1021 | 0.0135
freq_sweep | 2 | 3 | 3 | 1 | 120 | 1 | 0.999 | 51 | 120.0066 | 0.99998 | 99.9006 | 0.11988 | -0.0053409 | 0.12 | 119.754 | 1.001 | 0.11997 | -0.25261 | 0.001021 | 8.1e-05 | -0.2105 | 0.1021 | 0.0678
freq_sweep | 2 | 3 | 4 | 1 | 120 | 1 | 0.999 | 51 | 120.0066 | 0.99998 | 99.9006 | 0.11988 | -0.0053409 | 0.12 | 119.754 | 1 | 0.11984 | -0.25261 | 2.1e-05 | -4e-05 | -0.2105 | 0.0021 | -0.0332
freq_sweep | 2 | 3 | 5 | 1 | 120 | 1 | 0.999 | 51 | 120.0066 | 0.99998 | 99.9006 | 0.11988 | -0.0053409 | 0.12 | 119.754 | 1.001 | 0.11989 | -0.25261 | 0.001021 | 6e-06 | -0.2105 | 0.1021 | 0.0052
freq_sweep | 2 | 3 | 6 | 1 | 120 | 1 | 0.999 | 51 | 120.0066 | 0.99998 | 99.9006 | 0.11988 | -0.0053409 | 0.12 | 119.754 | 0.998 | 0.11957 | -0.25261 | -0.001979 | -0.000319 | -0.2105 | -0.1979 | -0.2659
freq_sweep | 2 | 3 | 7 | 2 | 120 | 1 | 0.999 | 51 | 119.9784 | 1.0002 | 99.9036 | 0.11988 | -0.0052662 | 0.12 | 119.759 | 2.978 | 0.3567 | -0.21936 | 1.9778 | 0.23681 | -0.1828 | -0.75106 | -0.82121
freq_sweep | 2 | 3 | 8 | 2 | 120 | 1 | 0.999 | 51 | 119.9784 | 1.0002 | 99.9036 | 0.11988 | -0.0052662 | 0.12 | 119.759 | 1 | 0.11983 | -0.21936 | -0.000179 | -5.6e-05 | -0.1828 | -0.0179 | -0.0468
freq_sweep | 2 | 3 | 9 | 2 | 120 | 1 | 0.999 | 51 | 119.9784 | 1.0002 | 99.9036 | 0.11988 | -0.0052662 | 0.12 | 119.759 | 1 | 0.11987 | -0.21936 | -0.000179 | -1.7e-05 | -0.1828 | -0.0179 | -0.0143
freq_sweep | 2 | 3 | 10 | 2 | 120 | 1 | 0.999 | 51 | 119.9784 | 1.0002 | 99.9036 | 0.11988 | -0.0052662 | 0.12 | 119.759 | 1 | 0.11984 | -0.21936 | -0.000179 | -4.6e-05 | -0.1828 | -0.0179 | -0.0385
freq_sweep | 2 | 3 | 11 | 2 | 120 | 1 | 0.999 | 51 | 119.9784 | 1.0002 | 99.9036 | 0.11988 | -0.0052662 | 0.12 | 119.759 | 1 | 0.11979 | -0.21936 | -0.000179 | -9.7e-05 | -0.1828 | -0.0179 | -0.081
freq_sweep | 2 | 3 | 12 | 2 | 120 | 1 | 0.999 | 51 | 119.9784 | 1.0002 | 99.9036 | 0.11988 | -0.0052662 | 0.12 | 119.759 | 1 | 0.1198 | -0.21936 | -0.000179 | -8e-05 | -0.1828 | -0.0179 | -0.0669
freq_sweep | 2 | 3 | 13 | 3 | 120 | 1 | 0.999 | 51 | 119.993 | 0.99991 | 99.9006 | 0.11986 | -0.0053491 | 0.11998 | 119.783 | 1 | 0.1198 | -0.20999 | 8.7e-05 | -6.3e-05 | -0.175 | 0.0087 | -0.0528
freq_sweep | 2 | 3 | 14 | 3 | 120 | 1 | 0.999 | 51 | 119.993 | 0.99991 | 99.9006 | 0.11986 | -0.0053491 | 0.11998 | 119.783 | 0.999 | 0.11971 | -0.20999 | -0.000913 | -0.000151 | -0.175 | -0.0913 | -0.1262
freq_sweep | 2 | 3 | 15 | 3 | 120 | 1 | 0.999 | 51 | 119.993 | 0.99991 | 99.9006 | 0.11986 | -0.0053491 | 0.11998 | 119.783 | 1 | 0.11989 | -0.20999 | 8.7e-05 | 2.4e-05 | -0.175 | 0.0087 | 0.0198
freq_sweep | 2 | 3 | 16 | 3 | 120 | 1 | 0.999 | 51 | 119.993 | 0.99991 | 99.9006 | 0.11986 | -0.0053491 | 0.11998 | 119.783 | 0.999 | 0.11976 | -0.20999 | -0.000913 | -9.9e-05 | -0.175 | -0.0913 | -0.0828
freq_sweep | 2 | 3 | 17 | 3 | 120 | 1 | 0.999 | 51 | 119.993 | 0.99991 | 99.9006 | 0.11986 | -0.0053491 | 0.11998 | 119.783 | 1 | 0.11977 | -0.20999 | 8.7e-05 | -9e-05 | -0.175 | 0.0087 | -0.0753
freq_sweep | 2 | 3 | 18 | 3 | 120 | 1 | 0.999 | 51 | 119.993 | 0.99991 | 99.9006 | 0.11986 | -0.0053491 | 0.11998 | 119.783 | 1 | 0.11978 | -0.20999 | 8.7e-05 | -8.6e-05 | -0.175 | 0.0087 | -0.072
freq_sweep | 3 | 1 | 1 | 1 | 120 | 1 | 0.999 | 52 | 120.013 | 0.99996 | 99.9005 | 0.11989 | -0.0053435 | 0.12001 | 119.721 | 2.968 | 0.35532 | -0.29203 | 1.968 | 0.23543 | -0.2433 | -1.063 | -1.2079
freq_sweep | 3 | 1 | 2 | 1 | 120 | 1 | 0.999 | 52 | 120.013 | 0.99996 | 99.9005 | 0.11989 | -0.0053435 | 0.12001 | 119.721 | 1.001 | 0.1199 | -0.29203 | 0.001037 | 1.6e-05 | -0.2433 | 0.1037 | 0.0132
freq_sweep | 3 | 1 | 3 | 1 | 120 | 1 | 0.999 | 52 | 120.013 | 0.99996 | 99.9005 | 0.11989 | -0.0053435 | 0.12001 | 119.721 | 1.001 | 0.11994 | -0.29203 | 0.001037 | 5.1e-05 | -0.2433 | 0.1037 | 0.0424
freq_sweep | 3 | 1 | 4 | 1 | 120 | 1 | 0.999 | 52 | 120.013 | 0.99996 | 99.9005 | 0.11989 | -0.0053435 | 0.12001 | 119.721 | 1 | 0.1198 | -0.29203 | 3.7e-05 | -8.5e-05 | -0.2433 | 0.0037 | -0.0711
freq_sweep | 3 | 1 | 5 | 1 | 120 | 1 | 0.999 | 52 | 120.013 | 0.99996 | 99.9005 | 0.11989 | -0.0053435 | 0.12001 | 119.721 | 1.001 | 0.11993 | -0.29203 | 0.001037 | 4.5e-05 | -0.2433 | 0.1037 | 0.0374
freq_sweep | 3 | 1 | 6 | 1 | 120 | 1 | 0.999 | 52 | 120.013 | 0.99996 | 99.9005 | 0.11989 | -0.0053435 | 0.12001 | 119.721 | 0.998 | 0.11961 | -0.29203 | -0.001963 | -0.000282 | -0.2433 | -0.1963 | -0.2354
freq_sweep | 3 | 1 | 7 | 2 | 120 | 1 | 0.999 | 52 | 119.9832 | 1.0002 | 99.9038 | 0.11989 | -0.005263 | 0.12 | 119.766 | 2.978 | 0.35678 | -0.21725 | 1.9778 | 0.23689 | -0.1811 | -0.75005 | -0.80158
freq_sweep | 3 | 1 | 8 | 2 | 120 | 1 | 0.999 | 52 | 119.9832 | 1.0002 | 99.9038 | 0.11989 | -0.005263 | 0.12 | 119.766 | 1 | 0.11982 | -0.21725 | -0.000168 | -6.7e-05 | -0.1811 | -0.0168 | -0.0559
freq_sweep | 3 | 1 | 9 | 2 | 120 | 1 | 0.999 | 52 | 119.9832 | 1.0002 | 99.9038 | 0.11989 | -0.005263 | 0.12 | 119.766 | 1.001 | 0.1199 | -0.21725 | 0.000832 | 1.4e-05 | -0.1811 | 0.0831 | 0.0117
freq_sweep | 3 | 1 | 10 | 2 | 120 | 1 | 0.999 | 52 | 119.9832 | 1.0002 | 99.9038 | 0.11989 | -0.005263 | 0.12 | 119.766 | 1 | 0.11985 | -0.21725 | -0.000168 | -3.4e-05 | -0.1811 | -0.0168 | -0.0284
freq_sweep | 3 | 1 | 11 | 2 | 120 | 1 | 0.999 | 52 | 119.9832 | 1.0002 | 99.9038 | 0.11989 | -0.005263 | 0.12 | 119.766 | 1 | 0.11977 | -0.21725 | -0.000168 | -0.000118 | -0.1811 | -0.0168 | -0.0984
freq_sweep | 3 | 1 | 12 | 2 | 120 | 1 | 0.999 | 52 | 119.9832 | 1.0002 | 99.9038 | 0.11989 | -0.005263 | 0.12 | 119.766 | 1 | 0.11982 | -0.21725 | -0.000168 | -6.6e-05 | -0.1811 | -0.0168 | -0.0551
freq_sweep | 3 | 1 | 13 | 3 | 120 | 1 | 0.999 | 52 | 119.9875 | 0.99992 | 99.9006 | 0.11986 | -0.0053482 | 0.11998 | 119.765 | 1 | 0.11978 | -0.22255 | 8e-05 | -7.5e-05 | -0.1855 | 0.008 | -0.0623
freq_sweep | 3 | 1 | 14 | 3 | 120 | 1 | 0.999 | 52 | 119.9875 | 0.99992 | 99.9006 | 0.11986 | -0.0053482 | 0.11998 | 119.765 | 0.999 | 0.11972 | -0.22255 | -0.00092 | -0.000144 | -0.1855 | -0.092 | -0.1199
freq_sweep | 3 | 1 | 15 | 3 | 120 | 1 | 0.999 | 52 | 119.9875 | 0.99992 | 99.9006 | 0.11986 | -0.0053482 | 0.11998 | 119.765 | 1 | 0.1199 | -0.22255 | 8e-05 | 3.7e-05 | -0.1855 | 0.008 | 0.0311
freq_sweep | 3 | 1 | 16 | 3 | 120 | 1 | 0.999 | 52 | 119.9875 | 0.99992 | 99.9006 | 0.11986 | -0.0053482 | 0.11998 | 119.765 | 0.999 | 0.11973 | -0.22255 | -0.00092 | -0.000125 | -0.1855 | -0.092 | -0.104
freq_sweep | 3 | 1 | 17 | 3 | 120 | 1 | 0.999 | 52 | 119.9875 | 0.99992 | 99.9006 | 0.11986 | -0.0053482 | 0.11998 | 119.765 | 1 | 0.11979 | -0.22255 | 8e-05 | -7.3e-05 | -0.1855 | 0.008 | -0.0606
freq_sweep | 3 | 1 | 18 | 3 | 120 | 1 | 0.999 | 52 | 119.9875 | 0.99992 | 99.9006 | 0.11986 | -0.0053482 | 0.11998 | 119.765 | 1 | 0.1198 | -0.22255 | 8e-05 | -6.1e-05 | -0.1855 | 0.008 | -0.0506
freq_sweep | 3 | 2 | 1 | 1 | 120 | 1 | 0.999 | 52 | 120.0194 | 0.99997 | 99.9006 | 0.1199 | -0.0053422 | 0.12002 | 119.752 | 2.968 | 0.35539 | -0.2674 | 1.968 | 0.2355 | -0.2228 | -1.0635 | -1.1938
freq_sweep | 3 | 2 | 2 | 1 | 120 | 1 | 0.999 | 52 | 120.0194 | 0.99997 | 99.9006 | 0.1199 | -0.0053422 | 0.12002 | 119.752 | 1.001 | 0.11993 | -0.2674 | 0.001032 | 2.9e-05 | -0.2228 | 0.1032 | 0.024
freq_sweep | 3 | 2 | 3 | 1 | 120 | 1 | 0.999 | 52 | 120.0194 | 0.99997 | 99.9006 | 0.1199 | -0.0053422 | 0.12002 | 119.752 | 1.001 | 0.11996 | -0.2674 | 0.001032 | 6.2e-05 | -0.2228 | 0.1032 | 0.0515
freq_sweep | 3 | 2 | 4 | 1 | 120 | 1 | 0.999 | 52 | 120.0194 | 0.99997 | 99.9006 | 0.1199 | -0.0053422 | 0.12002 | 119.752 | 1 | 0.11984 | -0.2674 | 3.2e-05 | -5.2e-05 | -0.2228 | 0.0032 | -0.0436
freq_sweep | 3 | 2 | 5 | 1 | 120 | 1 | 0.999 | 52 | 120.0194 | 0.99997 | 99.9006 | 0.1199 | -0.0053422 | 0.12002 | 119.752 | 1.001 | 0.11994 | -0.2674 | 0.001032 | 4.4e-05 | -0.2228 | 0.1032 | 0.0365
freq_sweep | 3 | 2 | 6 | 1 | 120 | 1 | 0.999 | 52 | 120.0194 | 0.99997 | 99.9006 | 0.1199 | -0.0053422 | 0.12002 | 119.752 | 0.998 | 0.11961 | -0.2674 | -0.001968 | -0.000281 | -0.2228 | -0.1968 | -0.2346
freq_sweep | 3 | 2 | 7 | 2 | 120 | 1 | 0.999 | 52 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052604 | 0.12 | 119.764 | 2.978 | 0.35672 | -0.21672 | 1.9778 | 0.23683 | -0.1806 | -0.74957 | -0.81657
freq_sweep | 3 | 2 | 8 | 2 | 120 | 1 | 0.999 | 52 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052604 | 0.12 | 119.764 | 1 | 0.11983 | -0.21672 | -0.000164 | -5.9e-05 | -0.1806 | -0.0164 | -0.0492
freq_sweep | 3 | 2 | 9 | 2 | 120 | 1 | 0.999 | 52 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052604 | 0.12 | 119.764 | 1 | 0.11989 | -0.21672 | -0.000164 | 0 | -0.1806 | -0.0164 | 0
freq_sweep | 3 | 2 | 10 | 2 | 120 | 1 | 0.999 | 52 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052604 | 0.12 | 119.764 | 1 | 0.11982 | -0.21672 | -0.000164 | -6.8e-05 | -0.1806 | -0.0164 | -0.0567
freq_sweep | 3 | 2 | 11 | 2 | 120 | 1 | 0.999 | 52 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052604 | 0.12 | 119.764 | 1 | 0.11979 | -0.21672 | -0.000164 | -9.3e-05 | -0.1806 | -0.0164 | -0.0775
freq_sweep | 3 | 2 | 12 | 2 | 120 | 1 | 0.999 | 52 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052604 | 0.12 | 119.764 | 1 | 0.11981 | -0.21672 | -0.000164 | -7.3e-05 | -0.1806 | -0.0164 | -0.0608
freq_sweep | 3 | 2 | 13 | 3 | 120 | 1 | 0.999 | 52 | 119.99 | 0.99992 | 99.9006 | 0.11986 | -0.0053473 | 0.11998 | 119.784 | 1 | 0.11978 | -0.206 | 8.3e-05 | -7.9e-05 | -0.1717 | 0.0083 | -0.0657
freq_sweep | 3 | 2 | 14 | 3 | 120 | 1 | 0.999 | 52 | 119.99 | 0.99992 | 99.9006 | 0.11986 | -0.0053473 | 0.11998 | 119.784 | 0.999 | 0.11976 | -0.206 | -0.000917 | -9.7e-05 | -0.1717 | -0.0917 | -0.0808
freq_sweep | 3 | 2 | 15 | 3 | 120 | 1 | 0.999 | 52 | 119.99 | 0.99992 | 99.9006 | 0.11986 | -0.0053473 | 0.11998 | 119.784 | 1 | 0.11987 | -0.206 | 8.3e-05 | 1.2e-05 | -0.1717 | 0.0083 | 0.0102
freq_sweep | 3 | 2 | 16 | 3 | 120 | 1 | 0.999 | 52 | 119.99 | 0.99992 | 99.9006 | 0.11986 | -0.0053473 | 0.11998 | 119.784 | 0.999 | 0.11975 | -0.206 | -0.000917 | -0.000114 | -0.1717 | -0.0917 | -0.0949
freq_sweep | 3 | 2 | 17 | 3 | 120 | 1 | 0.999 | 52 | 119.99 | 0.99992 | 99.9006 | 0.11986 | -0.0053473 | 0.11998 | 119.784 | 1 | 0.1198 | -0.206 | 8.3e-05 | -5.6e-05 | -0.1717 | 0.0083 | -0.0465
freq_sweep | 3 | 2 | 18 | 3 | 120 | 1 | 0.999 | 52 | 119.99 | 0.99992 | 99.9006 | 0.11986 | -0.0053473 | 0.11998 | 119.784 | 1 | 0.11977 | -0.206 | 8.3e-05 | -8.7e-05 | -0.1717 | 0.0083 | -0.0724
freq_sweep | 3 | 3 | 1 | 1 | 120 | 1 | 0.999 | 52 | 120.0096 | 0.99999 | 99.9006 | 0.11989 | -0.0053388 | 0.12001 | 119.721 | 2.967 | 0.3554 | -0.28856 | 1.967 | 0.23551 | -0.2404 | -1.0992 | -1.1864
freq_sweep | 3 | 3 | 2 | 1 | 120 | 1 | 0.999 | 52 | 120.0096 | 0.99999 | 99.9006 | 0.11989 | -0.0053388 | 0.12001 | 119.721 | 1.001 | 0.11988 | -0.28856 | 0.001008 | -1.1e-05 | -0.2404 | 0.1008 | -0.0094
freq_sweep | 3 | 3 | 3 | 1 | 120 | 1 | 0.999 | 52 | 120.0096 | 0.99999 | 99.9006 | 0.11989 | -0.0053388 | 0.12001 | 119.721 | 1.001 | 0.11995 | -0.28856 | 0.001008 | 5.8e-05 | -0.2404 | 0.1008 | 0.0481
freq_sweep | 3 | 3 | 4 | 1 | 120 | 1 | 0.999 | 52 | 120.0096 | 0.99999 | 99.9006 | 0.11989 | -0.0053388 | 0.12001 | 119.721 | 1 | 0.11984 | -0.28856 | 8e-06 | -5e-05 | -0.2404 | 0.0008 | -0.042
freq_sweep | 3 | 3 | 5 | 1 | 120 | 1 | 0.999 | 52 | 120.0096 | 0.99999 | 99.9006 | 0.11989 | -0.0053388 | 0.12001 | 119.721 | 1.001 | 0.11985 | -0.28856 | 0.001008 | -3.8e-05 | -0.2404 | 0.1008 | -0.032
freq_sweep | 3 | 3 | 6 | 1 | 120 | 1 | 0.999 | 52 | 120.0096 | 0.99999 | 99.9006 | 0.11989 | -0.0053388 | 0.12001 | 119.721 | 0.998 | 0.11959 | -0.28856 | -0.001992 | -0.000298 | -0.2404 | -0.1992 | -0.2488
freq_sweep | 3 | 3 | 7 | 2 | 120 | 1 | 0.999 | 52 | 119.9803 | 1.0002 | 99.9037 | 0.11988 | -0.0052646 | 0.12 | 119.757 | 2.978 | 0.3568 | -0.22326 | 1.9778 | 0.23692 | -0.1861 | -0.7489 | -0.79232
freq_sweep | 3 | 3 | 8 | 2 | 120 | 1 | 0.999 | 52 | 119.9803 | 1.0002 | 99.9037 | 0.11988 | -0.0052646 | 0.12 | 119.757 | 1 | 0.11982 | -0.22326 | -0.000157 | -6.5e-05 | -0.1861 | -0.0157 | -0.0538
freq_sweep | 3 | 3 | 9 | 2 | 120 | 1 | 0.999 | 52 | 119.9803 | 1.0002 | 99.9037 | 0.11988 | -0.0052646 | 0.12 | 119.757 | 1 | 0.11987 | -0.22326 | -0.000157 | -1.7e-05 | -0.1861 | -0.0157 | -0.0138
freq_sweep | 3 | 3 | 10 | 2 | 120 | 1 | 0.999 | 52 | 119.9803 | 1.0002 | 99.9037 | 0.11988 | -0.0052646 | 0.12 | 119.757 | 1 | 0.11986 | -0.22326 | -0.000157 | -2.8e-05 | -0.1861 | -0.0157 | -0.023
freq_sweep | 3 | 3 | 11 | 2 | 120 | 1 | 0.999 | 52 | 119.9803 | 1.0002 | 99.9037 | 0.11988 | -0.0052646 | 0.12 | 119.757 | 1 | 0.1198 | -0.22326 | -0.000157 | -8.8e-05 | -0.1861 | -0.0157 | -0.073
freq_sweep | 3 | 3 | 12 | 2 | 120 | 1 | 0.999 | 52 | 119.9803 | 1.0002 | 99.9037 | 0.11988 | -0.0052646 | 0.12 | 119.757 | 1 | 0.1198 | -0.22326 | -0.000157 | -8.3e-05 | -0.1861 | -0.0157 | -0.0688
freq_sweep | 3 | 3 | 13 | 3 | 120 | 1 | 0.999 | 52 | 119.9929 | 0.99994 | 99.9005 | 0.11987 | -0.0053506 | 0.11999 | 119.774 | 1 | 0.11978 | -0.21893 | 6.3e-05 | -8.8e-05 | -0.1825 | 0.0063 | -0.0734
freq_sweep | 3 | 3 | 14 | 3 | 120 | 1 | 0.999 | 52 | 119.9929 | 0.99994 | 99.9005 | 0.11987 | -0.0053506 | 0.11999 | 119.774 | 0.999 | 0.11972 | -0.21893 | -0.000937 | -0.000148 | -0.1825 | -0.0937 | -0.1235
freq_sweep | 3 | 3 | 15 | 3 | 120 | 1 | 0.999 | 52 | 119.9929 | 0.99994 | 99.9005 | 0.11987 | -0.0053506 | 0.11999 | 119.774 | 1 | 0.11989 | -0.21893 | 6.3e-05 | 2.1e-05 | -0.1825 | 0.0063 | 0.0175
freq_sweep | 3 | 3 | 16 | 3 | 120 | 1 | 0.999 | 52 | 119.9929 | 0.99994 | 99.9005 | 0.11987 | -0.0053506 | 0.11999 | 119.774 | 0.999 | 0.11973 | -0.21893 | -0.000937 | -0.00014 | -0.1825 | -0.0937 | -0.1168
freq_sweep | 3 | 3 | 17 | 3 | 120 | 1 | 0.999 | 52 | 119.9929 | 0.99994 | 99.9005 | 0.11987 | -0.0053506 | 0.11999 | 119.774 | 1 | 0.11977 | -0.21893 | 6.3e-05 | -0.0001 | -0.1825 | 0.0063 | -0.0835
freq_sweep | 3 | 3 | 18 | 3 | 120 | 1 | 0.999 | 52 | 119.9929 | 0.99994 | 99.9005 | 0.11987 | -0.0053506 | 0.11999 | 119.774 | 1 | 0.11978 | -0.21893 | 6.3e-05 | -8.8e-05 | -0.1825 | 0.0063 | -0.0734
freq_sweep | 4 | 1 | 1 | 1 | 120 | 1 | 0.999 | 53 | 120.0096 | 0.99998 | 99.9007 | 0.11989 | -0.005339 | 0.12001 | 119.74 | 2.968 | 0.35539 | -0.26959 | 1.968 | 0.2355 | -0.2246 | -1.0649 | -1.1886
freq_sweep | 4 | 1 | 2 | 1 | 120 | 1 | 0.999 | 53 | 120.0096 | 0.99998 | 99.9007 | 0.11989 | -0.005339 | 0.12001 | 119.74 | 1.001 | 0.11998 | -0.26959 | 0.001018 | 9.6e-05 | -0.2246 | 0.1018 | 0.0798
freq_sweep | 4 | 1 | 3 | 1 | 120 | 1 | 0.999 | 53 | 120.0096 | 0.99998 | 99.9007 | 0.11989 | -0.005339 | 0.12001 | 119.74 | 1.001 | 0.11996 | -0.26959 | 0.001018 | 6.9e-05 | -0.2246 | 0.1018 | 0.0573
freq_sweep | 4 | 1 | 4 | 1 | 120 | 1 | 0.999 | 53 | 120.0096 | 0.99998 | 99.9007 | 0.11989 | -0.005339 | 0.12001 | 119.74 | 1 | 0.11984 | -0.26959 | 1.8e-05 | -5e-05 | -0.2246 | 0.0018 | -0.042
freq_sweep | 4 | 1 | 5 | 1 | 120 | 1 | 0.999 | 53 | 120.0096 | 0.99998 | 99.9007 | 0.11989 | -0.005339 | 0.12001 | 119.74 | 1.001 | 0.11995 | -0.26959 | 0.001018 | 6.6e-05 | -0.2246 | 0.1018 | 0.0548
freq_sweep | 4 | 1 | 6 | 1 | 120 | 1 | 0.999 | 53 | 120.0096 | 0.99998 | 99.9007 | 0.11989 | -0.005339 | 0.12001 | 119.74 | 0.998 | 0.1196 | -0.26959 | -0.001982 | -0.000289 | -0.2246 | -0.1982 | -0.2414
freq_sweep | 4 | 1 | 7 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052628 | 0.12 | 119.775 | 2.978 | 0.35678 | -0.2057 | 1.9778 | 0.23689 | -0.1714 | -0.74835 | -0.79916
freq_sweep | 4 | 1 | 8 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052628 | 0.12 | 119.775 | 1 | 0.11981 | -0.2057 | -0.000151 | -7.7e-05 | -0.1714 | -0.0151 | -0.0646
freq_sweep | 4 | 1 | 9 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052628 | 0.12 | 119.775 | 1 | 0.11987 | -0.2057 | -0.000151 | -9e-06 | -0.1714 | -0.0151 | -0.0078
freq_sweep | 4 | 1 | 10 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052628 | 0.12 | 119.775 | 1 | 0.11986 | -0.2057 | -0.000151 | -1.8e-05 | -0.1714 | -0.0151 | -0.0153
freq_sweep | 4 | 1 | 11 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052628 | 0.12 | 119.775 | 1 | 0.11977 | -0.2057 | -0.000151 | -0.00011 | -0.1714 | -0.0151 | -0.0921
freq_sweep | 4 | 1 | 12 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052628 | 0.12 | 119.775 | 1 | 0.1198 | -0.2057 | -0.000151 | -7.8e-05 | -0.1714 | -0.0151 | -0.0654
freq_sweep | 4 | 1 | 13 | 3 | 120 | 1 | 0.999 | 53 | 119.9927 | 0.99991 | 99.9005 | 0.11986 | -0.0053499 | 0.11998 | 119.78 | 1 | 0.1198 | -0.21271 | 9e-05 | -6.1e-05 | -0.1773 | 0.009 | -0.0505
freq_sweep | 4 | 1 | 14 | 3 | 120 | 1 | 0.999 | 53 | 119.9927 | 0.99991 | 99.9005 | 0.11986 | -0.0053499 | 0.11998 | 119.78 | 0.999 | 0.11973 | -0.21271 | -0.00091 | -0.000131 | -0.1773 | -0.091 | -0.1089
freq_sweep | 4 | 1 | 15 | 3 | 120 | 1 | 0.999 | 53 | 119.9927 | 0.99991 | 99.9005 | 0.11986 | -0.0053499 | 0.11998 | 119.78 | 1.001 | 0.11991 | -0.21271 | 0.00109 | 4.6e-05 | -0.1773 | 0.109 | 0.0388
freq_sweep | 4 | 1 | 16 | 3 | 120 | 1 | 0.999 | 53 | 119.9927 | 0.99991 | 99.9005 | 0.11986 | -0.0053499 | 0.11998 | 119.78 | 0.999 | 0.11974 | -0.21271 | -0.00091 | -0.000126 | -0.1773 | -0.091 | -0.1047
freq_sweep | 4 | 1 | 17 | 3 | 120 | 1 | 0.999 | 53 | 119.9927 | 0.99991 | 99.9005 | 0.11986 | -0.0053499 | 0.11998 | 119.78 | 1 | 0.11979 | -0.21271 | 9e-05 | -7.5e-05 | -0.1773 | 0.009 | -0.0622
freq_sweep | 4 | 1 | 18 | 3 | 120 | 1 | 0.999 | 53 | 119.9927 | 0.99991 | 99.9005 | 0.11986 | -0.0053499 | 0.11998 | 119.78 | 1 | 0.1198 | -0.21271 | 9e-05 | -6.1e-05 | -0.1773 | 0.009 | -0.0505
freq_sweep | 4 | 2 | 1 | 1 | 120 | 1 | 0.999 | 53 | 120.0153 | 1 | 99.9009 | 0.1199 | -0.0053398 | 0.12002 | 119.707 | 2.968 | 0.35543 | -0.30834 | 1.968 | 0.23553 | -0.2569 | -1.0675 | -1.1861
freq_sweep | 4 | 2 | 2 | 1 | 120 | 1 | 0.999 | 53 | 120.0153 | 1 | 99.9009 | 0.1199 | -0.0053398 | 0.12002 | 119.707 | 1.001 | 0.11989 | -0.30834 | 0.000991 | -8e-06 | -0.2569 | 0.0991 | -0.0071
freq_sweep | 4 | 2 | 3 | 1 | 120 | 1 | 0.999 | 53 | 120.0153 | 1 | 99.9009 | 0.1199 | -0.0053398 | 0.12002 | 119.707 | 1.001 | 0.11993 | -0.30834 | 0.000991 | 3.8e-05 | -0.2569 | 0.0991 | 0.0313
freq_sweep | 4 | 2 | 4 | 1 | 120 | 1 | 0.999 | 53 | 120.0153 | 1 | 99.9009 | 0.1199 | -0.0053398 | 0.12002 | 119.707 | 1.001 | 0.11986 | -0.30834 | 0.000991 | -3.4e-05 | -0.2569 | 0.0991 | -0.0287
freq_sweep | 4 | 2 | 5 | 1 | 120 | 1 | 0.999 | 53 | 120.0153 | 1 | 99.9009 | 0.1199 | -0.0053398 | 0.12002 | 119.707 | 1.001 | 0.1199 | -0.30834 | 0.000991 | 3e-06 | -0.2569 | 0.0991 | 0.0021
freq_sweep | 4 | 2 | 6 | 1 | 120 | 1 | 0.999 | 53 | 120.0153 | 1 | 99.9009 | 0.1199 | -0.0053398 | 0.12002 | 119.707 | 0.998 | 0.11958 | -0.30834 | -0.002009 | -0.000317 | -0.2569 | -0.2009 | -0.2648
freq_sweep | 4 | 2 | 7 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052616 | 0.12 | 119.759 | 2.978 | 0.35673 | -0.22168 | 1.9779 | 0.23685 | -0.1848 | -0.74825 | -0.81187
freq_sweep | 4 | 2 | 8 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052616 | 0.12 | 119.759 | 1 | 0.11981 | -0.22168 | -0.00015 | -7.1e-05 | -0.1848 | -0.015 | -0.0595
freq_sweep | 4 | 2 | 9 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052616 | 0.12 | 119.759 | 1.001 | 0.11991 | -0.22168 | 0.00085 | 2.3e-05 | -0.1848 | 0.085 | 0.0189
freq_sweep | 4 | 2 | 10 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052616 | 0.12 | 119.759 | 1 | 0.11984 | -0.22168 | -0.00015 | -4.7e-05 | -0.1848 | -0.015 | -0.0395
freq_sweep | 4 | 2 | 11 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052616 | 0.12 | 119.759 | 1 | 0.11977 | -0.22168 | -0.00015 | -0.000115 | -0.1848 | -0.015 | -0.0962
freq_sweep | 4 | 2 | 12 | 2 | 120 | 1 | 0.999 | 53 | 119.9807 | 1.0002 | 99.9038 | 0.11988 | -0.0052616 | 0.12 | 119.759 | 1 | 0.11983 | -0.22168 | -0.00015 | -4.8e-05 | -0.1848 | -0.015 | -0.0403
freq_sweep | 4 | 2 | 13 | 3 | 120 | 1 | 0.999 | 53 | 119.9887 | 0.9999 | 99.9006 | 0.11986 | -0.0053481 | 0.11998 | 119.779 | 1 | 0.11978 | -0.20966 | 0.000102 | -8e-05 | -0.1747 | 0.0102 | -0.0669
freq_sweep | 4 | 2 | 14 | 3 | 120 | 1 | 0.999 | 53 | 119.9887 | 0.9999 | 99.9006 | 0.11986 | -0.0053481 | 0.11998 | 119.779 | 0.999 | 0.11973 | -0.20966 | -0.000898 | -0.000126 | -0.1747 | -0.0898 | -0.1053
freq_sweep | 4 | 2 | 15 | 3 | 120 | 1 | 0.999 | 53 | 119.9887 | 0.9999 | 99.9006 | 0.11986 | -0.0053481 | 0.11998 | 119.779 | 1 | 0.11987 | -0.20966 | 0.000102 | 1.4e-05 | -0.1747 | 0.0102 | 0.0115
freq_sweep | 4 | 2 | 16 | 3 | 120 | 1 | 0.999 | 53 | 119.9887 | 0.9999 | 99.9006 | 0.11986 | -0.0053481 | 0.11998 | 119.779 | 0.999 | 0.11972 | -0.20966 | -0.000898 | -0.000134 | -0.1747 | -0.0898 | -0.112
freq_sweep | 4 | 2 | 17 | 3 | 120 | 1 | 0.999 | 53 | 119.9887 | 0.9999 | 99.9006 | 0.11986 | -0.0053481 | 0.11998 | 119.779 | 1 | 0.11979 | -0.20966 | 0.000102 | -6.2e-05 | -0.1747 | 0.0102 | -0.0519
freq_sweep | 4 | 2 | 18 | 3 | 120 | 1 | 0.999 | 53 | 119.9887 | 0.9999 | 99.9006 | 0.11986 | -0.0053481 | 0.11998 | 119.779 | 1 | 0.11978 | -0.20966 | 0.000102 | -8.1e-05 | -0.1747 | 0.0102 | -0.0677
freq_sweep | 4 | 3 | 1 | 1 | 120 | 1 | 0.999 | 53 | 120.0121 | 0.99997 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.739 | 2.968 | 0.35536 | -0.27314 | 1.968 | 0.23547 | -0.2276 | -1.0641 | -1.1978
freq_sweep | 4 | 3 | 2 | 1 | 120 | 1 | 0.999 | 53 | 120.0121 | 0.99997 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.739 | 1.001 | 0.11997 | -0.27314 | 0.001026 | 7.8e-05 | -0.2276 | 0.1026 | 0.0651
freq_sweep | 4 | 3 | 3 | 1 | 120 | 1 | 0.999 | 53 | 120.0121 | 0.99997 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.739 | 1.001 | 0.11998 | -0.27314 | 0.001026 | 9.4e-05 | -0.2276 | 0.1026 | 0.0784
freq_sweep | 4 | 3 | 4 | 1 | 120 | 1 | 0.999 | 53 | 120.0121 | 0.99997 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.739 | 1 | 0.11982 | -0.27314 | 2.6e-05 | -6.7e-05 | -0.2276 | 0.0026 | -0.0559
freq_sweep | 4 | 3 | 5 | 1 | 120 | 1 | 0.999 | 53 | 120.0121 | 0.99997 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.739 | 1.001 | 0.1199 | -0.27314 | 0.001026 | 1.5e-05 | -0.2276 | 0.1026 | 0.0125
freq_sweep | 4 | 3 | 6 | 1 | 120 | 1 | 0.999 | 53 | 120.0121 | 0.99997 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.739 | 0.998 | 0.11962 | -0.27314 | -0.001974 | -0.000273 | -0.2276 | -0.1974 | -0.2277
freq_sweep | 4 | 3 | 7 | 2 | 120 | 1 | 0.999 | 53 | 119.9805 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 2.979 | 0.35681 | -0.20745 | 1.9788 | 0.23692 | -0.1729 | -0.71599 | -0.79124
freq_sweep | 4 | 3 | 8 | 2 | 120 | 1 | 0.999 | 53 | 119.9805 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.11984 | -0.20745 | -0.000161 | -4.1e-05 | -0.1729 | -0.0161 | -0.0344
freq_sweep | 4 | 3 | 9 | 2 | 120 | 1 | 0.999 | 53 | 119.9805 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.11987 | -0.20745 | -0.000161 | -1.4e-05 | -0.1729 | -0.0161 | -0.0119
freq_sweep | 4 | 3 | 10 | 2 | 120 | 1 | 0.999 | 53 | 119.9805 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.11986 | -0.20745 | -0.000161 | -2.3e-05 | -0.1729 | -0.0161 | -0.0194
freq_sweep | 4 | 3 | 11 | 2 | 120 | 1 | 0.999 | 53 | 119.9805 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.11979 | -0.20745 | -0.000161 | -8.9e-05 | -0.1729 | -0.0161 | -0.0744
freq_sweep | 4 | 3 | 12 | 2 | 120 | 1 | 0.999 | 53 | 119.9805 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.1198 | -0.20745 | -0.000161 | -8.3e-05 | -0.1729 | -0.0161 | -0.0694
freq_sweep | 4 | 3 | 13 | 3 | 120 | 1 | 0.999 | 53 | 119.9877 | 0.99992 | 99.9005 | 0.11986 | -0.0053497 | 0.11998 | 119.766 | 1 | 0.1198 | -0.22169 | 7.7e-05 | -5.9e-05 | -0.1848 | 0.0077 | -0.0493
freq_sweep | 4 | 3 | 14 | 3 | 120 | 1 | 0.999 | 53 | 119.9877 | 0.99992 | 99.9005 | 0.11986 | -0.0053497 | 0.11998 | 119.766 | 0.999 | 0.11972 | -0.22169 | -0.000923 | -0.000136 | -0.1848 | -0.0923 | -0.1135
freq_sweep | 4 | 3 | 15 | 3 | 120 | 1 | 0.999 | 53 | 119.9877 | 0.99992 | 99.9005 | 0.11986 | -0.0053497 | 0.11998 | 119.766 | 1.001 | 0.1199 | -0.22169 | 0.001077 | 4.4e-05 | -0.1848 | 0.1077 | 0.0367
freq_sweep | 4 | 3 | 16 | 3 | 120 | 1 | 0.999 | 53 | 119.9877 | 0.99992 | 99.9005 | 0.11986 | -0.0053497 | 0.11998 | 119.766 | 0.999 | 0.11975 | -0.22169 | -0.000923 | -0.000111 | -0.1848 | -0.0923 | -0.0927
freq_sweep | 4 | 3 | 17 | 3 | 120 | 1 | 0.999 | 53 | 119.9877 | 0.99992 | 99.9005 | 0.11986 | -0.0053497 | 0.11998 | 119.766 | 1 | 0.11977 | -0.22169 | 7.7e-05 | -9.2e-05 | -0.1848 | 0.0077 | -0.0768
freq_sweep | 4 | 3 | 18 | 3 | 120 | 1 | 0.999 | 53 | 119.9877 | 0.99992 | 99.9005 | 0.11986 | -0.0053497 | 0.11998 | 119.766 | 1 | 0.11979 | -0.22169 | 7.7e-05 | -7.3e-05 | -0.1848 | 0.0077 | -0.061
freq_sweep | 5 | 1 | 1 | 1 | 120 | 1 | 0.999 | 54 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053406 | 0.12001 | 119.74 | 2.967 | 0.35543 | -0.27139 | 1.967 | 0.23554 | -0.2261 | -1.0981 | -1.1786
freq_sweep | 5 | 1 | 2 | 1 | 120 | 1 | 0.999 | 54 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053406 | 0.12001 | 119.74 | 1.001 | 0.11985 | -0.27139 | 0.001019 | -4.2e-05 | -0.2261 | 0.1019 | -0.0351
freq_sweep | 5 | 1 | 3 | 1 | 120 | 1 | 0.999 | 54 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053406 | 0.12001 | 119.74 | 1.001 | 0.11996 | -0.27139 | 0.001019 | 7.1e-05 | -0.2261 | 0.1019 | 0.0592
freq_sweep | 5 | 1 | 4 | 1 | 120 | 1 | 0.999 | 54 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053406 | 0.12001 | 119.74 | 1 | 0.11986 | -0.27139 | 1.9e-05 | -2.9e-05 | -0.2261 | 0.0019 | -0.0242
freq_sweep | 5 | 1 | 5 | 1 | 120 | 1 | 0.999 | 54 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053406 | 0.12001 | 119.74 | 1.001 | 0.11985 | -0.27139 | 0.001019 | -4.3e-05 | -0.2261 | 0.1019 | -0.0359
freq_sweep | 5 | 1 | 6 | 1 | 120 | 1 | 0.999 | 54 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053406 | 0.12001 | 119.74 | 0.998 | 0.11961 | -0.27139 | -0.001981 | -0.000283 | -0.2261 | -0.1981 | -0.2361
freq_sweep | 5 | 1 | 7 | 2 | 120 | 1 | 0.999 | 54 | 119.9786 | 1.0002 | 99.9038 | 0.11988 | -0.005262 | 0.12 | 119.754 | 2.978 | 0.35679 | -0.22456 | 1.9778 | 0.23691 | -0.1872 | -0.74942 | -0.79376
freq_sweep | 5 | 1 | 8 | 2 | 120 | 1 | 0.999 | 54 | 119.9786 | 1.0002 | 99.9038 | 0.11988 | -0.005262 | 0.12 | 119.754 | 1 | 0.11981 | -0.22456 | -0.000162 | -7.2e-05 | -0.1872 | -0.0162 | -0.0597
freq_sweep | 5 | 1 | 9 | 2 | 120 | 1 | 0.999 | 54 | 119.9786 | 1.0002 | 99.9038 | 0.11988 | -0.005262 | 0.12 | 119.754 | 1.001 | 0.11988 | -0.22456 | 0.000838 | -6e-06 | -0.1872 | 0.0838 | -0.0047
freq_sweep | 5 | 1 | 10 | 2 | 120 | 1 | 0.999 | 54 | 119.9786 | 1.0002 | 99.9038 | 0.11988 | -0.005262 | 0.12 | 119.754 | 1 | 0.11986 | -0.22456 | -0.000162 | -1.9e-05 | -0.1872 | -0.0162 | -0.0155
freq_sweep | 5 | 1 | 11 | 2 | 120 | 1 | 0.999 | 54 | 119.9786 | 1.0002 | 99.9038 | 0.11988 | -0.005262 | 0.12 | 119.754 | 1 | 0.11979 | -0.22456 | -0.000162 | -9.3e-05 | -0.1872 | -0.0162 | -0.0772
freq_sweep | 5 | 1 | 12 | 2 | 120 | 1 | 0.999 | 54 | 119.9786 | 1.0002 | 99.9038 | 0.11988 | -0.005262 | 0.12 | 119.754 | 1 | 0.1198 | -0.22456 | -0.000162 | -8.2e-05 | -0.1872 | -0.0162 | -0.068
freq_sweep | 5 | 1 | 13 | 3 | 120 | 1 | 0.999 | 54 | 119.9896 | 0.99991 | 99.9006 | 0.11986 | -0.0053492 | 0.11998 | 119.776 | 1 | 0.11978 | -0.21357 | 8.6e-05 | -7.9e-05 | -0.178 | 0.0086 | -0.0659
freq_sweep | 5 | 1 | 14 | 3 | 120 | 1 | 0.999 | 54 | 119.9896 | 0.99991 | 99.9006 | 0.11986 | -0.0053492 | 0.11998 | 119.776 | 0.999 | 0.11972 | -0.21357 | -0.000914 | -0.00014 | -0.178 | -0.0915 | -0.1168
freq_sweep | 5 | 1 | 15 | 3 | 120 | 1 | 0.999 | 54 | 119.9896 | 0.99991 | 99.9006 | 0.11986 | -0.0053492 | 0.11998 | 119.776 | 1 | 0.11989 | -0.21357 | 8.6e-05 | 3.3e-05 | -0.178 | 0.0086 | 0.0275
freq_sweep | 5 | 1 | 16 | 3 | 120 | 1 | 0.999 | 54 | 119.9896 | 0.99991 | 99.9006 | 0.11986 | -0.0053492 | 0.11998 | 119.776 | 0.999 | 0.11973 | -0.21357 | -0.000914 | -0.000126 | -0.178 | -0.0915 | -0.1051
freq_sweep | 5 | 1 | 17 | 3 | 120 | 1 | 0.999 | 54 | 119.9896 | 0.99991 | 99.9006 | 0.11986 | -0.0053492 | 0.11998 | 119.776 | 1 | 0.11977 | -0.21357 | 8.6e-05 | -8.9e-05 | -0.178 | 0.0086 | -0.0743
freq_sweep | 5 | 1 | 18 | 3 | 120 | 1 | 0.999 | 54 | 119.9896 | 0.99991 | 99.9006 | 0.11986 | -0.0053492 | 0.11998 | 119.776 | 1 | 0.11979 | -0.21357 | 8.6e-05 | -7e-05 | -0.178 | 0.0086 | -0.0584
freq_sweep | 5 | 2 | 1 | 1 | 120 | 1 | 0.999 | 54 | 120.0146 | 0.99995 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.743 | 2.968 | 0.35548 | -0.27162 | 1.9681 | 0.23559 | -0.2263 | -1.0614 | -1.1652
freq_sweep | 5 | 2 | 2 | 1 | 120 | 1 | 0.999 | 54 | 120.0146 | 0.99995 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.743 | 1.001 | 0.11995 | -0.27162 | 0.001053 | 6.4e-05 | -0.2263 | 0.1053 | 0.0531
freq_sweep | 5 | 2 | 3 | 1 | 120 | 1 | 0.999 | 54 | 120.0146 | 0.99995 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.743 | 1.001 | 0.11997 | -0.27162 | 0.001053 | 7.9e-05 | -0.2263 | 0.1053 | 0.0656
freq_sweep | 5 | 2 | 4 | 1 | 120 | 1 | 0.999 | 54 | 120.0146 | 0.99995 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.743 | 1 | 0.11987 | -0.27162 | 5.3e-05 | -2.2e-05 | -0.2263 | 0.0053 | -0.0186
freq_sweep | 5 | 2 | 5 | 1 | 120 | 1 | 0.999 | 54 | 120.0146 | 0.99995 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.743 | 1.001 | 0.11991 | -0.27162 | 0.001053 | 1.9e-05 | -0.2263 | 0.1053 | 0.0156
freq_sweep | 5 | 2 | 6 | 1 | 120 | 1 | 0.999 | 54 | 120.0146 | 0.99995 | 99.9009 | 0.11989 | -0.0053402 | 0.12001 | 119.743 | 0.998 | 0.11959 | -0.27162 | -0.001947 | -0.000301 | -0.2263 | -0.1947 | -0.2513
freq_sweep | 5 | 2 | 7 | 2 | 120 | 1 | 0.999 | 54 | 119.979 | 1.0002 | 99.9037 | 0.11988 | -0.0052644 | 0.12 | 119.771 | 2.978 | 0.35675 | -0.208 | 1.9778 | 0.23687 | -0.1734 | -0.74984 | -0.8064
freq_sweep | 5 | 2 | 8 | 2 | 120 | 1 | 0.999 | 54 | 119.979 | 1.0002 | 99.9037 | 0.11988 | -0.0052644 | 0.12 | 119.771 | 1 | 0.1198 | -0.208 | -0.000166 | -8e-05 | -0.1734 | -0.0166 | -0.0671
freq_sweep | 5 | 2 | 9 | 2 | 120 | 1 | 0.999 | 54 | 119.979 | 1.0002 | 99.9037 | 0.11988 | -0.0052644 | 0.12 | 119.771 | 1 | 0.11989 | -0.208 | -0.000166 | 4e-06 | -0.1734 | -0.0166 | 0.003
freq_sweep | 5 | 2 | 10 | 2 | 120 | 1 | 0.999 | 54 | 119.979 | 1.0002 | 99.9037 | 0.11988 | -0.0052644 | 0.12 | 119.771 | 1 | 0.11984 | -0.208 | -0.000166 | -4e-05 | -0.1734 | -0.0166 | -0.0337
freq_sweep | 5 | 2 | 11 | 2 | 120 | 1 | 0.999 | 54 | 119.979 | 1.0002 | 99.9037 | 0.11988 | -0.0052644 | 0.12 | 119.771 | 0.999 | 0.11976 | -0.208 | -0.001166 | -0.000126 | -0.1734 | -0.1166 | -0.1054
freq_sweep | 5 | 2 | 12 | 2 | 120 | 1 | 0.999 | 54 | 119.979 | 1.0002 | 99.9037 | 0.11988 | -0.0052644 | 0.12 | 119.771 | 1 | 0.11982 | -0.208 | -0.000166 | -6.2e-05 | -0.1734 | -0.0166 | -0.0521
freq_sweep | 5 | 2 | 13 | 3 | 120 | 1 | 0.999 | 54 | 119.9906 | 0.99991 | 99.9005 | 0.11986 | -0.0053518 | 0.11998 | 119.779 | 0.999 | 0.11977 | -0.21162 | -0.000908 | -9.2e-05 | -0.1764 | -0.0908 | -0.0769
freq_sweep | 5 | 2 | 14 | 3 | 120 | 1 | 0.999 | 54 | 119.9906 | 0.99991 | 99.9005 | 0.11986 | -0.0053518 | 0.11998 | 119.779 | 0.999 | 0.11974 | -0.21162 | -0.000908 | -0.000117 | -0.1764 | -0.0908 | -0.0978
freq_sweep | 5 | 2 | 15 | 3 | 120 | 1 | 0.999 | 54 | 119.9906 | 0.99991 | 99.9005 | 0.11986 | -0.0053518 | 0.11998 | 119.779 | 1.001 | 0.11991 | -0.21162 | 0.001092 | 4.6e-05 | -0.1764 | 0.1092 | 0.0382
freq_sweep | 5 | 2 | 16 | 3 | 120 | 1 | 0.999 | 54 | 119.9906 | 0.99991 | 99.9005 | 0.11986 | -0.0053518 | 0.11998 | 119.779 | 0.999 | 0.11973 | -0.21162 | -0.000908 | -0.000125 | -0.1764 | -0.0908 | -0.1045
freq_sweep | 5 | 2 | 17 | 3 | 120 | 1 | 0.999 | 54 | 119.9906 | 0.99991 | 99.9005 | 0.11986 | -0.0053518 | 0.11998 | 119.779 | 1 | 0.1198 | -0.21162 | 9.2e-05 | -5.9e-05 | -0.1764 | 0.0092 | -0.0494
freq_sweep | 5 | 2 | 18 | 3 | 120 | 1 | 0.999 | 54 | 119.9906 | 0.99991 | 99.9005 | 0.11986 | -0.0053518 | 0.11998 | 119.779 | 1 | 0.11978 | -0.21162 | 9.2e-05 | -7.7e-05 | -0.1764 | 0.0092 | -0.0644
freq_sweep | 5 | 3 | 1 | 1 | 120 | 1 | 0.999 | 54 | 120.0088 | 0.99997 | 99.9007 | 0.11989 | -0.0053404 | 0.12 | 119.692 | 2.968 | 0.35532 | -0.31681 | 1.968 | 0.23543 | -0.264 | -1.0633 | -1.2069
freq_sweep | 5 | 3 | 2 | 1 | 120 | 1 | 0.999 | 54 | 120.0088 | 0.99997 | 99.9007 | 0.11989 | -0.0053404 | 0.12 | 119.692 | 1.001 | 0.11987 | -0.31681 | 0.001034 | -1.1e-05 | -0.264 | 0.1034 | -0.0088
freq_sweep | 5 | 3 | 3 | 1 | 120 | 1 | 0.999 | 54 | 120.0088 | 0.99997 | 99.9007 | 0.11989 | -0.0053404 | 0.12 | 119.692 | 1.001 | 0.11989 | -0.31681 | 0.001034 | -1e-06 | -0.264 | 0.1034 | -0.0005
freq_sweep | 5 | 3 | 4 | 1 | 120 | 1 | 0.999 | 54 | 120.0088 | 0.99997 | 99.9007 | 0.11989 | -0.0053404 | 0.12 | 119.692 | 1.001 | 0.11982 | -0.31681 | 0.001034 | -6.6e-05 | -0.264 | 0.1034 | -0.0547
freq_sweep | 5 | 3 | 5 | 1 | 120 | 1 | 0.999 | 54 | 120.0088 | 0.99997 | 99.9007 | 0.11989 | -0.0053404 | 0.12 | 119.692 | 1.001 | 0.1199 | -0.31681 | 0.001034 | 1.3e-05 | -0.264 | 0.1034 | 0.0112
freq_sweep | 5 | 3 | 6 | 1 | 120 | 1 | 0.999 | 54 | 120.0088 | 0.99997 | 99.9007 | 0.11989 | -0.0053404 | 0.12 | 119.692 | 0.999 | 0.11956 | -0.31681 | -0.000966 | -0.00033 | -0.264 | -0.0966 | -0.2749
freq_sweep | 5 | 3 | 7 | 2 | 120 | 1 | 0.999 | 54 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.005264 | 0.12 | 119.757 | 2.978 | 0.35673 | -0.22261 | 1.9778 | 0.23684 | -0.1855 | -0.74919 | -0.81266
freq_sweep | 5 | 3 | 8 | 2 | 120 | 1 | 0.999 | 54 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.005264 | 0.12 | 119.757 | 1 | 0.11983 | -0.22261 | -0.00016 | -5.2e-05 | -0.1855 | -0.016 | -0.0436
freq_sweep | 5 | 3 | 9 | 2 | 120 | 1 | 0.999 | 54 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.005264 | 0.12 | 119.757 | 1.001 | 0.11991 | -0.22261 | 0.00084 | 2.4e-05 | -0.1855 | 0.084 | 0.0198
freq_sweep | 5 | 3 | 10 | 2 | 120 | 1 | 0.999 | 54 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.005264 | 0.12 | 119.757 | 1 | 0.11983 | -0.22261 | -0.00016 | -5.8e-05 | -0.1855 | -0.016 | -0.0486
freq_sweep | 5 | 3 | 11 | 2 | 120 | 1 | 0.999 | 54 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.005264 | 0.12 | 119.757 | 1 | 0.1198 | -0.22261 | -0.00016 | -8.6e-05 | -0.1855 | -0.016 | -0.0719
freq_sweep | 5 | 3 | 12 | 2 | 120 | 1 | 0.999 | 54 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.005264 | 0.12 | 119.757 | 1 | 0.11983 | -0.22261 | -0.00016 | -5.5e-05 | -0.1855 | -0.016 | -0.0461
freq_sweep | 5 | 3 | 13 | 3 | 120 | 1 | 0.999 | 54 | 119.9904 | 0.99991 | 99.9006 | 0.11986 | -0.0053479 | 0.11998 | 119.769 | 1 | 0.11979 | -0.22136 | 9.2e-05 | -7.4e-05 | -0.1845 | 0.0092 | -0.0618
freq_sweep | 5 | 3 | 14 | 3 | 120 | 1 | 0.999 | 54 | 119.9904 | 0.99991 | 99.9006 | 0.11986 | -0.0053479 | 0.11998 | 119.769 | 0.999 | 0.11976 | -0.22136 | -0.000908 | -0.000104 | -0.1845 | -0.0908 | -0.0868
freq_sweep | 5 | 3 | 15 | 3 | 120 | 1 | 0.999 | 54 | 119.9904 | 0.99991 | 99.9006 | 0.11986 | -0.0053479 | 0.11998 | 119.769 | 1 | 0.11988 | -0.22136 | 9.2e-05 | 1.8e-05 | -0.1845 | 0.0092 | 0.0149
freq_sweep | 5 | 3 | 16 | 3 | 120 | 1 | 0.999 | 54 | 119.9904 | 0.99991 | 99.9006 | 0.11986 | -0.0053479 | 0.11998 | 119.769 | 0.999 | 0.11973 | -0.22136 | -0.000908 | -0.000131 | -0.1845 | -0.0908 | -0.1094
freq_sweep | 5 | 3 | 17 | 3 | 120 | 1 | 0.999 | 54 | 119.9904 | 0.99991 | 99.9006 | 0.11986 | -0.0053479 | 0.11998 | 119.769 | 1 | 0.11981 | -0.22136 | 9.2e-05 | -5.1e-05 | -0.1845 | 0.0092 | -0.0426
freq_sweep | 5 | 3 | 18 | 3 | 120 | 1 | 0.999 | 54 | 119.9904 | 0.99991 | 99.9006 | 0.11986 | -0.0053479 | 0.11998 | 119.769 | 1 | 0.11977 | -0.22136 | 9.2e-05 | -9e-05 | -0.1845 | 0.0092 | -0.0752
freq_sweep | 6 | 1 | 1 | 1 | 120 | 1 | 0.999 | 55 | 120.012 | 0.99999 | 99.9006 | 0.11989 | -0.0053397 | 0.12001 | 119.744 | 2.968 | 0.35547 | -0.26798 | 1.968 | 0.23558 | -0.2233 | -1.0655 | -1.1683
freq_sweep | 6 | 1 | 2 | 1 | 120 | 1 | 0.999 | 55 | 120.012 | 0.99999 | 99.9006 | 0.11989 | -0.0053397 | 0.12001 | 119.744 | 1.001 | 0.11994 | -0.26798 | 0.001012 | 4.8e-05 | -0.2233 | 0.1012 | 0.0398
freq_sweep | 6 | 1 | 3 | 1 | 120 | 1 | 0.999 | 55 | 120.012 | 0.99999 | 99.9006 | 0.11989 | -0.0053397 | 0.12001 | 119.744 | 1.001 | 0.11999 | -0.26798 | 0.001012 | 9.5e-05 | -0.2233 | 0.1012 | 0.079
freq_sweep | 6 | 1 | 4 | 1 | 120 | 1 | 0.999 | 55 | 120.012 | 0.99999 | 99.9006 | 0.11989 | -0.0053397 | 0.12001 | 119.744 | 1 | 0.11986 | -0.26798 | 1.2e-05 | -2.9e-05 | -0.2233 | 0.0012 | -0.0245
freq_sweep | 6 | 1 | 5 | 1 | 120 | 1 | 0.999 | 55 | 120.012 | 0.99999 | 99.9006 | 0.11989 | -0.0053397 | 0.12001 | 119.744 | 1.001 | 0.11995 | -0.26798 | 0.001012 | 5.9e-05 | -0.2233 | 0.1012 | 0.0489
freq_sweep | 6 | 1 | 6 | 1 | 120 | 1 | 0.999 | 55 | 120.012 | 0.99999 | 99.9006 | 0.11989 | -0.0053397 | 0.12001 | 119.744 | 0.999 | 0.11965 | -0.26798 | -0.000988 | -0.000245 | -0.2233 | -0.0988 | -0.2046
freq_sweep | 6 | 1 | 7 | 2 | 120 | 1 | 0.999 | 55 | 119.9791 | 1.0002 | 99.9038 | 0.11988 | -0.0052599 | 0.12 | 119.76 | 2.978 | 0.35673 | -0.21911 | 1.9778 | 0.23685 | -0.1826 | -0.74844 | -0.81079
freq_sweep | 6 | 1 | 8 | 2 | 120 | 1 | 0.999 | 55 | 119.9791 | 1.0002 | 99.9038 | 0.11988 | -0.0052599 | 0.12 | 119.76 | 1 | 0.1198 | -0.21911 | -0.000152 | -8.2e-05 | -0.1826 | -0.0152 | -0.0684
freq_sweep | 6 | 1 | 9 | 2 | 120 | 1 | 0.999 | 55 | 119.9791 | 1.0002 | 99.9038 | 0.11988 | -0.0052599 | 0.12 | 119.76 | 1.001 | 0.11989 | -0.21911 | 0.000848 | 6e-06 | -0.1826 | 0.0848 | 0.005
freq_sweep | 6 | 1 | 10 | 2 | 120 | 1 | 0.999 | 55 | 119.9791 | 1.0002 | 99.9038 | 0.11988 | -0.0052599 | 0.12 | 119.76 | 1 | 0.11984 | -0.21911 | -0.000152 | -4.5e-05 | -0.1826 | -0.0152 | -0.0375
freq_sweep | 6 | 1 | 11 | 2 | 120 | 1 | 0.999 | 55 | 119.9791 | 1.0002 | 99.9038 | 0.11988 | -0.0052599 | 0.12 | 119.76 | 1 | 0.11977 | -0.21911 | -0.000152 | -0.000111 | -0.1826 | -0.0152 | -0.0926
freq_sweep | 6 | 1 | 12 | 2 | 120 | 1 | 0.999 | 55 | 119.9791 | 1.0002 | 99.9038 | 0.11988 | -0.0052599 | 0.12 | 119.76 | 1 | 0.11981 | -0.21911 | -0.000152 | -7.1e-05 | -0.1826 | -0.0152 | -0.0592
freq_sweep | 6 | 1 | 13 | 3 | 120 | 1 | 0.999 | 55 | 119.9895 | 0.99991 | 99.9005 | 0.11986 | -0.0053503 | 0.11998 | 119.783 | 1 | 0.11981 | -0.20651 | 9.4e-05 | -5.1e-05 | -0.1721 | 0.0094 | -0.0424
freq_sweep | 6 | 1 | 14 | 3 | 120 | 1 | 0.999 | 55 | 119.9895 | 0.99991 | 99.9005 | 0.11986 | -0.0053503 | 0.11998 | 119.783 | 0.999 | 0.11976 | -0.20651 | -0.000906 | -0.000103 | -0.1721 | -0.0906 | -0.0858
freq_sweep | 6 | 1 | 15 | 3 | 120 | 1 | 0.999 | 55 | 119.9895 | 0.99991 | 99.9005 | 0.11986 | -0.0053503 | 0.11998 | 119.783 | 1 | 0.11989 | -0.20651 | 9.4e-05 | 3e-05 | -0.1721 | 0.0094 | 0.0252
freq_sweep | 6 | 1 | 16 | 3 | 120 | 1 | 0.999 | 55 | 119.9895 | 0.99991 | 99.9005 | 0.11986 | -0.0053503 | 0.11998 | 119.783 | 0.999 | 0.11977 | -0.20651 | -0.000906 | -8.9e-05 | -0.1721 | -0.0906 | -0.0741
freq_sweep | 6 | 1 | 17 | 3 | 120 | 1 | 0.999 | 55 | 119.9895 | 0.99991 | 99.9005 | 0.11986 | -0.0053503 | 0.11998 | 119.783 | 1 | 0.1198 | -0.20651 | 9.4e-05 | -5.8e-05 | -0.1721 | 0.0094 | -0.0483
freq_sweep | 6 | 1 | 18 | 3 | 120 | 1 | 0.999 | 55 | 119.9895 | 0.99991 | 99.9005 | 0.11986 | -0.0053503 | 0.11998 | 119.783 | 1 | 0.11977 | -0.20651 | 9.4e-05 | -8.7e-05 | -0.1721 | 0.0094 | -0.0725
freq_sweep | 6 | 2 | 1 | 1 | 120 | 1 | 0.999 | 55 | 120.014 | 0.99997 | 99.9007 | 0.11989 | -0.0053446 | 0.12001 | 119.733 | 2.968 | 0.35526 | -0.28102 | 1.968 | 0.23537 | -0.2342 | -1.064 | -1.2272
freq_sweep | 6 | 2 | 2 | 1 | 120 | 1 | 0.999 | 55 | 120.014 | 0.99997 | 99.9007 | 0.11989 | -0.0053446 | 0.12001 | 119.733 | 1.001 | 0.11994 | -0.28102 | 0.001027 | 5e-05 | -0.2342 | 0.1027 | 0.042
freq_sweep | 6 | 2 | 3 | 1 | 120 | 1 | 0.999 | 55 | 120.014 | 0.99997 | 99.9007 | 0.11989 | -0.0053446 | 0.12001 | 119.733 | 1.001 | 0.11994 | -0.28102 | 0.001027 | 4.8e-05 | -0.2342 | 0.1027 | 0.0403
freq_sweep | 6 | 2 | 4 | 1 | 120 | 1 | 0.999 | 55 | 120.014 | 0.99997 | 99.9007 | 0.11989 | -0.0053446 | 0.12001 | 119.733 | 1 | 0.11978 | -0.28102 | 2.7e-05 | -0.00011 | -0.2342 | 0.0027 | -0.0915
freq_sweep | 6 | 2 | 5 | 1 | 120 | 1 | 0.999 | 55 | 120.014 | 0.99997 | 99.9007 | 0.11989 | -0.0053446 | 0.12001 | 119.733 | 1.001 | 0.11992 | -0.28102 | 0.001027 | 3e-05 | -0.2342 | 0.1027 | 0.0253
freq_sweep | 6 | 2 | 6 | 1 | 120 | 1 | 0.999 | 55 | 120.014 | 0.99997 | 99.9007 | 0.11989 | -0.0053446 | 0.12001 | 119.733 | 0.999 | 0.11962 | -0.28102 | -0.000973 | -0.000268 | -0.2342 | -0.0973 | -0.2233
freq_sweep | 6 | 2 | 7 | 2 | 120 | 1 | 0.999 | 55 | 119.9851 | 1.0002 | 99.9039 | 0.11989 | -0.0052607 | 0.12 | 119.77 | 2.978 | 0.35677 | -0.21505 | 1.9778 | 0.23689 | -0.1792 | -0.7489 | -0.80395
freq_sweep | 6 | 2 | 8 | 2 | 120 | 1 | 0.999 | 55 | 119.9851 | 1.0002 | 99.9039 | 0.11989 | -0.0052607 | 0.12 | 119.77 | 1 | 0.11983 | -0.21505 | -0.000157 | -6.2e-05 | -0.1792 | -0.0157 | -0.0513
freq_sweep | 6 | 2 | 9 | 2 | 120 | 1 | 0.999 | 55 | 119.9851 | 1.0002 | 99.9039 | 0.11989 | -0.0052607 | 0.12 | 119.77 | 1.001 | 0.11988 | -0.21505 | 0.000843 | -9e-06 | -0.1792 | 0.0843 | -0.0071
freq_sweep | 6 | 2 | 10 | 2 | 120 | 1 | 0.999 | 55 | 119.9851 | 1.0002 | 99.9039 | 0.11989 | -0.0052607 | 0.12 | 119.77 | 1 | 0.11985 | -0.21505 | -0.000157 | -3.7e-05 | -0.1792 | -0.0157 | -0.0305
freq_sweep | 6 | 2 | 11 | 2 | 120 | 1 | 0.999 | 55 | 119.9851 | 1.0002 | 99.9039 | 0.11989 | -0.0052607 | 0.12 | 119.77 | 1 | 0.11979 | -0.21505 | -0.000157 | -0.000103 | -0.1792 | -0.0157 | -0.0855
freq_sweep | 6 | 2 | 12 | 2 | 120 | 1 | 0.999 | 55 | 119.9851 | 1.0002 | 99.9039 | 0.11989 | -0.0052607 | 0.12 | 119.77 | 1 | 0.11981 | -0.21505 | -0.000157 | -8.2e-05 | -0.1792 | -0.0157 | -0.068
freq_sweep | 6 | 2 | 13 | 3 | 120 | 1 | 0.999 | 55 | 119.9879 | 0.99991 | 99.9006 | 0.11986 | -0.0053496 | 0.11998 | 119.765 | 1 | 0.11979 | -0.22287 | 9.2e-05 | -6.5e-05 | -0.1857 | 0.0092 | -0.0538
freq_sweep | 6 | 2 | 14 | 3 | 120 | 1 | 0.999 | 55 | 119.9879 | 0.99991 | 99.9006 | 0.11986 | -0.0053496 | 0.11998 | 119.765 | 0.999 | 0.11971 | -0.22287 | -0.000908 | -0.000152 | -0.1857 | -0.0908 | -0.1264
freq_sweep | 6 | 2 | 15 | 3 | 120 | 1 | 0.999 | 55 | 119.9879 | 0.99991 | 99.9006 | 0.11986 | -0.0053496 | 0.11998 | 119.765 | 1.001 | 0.11991 | -0.22287 | 0.001092 | 5.1e-05 | -0.1857 | 0.1092 | 0.043
freq_sweep | 6 | 2 | 16 | 3 | 120 | 1 | 0.999 | 55 | 119.9879 | 0.99991 | 99.9006 | 0.11986 | -0.0053496 | 0.11998 | 119.765 | 0.999 | 0.11975 | -0.22287 | -0.000908 | -0.000109 | -0.1857 | -0.0908 | -0.0905
freq_sweep | 6 | 2 | 17 | 3 | 120 | 1 | 0.999 | 55 | 119.9879 | 0.99991 | 99.9006 | 0.11986 | -0.0053496 | 0.11998 | 119.765 | 1 | 0.11978 | -0.22287 | 9.2e-05 | -8.2e-05 | -0.1857 | 0.0092 | -0.068
freq_sweep | 6 | 2 | 18 | 3 | 120 | 1 | 0.999 | 55 | 119.9879 | 0.99991 | 99.9006 | 0.11986 | -0.0053496 | 0.11998 | 119.765 | 1 | 0.11982 | -0.22287 | 9.2e-05 | -4.2e-05 | -0.1857 | 0.0092 | -0.0346
freq_sweep | 6 | 3 | 1 | 1 | 120 | 1 | 0.999 | 55 | 120.0134 | 0.99999 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.693 | 2.968 | 0.35555 | -0.32039 | 1.968 | 0.23566 | -0.267 | -1.0657 | -1.1484
freq_sweep | 6 | 3 | 2 | 1 | 120 | 1 | 0.999 | 55 | 120.0134 | 0.99999 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.693 | 1.001 | 0.11992 | -0.32039 | 0.00101 | 2.4e-05 | -0.267 | 0.101 | 0.0201
freq_sweep | 6 | 3 | 3 | 1 | 120 | 1 | 0.999 | 55 | 120.0134 | 0.99999 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.693 | 1.001 | 0.11991 | -0.32039 | 0.00101 | 1.6e-05 | -0.267 | 0.101 | 0.0135
freq_sweep | 6 | 3 | 4 | 1 | 120 | 1 | 0.999 | 55 | 120.0134 | 0.99999 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.693 | 1 | 0.11988 | -0.32039 | 1e-05 | -9e-06 | -0.267 | 0.001 | -0.0074
freq_sweep | 6 | 3 | 5 | 1 | 120 | 1 | 0.999 | 55 | 120.0134 | 0.99999 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.693 | 1.001 | 0.1199 | -0.32039 | 0.00101 | 8e-06 | -0.267 | 0.101 | 0.0068
freq_sweep | 6 | 3 | 6 | 1 | 120 | 1 | 0.999 | 55 | 120.0134 | 0.99999 | 99.9006 | 0.11989 | -0.0053414 | 0.12001 | 119.693 | 0.998 | 0.11954 | -0.32039 | -0.00199 | -0.000352 | -0.267 | -0.199 | -0.2935
freq_sweep | 6 | 3 | 7 | 2 | 120 | 1 | 0.999 | 55 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052649 | 0.12 | 119.77 | 2.978 | 0.35668 | -0.21375 | 1.9778 | 0.2368 | -0.1781 | -0.74875 | -0.82787
freq_sweep | 6 | 3 | 8 | 2 | 120 | 1 | 0.999 | 55 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052649 | 0.12 | 119.77 | 1 | 0.1198 | -0.21375 | -0.000155 | -8.4e-05 | -0.1781 | -0.0155 | -0.0699
freq_sweep | 6 | 3 | 9 | 2 | 120 | 1 | 0.999 | 55 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052649 | 0.12 | 119.77 | 1.001 | 0.11991 | -0.21375 | 0.000845 | 2e-05 | -0.1781 | 0.0845 | 0.0168
freq_sweep | 6 | 3 | 10 | 2 | 120 | 1 | 0.999 | 55 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052649 | 0.12 | 119.77 | 1 | 0.11983 | -0.21375 | -0.000155 | -5.4e-05 | -0.1781 | -0.0155 | -0.0449
freq_sweep | 6 | 3 | 11 | 2 | 120 | 1 | 0.999 | 55 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052649 | 0.12 | 119.77 | 1 | 0.11977 | -0.21375 | -0.000155 | -0.000118 | -0.1781 | -0.0155 | -0.0983
freq_sweep | 6 | 3 | 12 | 2 | 120 | 1 | 0.999 | 55 | 119.9837 | 1.0002 | 99.9037 | 0.11989 | -0.0052649 | 0.12 | 119.77 | 1 | 0.11983 | -0.21375 | -0.000155 | -5.8e-05 | -0.1781 | -0.0155 | -0.0482
freq_sweep | 6 | 3 | 13 | 3 | 120 | 1 | 0.999 | 55 | 119.9906 | 0.99992 | 99.9006 | 0.11986 | -0.0053467 | 0.11998 | 119.779 | 1 | 0.11979 | -0.2116 | 8.4e-05 | -7.1e-05 | -0.1763 | 0.0084 | -0.0595
freq_sweep | 6 | 3 | 14 | 3 | 120 | 1 | 0.999 | 55 | 119.9906 | 0.99992 | 99.9006 | 0.11986 | -0.0053467 | 0.11998 | 119.779 | 0.999 | 0.11975 | -0.2116 | -0.000916 | -0.000113 | -0.1763 | -0.0916 | -0.0945
freq_sweep | 6 | 3 | 15 | 3 | 120 | 1 | 0.999 | 55 | 119.9906 | 0.99992 | 99.9006 | 0.11986 | -0.0053467 | 0.11998 | 119.779 | 1 | 0.11989 | -0.2116 | 8.4e-05 | 2.6e-05 | -0.1763 | 0.0084 | 0.0214
freq_sweep | 6 | 3 | 16 | 3 | 120 | 1 | 0.999 | 55 | 119.9906 | 0.99992 | 99.9006 | 0.11986 | -0.0053467 | 0.11998 | 119.779 | 0.999 | 0.11974 | -0.2116 | -0.000916 | -0.000119 | -0.1763 | -0.0916 | -0.0996
freq_sweep | 6 | 3 | 17 | 3 | 120 | 1 | 0.999 | 55 | 119.9906 | 0.99992 | 99.9006 | 0.11986 | -0.0053467 | 0.11998 | 119.779 | 1 | 0.11982 | -0.2116 | 8.4e-05 | -4.4e-05 | -0.1763 | 0.0084 | -0.037
freq_sweep | 6 | 3 | 18 | 3 | 120 | 1 | 0.999 | 55 | 119.9906 | 0.99992 | 99.9006 | 0.11986 | -0.0053467 | 0.11998 | 119.779 | 1 | 0.1198 | -0.2116 | 8.4e-05 | -6.5e-05 | -0.1763 | 0.0084 | -0.0545
freq_sweep | 7 | 1 | 1 | 1 | 120 | 1 | 0.999 | 56 | 120.0301 | 0.99996 | 99.9005 | 0.11991 | -0.0053402 | 0.12003 | 119.791 | 2.968 | 0.35557 | -0.23907 | 1.968 | 0.23567 | -0.1992 | -1.0625 | -1.1514
freq_sweep | 7 | 1 | 2 | 1 | 120 | 1 | 0.999 | 56 | 120.0301 | 0.99996 | 99.9005 | 0.11991 | -0.0053402 | 0.12003 | 119.791 | 1.001 | 0.11995 | -0.23907 | 0.001042 | 4.8e-05 | -0.1992 | 0.1042 | 0.0404
freq_sweep | 7 | 1 | 3 | 1 | 120 | 1 | 0.999 | 56 | 120.0301 | 0.99996 | 99.9005 | 0.11991 | -0.0053402 | 0.12003 | 119.791 | 1.002 | 0.12 | -0.23907 | 0.002042 | 9.1e-05 | -0.1992 | 0.2042 | 0.0762
freq_sweep | 7 | 1 | 4 | 1 | 120 | 1 | 0.999 | 56 | 120.0301 | 0.99996 | 99.9005 | 0.11991 | -0.0053402 | 0.12003 | 119.791 | 1 | 0.11989 | -0.23907 | 4.2e-05 | -1.4e-05 | -0.1992 | 0.0042 | -0.0113
freq_sweep | 7 | 1 | 5 | 1 | 120 | 1 | 0.999 | 56 | 120.0301 | 0.99996 | 99.9005 | 0.11991 | -0.0053402 | 0.12003 | 119.791 | 1.001 | 0.11995 | -0.23907 | 0.001042 | 4.9e-05 | -0.1992 | 0.1042 | 0.0412
freq_sweep | 7 | 1 | 6 | 1 | 120 | 1 | 0.999 | 56 | 120.0301 | 0.99996 | 99.9005 | 0.11991 | -0.0053402 | 0.12003 | 119.791 | 0.998 | 0.11961 | -0.23907 | -0.001958 | -0.000296 | -0.1992 | -0.1958 | -0.2465
freq_sweep | 7 | 1 | 7 | 2 | 120 | 1 | 0.999 | 56 | 119.9829 | 1.0002 | 99.9036 | 0.11989 | -0.0052674 | 0.12 | 119.763 | 2.978 | 0.35673 | -0.21992 | 1.9778 | 0.23684 | -0.1833 | -0.74977 | -0.81557
freq_sweep | 7 | 1 | 8 | 2 | 120 | 1 | 0.999 | 56 | 119.9829 | 1.0002 | 99.9036 | 0.11989 | -0.0052674 | 0.12 | 119.763 | 1 | 0.11983 | -0.21992 | -0.000166 | -6.1e-05 | -0.1833 | -0.0166 | -0.051
freq_sweep | 7 | 1 | 9 | 2 | 120 | 1 | 0.999 | 56 | 119.9829 | 1.0002 | 99.9036 | 0.11989 | -0.0052674 | 0.12 | 119.763 | 1 | 0.11987 | -0.21992 | -0.000166 | -1.7e-05 | -0.1833 | -0.0166 | -0.0143
freq_sweep | 7 | 1 | 10 | 2 | 120 | 1 | 0.999 | 56 | 119.9829 | 1.0002 | 99.9036 | 0.11989 | -0.0052674 | 0.12 | 119.763 | 1 | 0.11983 | -0.21992 | -0.000166 | -5.6e-05 | -0.1833 | -0.0166 | -0.0468
freq_sweep | 7 | 1 | 11 | 2 | 120 | 1 | 0.999 | 56 | 119.9829 | 1.0002 | 99.9036 | 0.11989 | -0.0052674 | 0.12 | 119.763 | 1 | 0.11979 | -0.21992 | -0.000166 | -9.7e-05 | -0.1833 | -0.0166 | -0.081
freq_sweep | 7 | 1 | 12 | 2 | 120 | 1 | 0.999 | 56 | 119.9829 | 1.0002 | 99.9036 | 0.11989 | -0.0052674 | 0.12 | 119.763 | 1 | 0.1198 | -0.21992 | -0.000166 | -9.1e-05 | -0.1833 | -0.0166 | -0.076
freq_sweep | 7 | 1 | 13 | 3 | 120 | 1 | 0.999 | 56 | 119.9914 | 0.99993 | 99.9006 | 0.11986 | -0.0053484 | 0.11998 | 119.784 | 1 | 0.11981 | -0.20742 | 7.3e-05 | -5.4e-05 | -0.1729 | 0.0073 | -0.0454
freq_sweep | 7 | 1 | 14 | 3 | 120 | 1 | 0.999 | 56 | 119.9914 | 0.99993 | 99.9006 | 0.11986 | -0.0053484 | 0.11998 | 119.784 | 0.999 | 0.11972 | -0.20742 | -0.000927 | -0.000148 | -0.1729 | -0.0927 | -0.1238
freq_sweep | 7 | 1 | 15 | 3 | 120 | 1 | 0.999 | 56 | 119.9914 | 0.99993 | 99.9006 | 0.11986 | -0.0053484 | 0.11998 | 119.784 | 1 | 0.11988 | -0.20742 | 7.3e-05 | 2e-05 | -0.1729 | 0.0073 | 0.0163
freq_sweep | 7 | 1 | 16 | 3 | 120 | 1 | 0.999 | 56 | 119.9914 | 0.99993 | 99.9006 | 0.11986 | -0.0053484 | 0.11998 | 119.784 | 0.999 | 0.11976 | -0.20742 | -0.000927 | -0.000104 | -0.1729 | -0.0927 | -0.0871
freq_sweep | 7 | 1 | 17 | 3 | 120 | 1 | 0.999 | 56 | 119.9914 | 0.99993 | 99.9006 | 0.11986 | -0.0053484 | 0.11998 | 119.784 | 1 | 0.11978 | -0.20742 | 7.3e-05 | -8.5e-05 | -0.1729 | 0.0073 | -0.0713
freq_sweep | 7 | 1 | 18 | 3 | 120 | 1 | 0.999 | 56 | 119.9914 | 0.99993 | 99.9006 | 0.11986 | -0.0053484 | 0.11998 | 119.784 | 1 | 0.11978 | -0.20742 | 7.3e-05 | -8.1e-05 | -0.1729 | 0.0073 | -0.0679
freq_sweep | 7 | 2 | 1 | 1 | 120 | 1 | 0.999 | 56 | 120.0126 | 0.99998 | 99.9009 | 0.11989 | -0.0053397 | 0.12001 | 119.753 | 2.968 | 0.35545 | -0.25963 | 1.968 | 0.23556 | -0.2163 | -1.0649 | -1.1746
freq_sweep | 7 | 2 | 2 | 1 | 120 | 1 | 0.999 | 56 | 120.0126 | 0.99998 | 99.9009 | 0.11989 | -0.0053397 | 0.12001 | 119.753 | 1.001 | 0.11995 | -0.25963 | 0.001018 | 6.3e-05 | -0.2163 | 0.1018 | 0.0529
freq_sweep | 7 | 2 | 3 | 1 | 120 | 1 | 0.999 | 56 | 120.0126 | 0.99998 | 99.9009 | 0.11989 | -0.0053397 | 0.12001 | 119.753 | 1.001 | 0.11999 | -0.25963 | 0.001018 | 9.7e-05 | -0.2163 | 0.1018 | 0.0813
freq_sweep | 7 | 2 | 4 | 1 | 120 | 1 | 0.999 | 56 | 120.0126 | 0.99998 | 99.9009 | 0.11989 | -0.0053397 | 0.12001 | 119.753 | 1 | 0.11986 | -0.25963 | 1.8e-05 | -3.5e-05 | -0.2163 | 0.0018 | -0.0288
freq_sweep | 7 | 2 | 5 | 1 | 120 | 1 | 0.999 | 56 | 120.0126 | 0.99998 | 99.9009 | 0.11989 | -0.0053397 | 0.12001 | 119.753 | 1.001 | 0.11994 | -0.25963 | 0.001018 | 4.8e-05 | -0.2163 | 0.1018 | 0.0404
freq_sweep | 7 | 2 | 6 | 1 | 120 | 1 | 0.999 | 56 | 120.0126 | 0.99998 | 99.9009 | 0.11989 | -0.0053397 | 0.12001 | 119.741 | 0.998 | 0.11962 | -0.27163 | -0.001982 | -0.000272 | -0.2263 | -0.1982 | -0.2265
freq_sweep | 7 | 2 | 7 | 2 | 120 | 1 | 0.999 | 56 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.0052645 | 0.12 | 119.774 | 2.979 | 0.35683 | -0.20561 | 1.9788 | 0.23695 | -0.1714 | -0.7161 | -0.78342
freq_sweep | 7 | 2 | 8 | 2 | 120 | 1 | 0.999 | 56 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.0052645 | 0.12 | 119.774 | 1 | 0.11979 | -0.20561 | -0.000162 | -9.1e-05 | -0.1714 | -0.0162 | -0.0755
freq_sweep | 7 | 2 | 9 | 2 | 120 | 1 | 0.999 | 56 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.0052645 | 0.12 | 119.774 | 1 | 0.11987 | -0.20561 | -0.000162 | -1e-05 | -0.1714 | -0.0162 | -0.0079
freq_sweep | 7 | 2 | 10 | 2 | 120 | 1 | 0.999 | 56 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.0052645 | 0.12 | 119.774 | 1 | 0.11986 | -0.20561 | -0.000162 | -2e-05 | -0.1714 | -0.0162 | -0.0163
freq_sweep | 7 | 2 | 11 | 2 | 120 | 1 | 0.999 | 56 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.0052645 | 0.12 | 119.774 | 1 | 0.11977 | -0.20561 | -0.000162 | -0.000116 | -0.1714 | -0.0162 | -0.0964
freq_sweep | 7 | 2 | 12 | 2 | 120 | 1 | 0.999 | 56 | 119.9796 | 1.0002 | 99.9037 | 0.11988 | -0.0052645 | 0.12 | 119.768 | 1 | 0.1198 | -0.21161 | -0.000162 | -8.3e-05 | -0.1764 | -0.0162 | -0.0688
freq_sweep | 7 | 2 | 13 | 3 | 120 | 1 | 0.999 | 56 | 119.992 | 0.99991 | 99.9005 | 0.11986 | -0.0053506 | 0.11998 | 119.773 | 1 | 0.11979 | -0.21902 | 8.8e-05 | -7.4e-05 | -0.1825 | 0.0088 | -0.0618
freq_sweep | 7 | 2 | 14 | 3 | 120 | 1 | 0.999 | 56 | 119.992 | 0.99991 | 99.9005 | 0.11986 | -0.0053506 | 0.11998 | 119.773 | 0.999 | 0.11974 | -0.21902 | -0.000912 | -0.000125 | -0.1825 | -0.0912 | -0.1044
freq_sweep | 7 | 2 | 15 | 3 | 120 | 1 | 0.999 | 56 | 119.992 | 0.99991 | 99.9005 | 0.11986 | -0.0053506 | 0.11998 | 119.773 | 1 | 0.1199 | -0.21902 | 8.8e-05 | 3.9e-05 | -0.1825 | 0.0088 | 0.0324
freq_sweep | 7 | 2 | 16 | 3 | 120 | 1 | 0.999 | 56 | 119.992 | 0.99991 | 99.9005 | 0.11986 | -0.0053506 | 0.11998 | 119.773 | 0.999 | 0.11975 | -0.21902 | -0.000912 | -0.000114 | -0.1825 | -0.0912 | -0.0952
freq_sweep | 7 | 2 | 17 | 3 | 120 | 1 | 0.999 | 56 | 119.992 | 0.99991 | 99.9005 | 0.11986 | -0.0053506 | 0.11998 | 119.764 | 1 | 0.11981 | -0.22802 | 8.8e-05 | -5.6e-05 | -0.19 | 0.0088 | -0.0468
freq_sweep | 7 | 2 | 18 | 3 | 120 | 1 | 0.999 | 56 | 119.992 | 0.99991 | 99.9005 | 0.11986 | -0.0053506 | 0.11998 | 119.764 | 1 | 0.11981 | -0.22802 | 8.8e-05 | -5.1e-05 | -0.19 | 0.0088 | -0.0427
freq_sweep | 7 | 3 | 1 | 1 | 120 | 1 | 0.999 | 56 | 120.0113 | 0.99997 | 99.9008 | 0.11989 | -0.0053423 | 0.12001 | 119.722 | 2.968 | 0.35541 | -0.28927 | 1.968 | 0.23553 | -0.241 | -1.0632 | -1.1819
freq_sweep | 7 | 3 | 2 | 1 | 120 | 1 | 0.999 | 56 | 120.0113 | 0.99997 | 99.9008 | 0.11989 | -0.0053423 | 0.12001 | 119.722 | 1.001 | 0.1199 | -0.28927 | 0.001035 | 1.6e-05 | -0.241 | 0.1035 | 0.0134
freq_sweep | 7 | 3 | 3 | 1 | 120 | 1 | 0.999 | 56 | 120.0113 | 0.99997 | 99.9008 | 0.11989 | -0.0053423 | 0.12001 | 119.722 | 1.001 | 0.11996 | -0.28927 | 0.001035 | 7.2e-05 | -0.241 | 0.1035 | 0.0601
freq_sweep | 7 | 3 | 4 | 1 | 120 | 1 | 0.999 | 56 | 120.0113 | 0.99997 | 99.9008 | 0.11989 | -0.0053423 | 0.12001 | 119.722 | 1.001 | 0.11986 | -0.28927 | 0.001035 | -2.6e-05 | -0.241 | 0.1035 | -0.0217
freq_sweep | 7 | 3 | 5 | 1 | 120 | 1 | 0.999 | 56 | 120.0113 | 0.99997 | 99.9008 | 0.11989 | -0.0053423 | 0.12001 | 119.722 | 1.001 | 0.11992 | -0.28927 | 0.001035 | 3e-05 | -0.241 | 0.1035 | 0.025
freq_sweep | 7 | 3 | 6 | 1 | 120 | 1 | 0.999 | 56 | 120.0113 | 0.99997 | 99.9008 | 0.11989 | -0.0053423 | 0.12001 | 119.722 | 0.999 | 0.11962 | -0.28927 | -0.000965 | -0.000266 | -0.241 | -0.0965 | -0.2219
freq_sweep | 7 | 3 | 7 | 2 | 120 | 1 | 0.999 | 56 | 119.9827 | 1.0002 | 99.9039 | 0.11989 | -0.0052599 | 0.12 | 119.765 | 2.978 | 0.35672 | -0.21766 | 1.9778 | 0.23683 | -0.1814 | -0.75007 | -0.81873
freq_sweep | 7 | 3 | 8 | 2 | 120 | 1 | 0.999 | 56 | 119.9827 | 1.0002 | 99.9039 | 0.11989 | -0.0052599 | 0.12 | 119.765 | 1 | 0.1198 | -0.21766 | -0.000169 | -8.5e-05 | -0.1814 | -0.0169 | -0.0705
freq_sweep | 7 | 3 | 9 | 2 | 120 | 1 | 0.999 | 56 | 119.9827 | 1.0002 | 99.9039 | 0.11989 | -0.0052599 | 0.12 | 119.765 | 1.001 | 0.11989 | -0.21766 | 0.000831 | 6e-06 | -0.1814 | 0.0831 | 0.0054
freq_sweep | 7 | 3 | 10 | 2 | 120 | 1 | 0.999 | 56 | 119.9827 | 1.0002 | 99.9039 | 0.11989 | -0.0052599 | 0.12 | 119.765 | 1 | 0.11982 | -0.21766 | -0.000169 | -6.9e-05 | -0.1814 | -0.0169 | -0.0572
freq_sweep | 7 | 3 | 11 | 2 | 120 | 1 | 0.999 | 56 | 119.9827 | 1.0002 | 99.9039 | 0.11989 | -0.0052599 | 0.12 | 119.765 | 1 | 0.11979 | -0.21766 | -0.000169 | -0.000101 | -0.1814 | -0.0169 | -0.0839
freq_sweep | 7 | 3 | 12 | 2 | 120 | 1 | 0.999 | 56 | 119.9827 | 1.0002 | 99.9039 | 0.11989 | -0.0052599 | 0.12 | 119.765 | 1 | 0.11982 | -0.21766 | -0.000169 | -6.6e-05 | -0.1814 | -0.0169 | -0.0547
freq_sweep | 7 | 3 | 13 | 3 | 120 | 1 | 0.999 | 56 | 119.9894 | 0.99992 | 99.9006 | 0.11986 | -0.0053487 | 0.11998 | 119.784 | 1 | 0.11979 | -0.20543 | 7.5e-05 | -7.4e-05 | -0.1712 | 0.0075 | -0.0618
freq_sweep | 7 | 3 | 14 | 3 | 120 | 1 | 0.999 | 56 | 119.9894 | 0.99992 | 99.9006 | 0.11986 | -0.0053487 | 0.11998 | 119.784 | 0.999 | 0.11975 | -0.20543 | -0.000925 | -0.000107 | -0.1712 | -0.0925 | -0.0894
freq_sweep | 7 | 3 | 15 | 3 | 120 | 1 | 0.999 | 56 | 119.9894 | 0.99992 | 99.9006 | 0.11986 | -0.0053487 | 0.11998 | 119.784 | 1 | 0.11987 | -0.20543 | 7.5e-05 | 1.2e-05 | -0.1712 | 0.0075 | 0.0099
freq_sweep | 7 | 3 | 16 | 3 | 120 | 1 | 0.999 | 56 | 119.9894 | 0.99992 | 99.9006 | 0.11986 | -0.0053487 | 0.11998 | 119.784 | 0.999 | 0.11976 | -0.20543 | -0.000925 | -0.000105 | -0.1712 | -0.0925 | -0.0877
freq_sweep | 7 | 3 | 17 | 3 | 120 | 1 | 0.999 | 56 | 119.9894 | 0.99992 | 99.9006 | 0.11986 | -0.0053487 | 0.11998 | 119.784 | 1 | 0.1198 | -0.20543 | 7.5e-05 | -5.7e-05 | -0.1712 | 0.0075 | -0.0477
freq_sweep | 7 | 3 | 18 | 3 | 120 | 1 | 0.999 | 56 | 119.9894 | 0.99992 | 99.9006 | 0.11986 | -0.0053487 | 0.11998 | 119.784 | 1 | 0.11976 | -0.20543 | 7.5e-05 | -9.6e-05 | -0.1712 | 0.0075 | -0.0802
freq_sweep | 8 | 1 | 1 | 1 | 120 | 1 | 0.999 | 57 | 120.0235 | 0.99996 | 99.901 | 0.1199 | -0.0053373 | 0.12002 | 119.713 | 2.968 | 0.35533 | -0.31051 | 1.968 | 0.23542 | -0.2587 | -1.0626 | -1.2161
freq_sweep | 8 | 1 | 2 | 1 | 120 | 1 | 0.999 | 57 | 120.0235 | 0.99996 | 99.901 | 0.1199 | -0.0053373 | 0.12002 | 119.713 | 1.001 | 0.11994 | -0.31051 | 0.001041 | 4.1e-05 | -0.2587 | 0.1041 | 0.0344
freq_sweep | 8 | 1 | 3 | 1 | 120 | 1 | 0.999 | 57 | 120.0235 | 0.99996 | 99.901 | 0.1199 | -0.0053373 | 0.12002 | 119.713 | 1.001 | 0.11995 | -0.31051 | 0.001041 | 5.1e-05 | -0.2587 | 0.1041 | 0.0427
freq_sweep | 8 | 1 | 4 | 1 | 120 | 1 | 0.999 | 57 | 120.0235 | 0.99996 | 99.901 | 0.1199 | -0.0053373 | 0.12002 | 119.713 | 1 | 0.11982 | -0.31051 | 4.1e-05 | -8.1e-05 | -0.2587 | 0.0041 | -0.0674
freq_sweep | 8 | 1 | 5 | 1 | 120 | 1 | 0.999 | 57 | 120.0235 | 0.99996 | 99.901 | 0.1199 | -0.0053373 | 0.12002 | 119.713 | 1.001 | 0.11992 | -0.31051 | 0.001041 | 2e-05 | -0.2587 | 0.1041 | 0.0169
freq_sweep | 8 | 1 | 6 | 1 | 120 | 1 | 0.999 | 57 | 120.0235 | 0.99996 | 99.901 | 0.1199 | -0.0053373 | 0.12002 | 119.713 | 0.998 | 0.11958 | -0.31051 | -0.001959 | -0.000322 | -0.2587 | -0.1959 | -0.2684
freq_sweep | 8 | 1 | 7 | 2 | 120 | 1 | 0.999 | 57 | 119.9817 | 1.0002 | 99.9035 | 0.11989 | -0.0052698 | 0.12 | 119.758 | 2.978 | 0.35678 | -0.22368 | 1.9778 | 0.23689 | -0.1864 | -0.7496 | -0.80093
freq_sweep | 8 | 1 | 8 | 2 | 120 | 1 | 0.999 | 57 | 119.9817 | 1.0002 | 99.9035 | 0.11989 | -0.0052698 | 0.12 | 119.758 | 1 | 0.11983 | -0.22368 | -0.000164 | -5.3e-05 | -0.1864 | -0.0164 | -0.0438
freq_sweep | 8 | 1 | 9 | 2 | 120 | 1 | 0.999 | 57 | 119.9817 | 1.0002 | 99.9035 | 0.11989 | -0.0052698 | 0.12 | 119.758 | 1 | 0.11987 | -0.22368 | -0.000164 | -1.5e-05 | -0.1864 | -0.0164 | -0.0121
freq_sweep | 8 | 1 | 10 | 2 | 120 | 1 | 0.999 | 57 | 119.9817 | 1.0002 | 99.9035 | 0.11989 | -0.0052698 | 0.12 | 119.758 | 1 | 0.11985 | -0.22368 | -0.000164 | -3.9e-05 | -0.1864 | -0.0164 | -0.0321
freq_sweep | 8 | 1 | 11 | 2 | 120 | 1 | 0.999 | 57 | 119.9817 | 1.0002 | 99.9035 | 0.11989 | -0.0052698 | 0.12 | 119.758 | 1 | 0.11981 | -0.22368 | -0.000164 | -7.9e-05 | -0.1864 | -0.0164 | -0.0655
freq_sweep | 8 | 1 | 12 | 2 | 120 | 1 | 0.999 | 57 | 119.9817 | 1.0002 | 99.9035 | 0.11989 | -0.0052698 | 0.12 | 119.758 | 1 | 0.11981 | -0.22368 | -0.000164 | -7.6e-05 | -0.1864 | -0.0164 | -0.063
freq_sweep | 8 | 1 | 13 | 3 | 120 | 1 | 0.999 | 57 | 119.9939 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.771 | 1 | 0.11979 | -0.22287 | 8.6e-05 | -7.8e-05 | -0.1857 | 0.0086 | -0.0652
freq_sweep | 8 | 1 | 14 | 3 | 120 | 1 | 0.999 | 57 | 119.9939 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.771 | 0.999 | 0.11973 | -0.22287 | -0.000914 | -0.000137 | -0.1857 | -0.0914 | -0.1145
freq_sweep | 8 | 1 | 15 | 3 | 120 | 1 | 0.999 | 57 | 119.9939 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.771 | 1 | 0.11989 | -0.22287 | 8.6e-05 | 2.2e-05 | -0.1857 | 0.0086 | 0.0182
freq_sweep | 8 | 1 | 16 | 3 | 120 | 1 | 0.999 | 57 | 119.9939 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.771 | 0.999 | 0.11975 | -0.22287 | -0.000914 | -0.000116 | -0.1857 | -0.0914 | -0.0969
freq_sweep | 8 | 1 | 17 | 3 | 120 | 1 | 0.999 | 57 | 119.9939 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.771 | 1 | 0.11978 | -0.22287 | 8.6e-05 | -8.6e-05 | -0.1857 | 0.0086 | -0.0719
freq_sweep | 8 | 1 | 18 | 3 | 120 | 1 | 0.999 | 57 | 119.9939 | 0.99991 | 99.9005 | 0.11986 | -0.0053498 | 0.11998 | 119.771 | 0.999 | 0.11976 | -0.22287 | -0.000914 | -0.0001 | -0.1857 | -0.0914 | -0.0836
freq_sweep | 8 | 2 | 1 | 1 | 120 | 1 | 0.999 | 57 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053401 | 0.12001 | 119.733 | 2.968 | 0.35539 | -0.27844 | 1.968 | 0.2355 | -0.232 | -1.0649 | -1.1902
freq_sweep | 8 | 2 | 2 | 1 | 120 | 1 | 0.999 | 57 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053401 | 0.12001 | 119.733 | 1.001 | 0.11994 | -0.27844 | 0.001017 | 5e-05 | -0.232 | 0.1017 | 0.0414
freq_sweep | 8 | 2 | 3 | 1 | 120 | 1 | 0.999 | 57 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053401 | 0.12001 | 119.733 | 1.002 | 0.11997 | -0.27844 | 0.002017 | 8.4e-05 | -0.232 | 0.2017 | 0.0698
freq_sweep | 8 | 2 | 4 | 1 | 120 | 1 | 0.999 | 57 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053401 | 0.12001 | 119.733 | 1 | 0.11983 | -0.27844 | 1.7e-05 | -5.9e-05 | -0.232 | 0.0017 | -0.0495
freq_sweep | 8 | 2 | 5 | 1 | 120 | 1 | 0.999 | 57 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053401 | 0.12001 | 119.733 | 1.001 | 0.11994 | -0.27844 | 0.001017 | 4.8e-05 | -0.232 | 0.1017 | 0.0398
freq_sweep | 8 | 2 | 6 | 1 | 120 | 1 | 0.999 | 57 | 120.0114 | 0.99998 | 99.9008 | 0.11989 | -0.0053401 | 0.12001 | 119.733 | 0.998 | 0.11957 | -0.27844 | -0.001983 | -0.000316 | -0.232 | -0.1983 | -0.2638
freq_sweep | 8 | 2 | 7 | 2 | 120 | 1 | 0.999 | 57 | 119.9788 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 2.978 | 0.35679 | -0.20581 | 1.9778 | 0.23691 | -0.1715 | -0.75062 | -0.79538
freq_sweep | 8 | 2 | 8 | 2 | 120 | 1 | 0.999 | 57 | 119.9788 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.11981 | -0.20581 | -0.000174 | -7.3e-05 | -0.1715 | -0.0174 | -0.0611
freq_sweep | 8 | 2 | 9 | 2 | 120 | 1 | 0.999 | 57 | 119.9788 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.11987 | -0.20581 | -0.000174 | -1.7e-05 | -0.1715 | -0.0174 | -0.0143
freq_sweep | 8 | 2 | 10 | 2 | 120 | 1 | 0.999 | 57 | 119.9788 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.11986 | -0.20581 | -0.000174 | -2.4e-05 | -0.1715 | -0.0174 | -0.0202
freq_sweep | 8 | 2 | 11 | 2 | 120 | 1 | 0.999 | 57 | 119.9788 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.11979 | -0.20581 | -0.000174 | -9.3e-05 | -0.1715 | -0.0174 | -0.0777
freq_sweep | 8 | 2 | 12 | 2 | 120 | 1 | 0.999 | 57 | 119.9788 | 1.0002 | 99.9037 | 0.11988 | -0.0052635 | 0.12 | 119.773 | 1 | 0.1198 | -0.20581 | -0.000174 | -8e-05 | -0.1715 | -0.0174 | -0.0669
freq_sweep | 8 | 2 | 13 | 3 | 120 | 1 | 0.999 | 57 | 119.9915 | 0.99992 | 99.9006 | 0.11986 | -0.0053485 | 0.11998 | 119.779 | 1 | 0.1198 | -0.21253 | 7.9e-05 | -6.4e-05 | -0.1771 | 0.0079 | -0.0532
freq_sweep | 8 | 2 | 14 | 3 | 120 | 1 | 0.999 | 57 | 119.9915 | 0.99992 | 99.9006 | 0.11986 | -0.0053485 | 0.11998 | 119.779 | 0.999 | 0.11974 | -0.21253 | -0.000921 | -0.000127 | -0.1771 | -0.0921 | -0.1058
freq_sweep | 8 | 2 | 15 | 3 | 120 | 1 | 0.999 | 57 | 119.9915 | 0.99992 | 99.9006 | 0.11986 | -0.0053485 | 0.11998 | 119.779 | 1.001 | 0.11991 | -0.21253 | 0.001079 | 5.1e-05 | -0.1771 | 0.1079 | 0.0427
freq_sweep | 8 | 2 | 16 | 3 | 120 | 1 | 0.999 | 57 | 119.9915 | 0.99992 | 99.9006 | 0.11986 | -0.0053485 | 0.11998 | 119.779 | 0.999 | 0.11976 | -0.21253 | -0.000921 | -0.000105 | -0.1771 | -0.0921 | -0.0874
freq_sweep | 8 | 2 | 17 | 3 | 120 | 1 | 0.999 | 57 | 119.9915 | 0.99992 | 99.9006 | 0.11986 | -0.0053485 | 0.11998 | 119.779 | 1 | 0.11979 | -0.21253 | 7.9e-05 | -7.8e-05 | -0.1771 | 0.0079 | -0.0649
freq_sweep | 8 | 2 | 18 | 3 | 120 | 1 | 0.999 | 57 | 119.9915 | 0.99992 | 99.9006 | 0.11986 | -0.0053485 | 0.11998 | 119.779 | 1 | 0.1198 | -0.21253 | 7.9e-05 | -6.1e-05 | -0.1771 | 0.0079 | -0.0507
freq_sweep | 8 | 3 | 1 | 1 | 120 | 1 | 0.999 | 57 | 119.9986 | 0.99998 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.753 | 2.968 | 0.3555 | -0.2456 | 1.968 | 0.23562 | -0.2047 | -1.0645 | -1.1481
freq_sweep | 8 | 3 | 2 | 1 | 120 | 1 | 0.999 | 57 | 119.9986 | 0.99998 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.753 | 1.001 | 0.11991 | -0.2456 | 0.001022 | 3.8e-05 | -0.2047 | 0.1022 | 0.032
freq_sweep | 8 | 3 | 3 | 1 | 120 | 1 | 0.999 | 57 | 119.9986 | 0.99998 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.753 | 1.001 | 0.11996 | -0.2456 | 0.001022 | 8.2e-05 | -0.2047 | 0.1022 | 0.0687
freq_sweep | 8 | 3 | 4 | 1 | 120 | 1 | 0.999 | 57 | 119.9986 | 0.99998 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.753 | 1 | 0.11987 | -0.2456 | 2.2e-05 | -6e-06 | -0.2047 | 0.0022 | -0.0047
freq_sweep | 8 | 3 | 5 | 1 | 120 | 1 | 0.999 | 57 | 119.9986 | 0.99998 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.753 | 1.001 | 0.11991 | -0.2456 | 0.001022 | 3.6e-05 | -0.2047 | 0.1022 | 0.0303
freq_sweep | 8 | 3 | 6 | 1 | 120 | 1 | 0.999 | 57 | 119.9986 | 0.99998 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.753 | 0.998 | 0.11958 | -0.2456 | -0.001978 | -0.000293 | -0.2047 | -0.1978 | -0.2441
freq_sweep | 8 | 3 | 7 | 2 | 120 | 1 | 0.999 | 57 | 119.9851 | 1.0002 | 99.9038 | 0.11989 | -0.0052609 | 0.12 | 119.762 | 2.978 | 0.35673 | -0.22306 | 1.9778 | 0.23684 | -0.1859 | -0.74964 | -0.81719
freq_sweep | 8 | 3 | 8 | 2 | 120 | 1 | 0.999 | 57 | 119.9851 | 1.0002 | 99.9038 | 0.11989 | -0.0052609 | 0.12 | 119.762 | 1 | 0.11982 | -0.22306 | -0.000164 | -6.9e-05 | -0.1859 | -0.0164 | -0.0579
freq_sweep | 8 | 3 | 9 | 2 | 120 | 1 | 0.999 | 57 | 119.9851 | 1.0002 | 99.9038 | 0.11989 | -0.0052609 | 0.12 | 119.762 | 1.001 | 0.1199 | -0.22306 | 0.000836 | 1.5e-05 | -0.1859 | 0.0836 | 0.0122
freq_sweep | 8 | 3 | 10 | 2 | 120 | 1 | 0.999 | 57 | 119.9851 | 1.0002 | 99.9038 | 0.11989 | -0.0052609 | 0.12 | 119.762 | 1 | 0.11983 | -0.22306 | -0.000164 | -5.5e-05 | -0.1859 | -0.0164 | -0.0462
freq_sweep | 8 | 3 | 11 | 2 | 120 | 1 | 0.999 | 57 | 119.9851 | 1.0002 | 99.9038 | 0.11989 | -0.0052609 | 0.12 | 119.762 | 1 | 0.11978 | -0.22306 | -0.000164 | -0.00011 | -0.1859 | -0.0164 | -0.0921
freq_sweep | 8 | 3 | 12 | 2 | 120 | 1 | 0.999 | 57 | 119.9851 | 1.0002 | 99.9038 | 0.11989 | -0.0052609 | 0.12 | 119.762 | 1 | 0.11984 | -0.22306 | -0.000164 | -4.9e-05 | -0.1859 | -0.0164 | -0.0412
freq_sweep | 8 | 3 | 13 | 3 | 120 | 1 | 0.999 | 57 | 119.9872 | 0.99992 | 99.9006 | 0.11986 | -0.0053468 | 0.11998 | 119.768 | 1 | 0.11977 | -0.21921 | 7.7e-05 | -9.1e-05 | -0.1827 | 0.0077 | -0.0757
freq_sweep | 8 | 3 | 14 | 3 | 120 | 1 | 0.999 | 57 | 119.9872 | 0.99992 | 99.9006 | 0.11986 | -0.0053468 | 0.11998 | 119.768 | 0.999 | 0.11975 | -0.21921 | -0.000923 | -0.000112 | -0.1827 | -0.0923 | -0.0932
freq_sweep | 8 | 3 | 15 | 3 | 120 | 1 | 0.999 | 57 | 119.9872 | 0.99992 | 99.9006 | 0.11986 | -0.0053468 | 0.11998 | 119.768 | 1 | 0.11988 | -0.21921 | 7.7e-05 | 2.1e-05 | -0.1827 | 0.0077 | 0.0178
freq_sweep | 8 | 3 | 16 | 3 | 120 | 1 | 0.999 | 57 | 119.9872 | 0.99992 | 99.9006 | 0.11986 | -0.0053468 | 0.11998 | 119.768 | 0.999 | 0.11973 | -0.21921 | -0.000923 | -0.000126 | -0.1827 | -0.0923 | -0.1049
freq_sweep | 8 | 3 | 17 | 3 | 120 | 1 | 0.999 | 57 | 119.9872 | 0.99992 | 99.9006 | 0.11986 | -0.0053468 | 0.11998 | 119.768 | 1 | 0.11981 | -0.21921 | 7.7e-05 | -4.4e-05 | -0.1827 | 0.0077 | -0.0365
freq_sweep | 8 | 3 | 18 | 3 | 120 | 1 | 0.999 | 57 | 119.9872 | 0.99992 | 99.9006 | 0.11986 | -0.0053468 | 0.11998 | 119.768 | 1 | 0.11977 | -0.21921 | 7.7e-05 | -8.9e-05 | -0.1827 | 0.0077 | -0.074
freq_sweep | 9 | 1 | 1 | 1 | 120 | 1 | 0.999 | 58 | 120.0179 | 0.99999 | 99.9009 | 0.1199 | -0.00534 | 0.12002 | 119.724 | 2.968 | 0.35544 | -0.29385 | 1.968 | 0.23554 | -0.2448 | -1.0657 | -1.1836
freq_sweep | 9 | 1 | 2 | 1 | 120 | 1 | 0.999 | 58 | 120.0179 | 0.99999 | 99.9009 | 0.1199 | -0.00534 | 0.12002 | 119.724 | 1.001 | 0.11992 | -0.29385 | 0.00101 | 2.1e-05 | -0.2448 | 0.101 | 0.0177
freq_sweep | 9 | 1 | 3 | 1 | 120 | 1 | 0.999 | 58 | 120.0179 | 0.99999 | 99.9009 | 0.1199 | -0.00534 | 0.12002 | 119.724 | 1.001 | 0.11996 | -0.29385 | 0.00101 | 6.6e-05 | -0.2448 | 0.101 | 0.0553
freq_sweep | 9 | 1 | 4 | 1 | 120 | 1 | 0.999 | 58 | 120.0179 | 0.99999 | 99.9009 | 0.1199 | -0.00534 | 0.12002 | 119.724 | 1 | 0.11984 | -0.29385 | 1e-05 | -6.2e-05 | -0.2448 | 0.001 | -0.0515
freq_sweep | 9 | 1 | 5 | 1 | 120 | 1 | 0.999 | 58 | 120.0179 | 0.99999 | 99.9009 | 0.1199 | -0.00534 | 0.12002 | 119.724 | 1.001 | 0.11991 | -0.29385 | 0.00101 | 9e-06 | -0.2448 | 0.101 | 0.0077
freq_sweep | 9 | 1 | 6 | 1 | 120 | 1 | 0.999 | 58 | 120.0179 | 0.99999 | 99.9009 | 0.1199 | -0.00534 | 0.12002 | 119.724 | 0.998 | 0.11961 | -0.29385 | -0.00199 | -0.000287 | -0.2448 | -0.199 | -0.2391
freq_sweep | 9 | 1 | 7 | 2 | 120 | 1 | 0.999 | 58 | 119.9806 | 1.0002 | 99.9038 | 0.11988 | -0.0052627 | 0.12 | 119.771 | 2.978 | 0.35673 | -0.20958 | 1.9778 | 0.23685 | -0.1747 | -0.74955 | -0.81249
freq_sweep | 9 | 1 | 8 | 2 | 120 | 1 | 0.999 | 58 | 119.9806 | 1.0002 | 99.9038 | 0.11988 | -0.0052627 | 0.12 | 119.771 | 1 | 0.11984 | -0.20958 | -0.000163 | -4.9e-05 | -0.1747 | -0.0163 | -0.0406
freq_sweep | 9 | 1 | 9 | 2 | 120 | 1 | 0.999 | 58 | 119.9806 | 1.0002 | 99.9038 | 0.11988 | -0.0052627 | 0.12 | 119.771 | 1 | 0.11988 | -0.20958 | -0.000163 | -9e-06 | -0.1747 | -0.0163 | -0.0073
freq_sweep | 9 | 1 | 10 | 2 | 120 | 1 | 0.999 | 58 | 119.9806 | 1.0002 | 99.9038 | 0.11988 | -0.0052627 | 0.12 | 119.771 | 1 | 0.11983 | -0.20958 | -0.000163 | -5e-05 | -0.1747 | -0.0163 | -0.0415
freq_sweep | 9 | 1 | 11 | 2 | 120 | 1 | 0.999 | 58 | 119.9806 | 1.0002 | 99.9038 | 0.11988 | -0.0052627 | 0.12 | 119.771 | 1 | 0.11979 | -0.20958 | -0.000163 | -9e-05 | -0.1747 | -0.0163 | -0.0748
freq_sweep | 9 | 1 | 12 | 2 | 120 | 1 | 0.999 | 58 | 119.9806 | 1.0002 | 99.9038 | 0.11988 | -0.0052627 | 0.12 | 119.771 | 1 | 0.11981 | -0.20958 | -0.000163 | -7.3e-05 | -0.1747 | -0.0163 | -0.0607
freq_sweep | 9 | 1 | 13 | 3 | 120 | 1 | 0.999 | 58 | 119.99 | 0.99992 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.779 | 1 | 0.11981 | -0.21095 | 8.2e-05 | -5.4e-05 | -0.1758 | 0.0082 | -0.0449
freq_sweep | 9 | 1 | 14 | 3 | 120 | 1 | 0.999 | 58 | 119.99 | 0.99992 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.779 | 0.999 | 0.11975 | -0.21095 | -0.000918 | -0.000114 | -0.1758 | -0.0918 | -0.0949
freq_sweep | 9 | 1 | 15 | 3 | 120 | 1 | 0.999 | 58 | 119.99 | 0.99992 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.779 | 1 | 0.1199 | -0.21095 | 8.2e-05 | 4.1e-05 | -0.1758 | 0.0082 | 0.0344
freq_sweep | 9 | 1 | 16 | 3 | 120 | 1 | 0.999 | 58 | 119.99 | 0.99992 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.779 | 0.999 | 0.11976 | -0.21095 | -0.000918 | -0.000101 | -0.1758 | -0.0918 | -0.0841
freq_sweep | 9 | 1 | 17 | 3 | 120 | 1 | 0.999 | 58 | 119.99 | 0.99992 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.779 | 1 | 0.11981 | -0.21095 | 8.2e-05 | -5.4e-05 | -0.1758 | 0.0082 | -0.0449
freq_sweep | 9 | 1 | 18 | 3 | 120 | 1 | 0.999 | 58 | 119.99 | 0.99992 | 99.9005 | 0.11986 | -0.00535 | 0.11998 | 119.779 | 1 | 0.11979 | -0.21095 | 8.2e-05 | -7.6e-05 | -0.1758 | 0.0082 | -0.0632
freq_sweep | 9 | 2 | 1 | 1 | 120 | 1 | 0.999 | 58 | 120.0135 | 0.99998 | 99.9008 | 0.11989 | -0.0053419 | 0.12001 | 119.737 | 2.968 | 0.35546 | -0.2765 | 1.968 | 0.23557 | -0.2304 | -1.0642 | -1.1712
freq_sweep | 9 | 2 | 2 | 1 | 120 | 1 | 0.999 | 58 | 120.0135 | 0.99998 | 99.9008 | 0.11989 | -0.0053419 | 0.12001 | 119.737 | 1.001 | 0.11997 | -0.2765 | 0.001025 | 8e-05 | -0.2304 | 0.1025 | 0.0664
freq_sweep | 9 | 2 | 3 | 1 | 120 | 1 | 0.999 | 58 | 120.0135 | 0.99998 | 99.9008 | 0.11989 | -0.0053419 | 0.12001 | 119.737 | 1.001 | 0.11997 | -0.2765 | 0.001025 | 7.9e-05 | -0.2304 | 0.1025 | 0.0655
freq_sweep | 9 | 2 | 4 | 1 | 120 | 1 | 0.999 | 58 | 120.0135 | 0.99998 | 99.9008 | 0.11989 | -0.0053419 | 0.12001 | 119.737 | 1 | 0.11987 | -0.2765 | 2.5e-05 | -2.5e-05 | -0.2304 | 0.0025 | -0.0212
freq_sweep | 9 | 2 | 5 | 1 | 120 | 1 | 0.999 | 58 | 120.0135 | 0.99998 | 99.9008 | 0.11989 | -0.0053419 | 0.12001 | 119.737 | 1.001 | 0.11995 | -0.2765 | 0.001025 | 6.2e-05 | -0.2304 | 0.1025 | 0.0513
freq_sweep | 9 | 2 | 6 | 1 | 120 | 1 | 0.999 | 58 | 120.0135 | 0.99998 | 99.9008 | 0.11989 | -0.0053419 | 0.12001 | 119.737 | 0.998 | 0.11959 | -0.2765 | -0.001975 | -0.000299 | -0.2304 | -0.1975 | -0.2498
freq_sweep | 9 | 2 | 7 | 2 | 120 | 1 | 0.999 | 58 | 119.983 | 1.0002 | 99.9038 | 0.11989 | -0.005262 | 0.12 | 119.764 | 2.979 | 0.35682 | -0.21903 | 1.9788 | 0.23693 | -0.1826 | -0.71806 | -0.79223
freq_sweep | 9 | 2 | 8 | 2 | 120 | 1 | 0.999 | 58 | 119.983 | 1.0002 | 99.9038 | 0.11989 | -0.005262 | 0.12 | 119.764 | 1 | 0.1198 | -0.21903 | -0.000182 | -9.2e-05 | -0.1826 | -0.0182 | -0.0771
freq_sweep | 9 | 2 | 9 | 2 | 120 | 1 | 0.999 | 58 | 119.983 | 1.0002 | 99.9038 | 0.11989 | -0.005262 | 0.12 | 119.764 | 1 | 0.11988 | -0.21903 | -0.000182 | -9e-06 | -0.1826 | -0.0182 | -0.0079
freq_sweep | 9 | 2 | 10 | 2 | 120 | 1 | 0.999 | 58 | 119.983 | 1.0002 | 99.9038 | 0.11989 | -0.005262 | 0.12 | 119.764 | 1 | 0.11984 | -0.21903 | -0.000182 | -4.5e-05 | -0.1826 | -0.0182 | -0.0379
freq_sweep | 9 | 2 | 11 | 2 | 120 | 1 | 0.999 | 58 | 119.983 | 1.0002 | 99.9038 | 0.11989 | -0.005262 | 0.12 | 119.764 | 1 | 0.11976 | -0.21903 | -0.000182 | -0.000127 | -0.1826 | -0.0182 | -0.1063
freq_sweep | 9 | 2 | 12 | 2 | 120 | 1 | 0.999 | 58 | 119.983 | 1.0002 | 99.9038 | 0.11989 | -0.005262 | 0.12 | 119.764 | 1 | 0.11982 | -0.21903 | -0.000182 | -7e-05 | -0.1826 | -0.0182 | -0.0588
freq_sweep | 9 | 2 | 13 | 3 | 120 | 1 | 0.999 | 58 | 119.9897 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.783 | 0.999 | 0.11976 | -0.20668 | -0.000923 | -9.6e-05 | -0.1722 | -0.0923 | -0.0802
freq_sweep | 9 | 2 | 14 | 3 | 120 | 1 | 0.999 | 58 | 119.9897 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.783 | 0.999 | 0.11975 | -0.20668 | -0.000923 | -0.000114 | -0.1722 | -0.0923 | -0.0952
freq_sweep | 9 | 2 | 15 | 3 | 120 | 1 | 0.999 | 58 | 119.9897 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.783 | 1.001 | 0.11991 | -0.20668 | 0.001077 | 4.7e-05 | -0.1722 | 0.1077 | 0.0391
freq_sweep | 9 | 2 | 16 | 3 | 120 | 1 | 0.999 | 58 | 119.9897 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.783 | 0.999 | 0.11972 | -0.20668 | -0.000923 | -0.000137 | -0.1722 | -0.0923 | -0.1144
freq_sweep | 9 | 2 | 17 | 3 | 120 | 1 | 0.999 | 58 | 119.9897 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.783 | 1 | 0.1198 | -0.20668 | 7.7e-05 | -5.6e-05 | -0.1722 | 0.0077 | -0.0468
freq_sweep | 9 | 2 | 18 | 3 | 120 | 1 | 0.999 | 58 | 119.9897 | 0.99992 | 99.9006 | 0.11986 | -0.0053472 | 0.11998 | 119.783 | 1 | 0.1198 | -0.20668 | 7.7e-05 | -6.1e-05 | -0.1722 | 0.0077 | -0.051
freq_sweep | 9 | 3 | 1 | 1 | 120 | 1 | 0.999 | 58 | 119.9854 | 0.99995 | 99.9 | 0.11986 | -0.0053497 | 0.11998 | 119.7 | 2.968 | 0.35528 | -0.28543 | 1.968 | 0.23542 | -0.2379 | -1.062 | -1.1954
freq_sweep | 9 | 3 | 2 | 1 | 120 | 1 | 0.999 | 58 | 119.9854 | 0.99995 | 99.9 | 0.11986 | -0.0053497 | 0.11998 | 119.7 | 1.001 | 0.11987 | -0.28543 | 0.001047 | 1.3e-05 | -0.2379 | 0.1047 | 0.011
freq_sweep | 9 | 3 | 3 | 1 | 120 | 1 | 0.999 | 58 | 119.9854 | 0.99995 | 99.9 | 0.11986 | -0.0053497 | 0.11998 | 119.7 | 1.002 | 0.11995 | -0.28543 | 0.002047 | 9e-05 | -0.2379 | 0.2047 | 0.0752
freq_sweep | 9 | 3 | 4 | 1 | 120 | 1 | 0.999 | 58 | 119.9854 | 0.99995 | 99.9 | 0.11986 | -0.0053497 | 0.11998 | 119.7 | 1 | 0.11979 | -0.28543 | 4.7e-05 | -6.5e-05 | -0.2379 | 0.0047 | -0.0541
freq_sweep | 9 | 3 | 5 | 1 | 120 | 1 | 0.999 | 58 | 119.9854 | 0.99995 | 99.9 | 0.11986 | -0.0053497 | 0.11998 | 119.7 | 1.001 | 0.11986 | -0.28543 | 0.001047 | 5e-06 | -0.2379 | 0.1047 | 0.0043
freq_sweep | 9 | 3 | 6 | 1 | 120 | 1 | 0.999 | 58 | 119.9854 | 0.99995 | 99.9 | 0.11986 | -0.0053497 | 0.11998 | 119.7 | 0.998 | 0.11954 | -0.28543 | -0.001953 | -0.000324 | -0.2379 | -0.1953 | -0.2702
freq_sweep | 9 | 3 | 7 | 2 | 120 | 1 | 0.999 | 58 | 119.9821 | 1.0002 | 99.9038 | 0.11989 | -0.0052617 | 0.12 | 119.763 | 2.978 | 0.35678 | -0.21909 | 1.9778 | 0.23689 | -0.1826 | -0.74939 | -0.8011
freq_sweep | 9 | 3 | 8 | 2 | 120 | 1 | 0.999 | 58 | 119.9821 | 1.0002 | 99.9038 | 0.11989 | -0.0052617 | 0.12 | 119.763 | 1 | 0.11983 | -0.21909 | -0.000162 | -5.3e-05 | -0.1826 | -0.0162 | -0.0443
freq_sweep | 9 | 3 | 9 | 2 | 120 | 1 | 0.999 | 58 | 119.9821 | 1.0002 | 99.9038 | 0.11989 | -0.0052617 | 0.12 | 119.763 | 1 | 0.11988 | -0.21909 | -0.000162 | -5e-06 | -0.1826 | -0.0162 | -0.0042
freq_sweep | 9 | 3 | 10 | 2 | 120 | 1 | 0.999 | 58 | 119.9821 | 1.0002 | 99.9038 | 0.11989 | -0.0052617 | 0.12 | 119.763 | 1 | 0.11984 | -0.21909 | -0.000162 | -5e-05 | -0.1826 | -0.0162 | -0.0418
freq_sweep | 9 | 3 | 11 | 2 | 120 | 1 | 0.999 | 58 | 119.9821 | 1.0002 | 99.9038 | 0.11989 | -0.0052617 | 0.12 | 119.763 | 1 | 0.11981 | -0.21909 | -0.000162 | -8e-05 | -0.1826 | -0.0162 | -0.0668
freq_sweep | 9 | 3 | 12 | 2 | 120 | 1 | 0.999 | 58 | 119.9821 | 1.0002 | 99.9038 | 0.11989 | -0.0052617 | 0.12 | 119.763 | 1 | 0.11982 | -0.21909 | -0.000162 | -6.4e-05 | -0.1826 | -0.0162 | -0.0534
freq_sweep | 9 | 3 | 13 | 3 | 120 | 1 | 0.999 | 58 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053477 | 0.11998 | 119.766 | 1 | 0.11981 | -0.22079 | 8.7e-05 | -4.6e-05 | -0.184 | 0.0087 | -0.0384
freq_sweep | 9 | 3 | 14 | 3 | 120 | 1 | 0.999 | 58 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053477 | 0.11998 | 119.766 | 0.999 | 0.11972 | -0.22079 | -0.000913 | -0.000138 | -0.184 | -0.0913 | -0.1152
freq_sweep | 9 | 3 | 15 | 3 | 120 | 1 | 0.999 | 58 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053477 | 0.11998 | 119.766 | 1 | 0.11987 | -0.22079 | 8.7e-05 | 1.7e-05 | -0.184 | 0.0087 | 0.0141
freq_sweep | 9 | 3 | 16 | 3 | 120 | 1 | 0.999 | 58 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053477 | 0.11998 | 119.766 | 0.999 | 0.11976 | -0.22079 | -0.000913 | -0.0001 | -0.184 | -0.0913 | -0.0835
freq_sweep | 9 | 3 | 17 | 3 | 120 | 1 | 0.999 | 58 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053477 | 0.11998 | 119.766 | 1 | 0.11977 | -0.22079 | 8.7e-05 | -8.7e-05 | -0.184 | 0.0087 | -0.0726
freq_sweep | 9 | 3 | 18 | 3 | 120 | 1 | 0.999 | 58 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053477 | 0.11998 | 119.766 | 0.999 | 0.11976 | -0.22079 | -0.000913 | -9.6e-05 | -0.184 | -0.0913 | -0.0801
freq_sweep | 10 | 1 | 1 | 1 | 120 | 1 | 0.999 | 59 | 120.019 | 0.99996 | 99.9009 | 0.1199 | -0.0053388 | 0.12001 | 119.72 | 2.968 | 0.35535 | -0.29903 | 1.968 | 0.23545 | -0.2492 | -1.0626 | -1.2059
freq_sweep | 10 | 1 | 2 | 1 | 120 | 1 | 0.999 | 59 | 120.019 | 0.99996 | 99.9009 | 0.1199 | -0.0053388 | 0.12001 | 119.72 | 1.001 | 0.11994 | -0.29903 | 0.001041 | 4.8e-05 | -0.2492 | 0.1042 | 0.0399
freq_sweep | 10 | 1 | 3 | 1 | 120 | 1 | 0.999 | 59 | 120.019 | 0.99996 | 99.9009 | 0.1199 | -0.0053388 | 0.12001 | 119.72 | 1.001 | 0.11996 | -0.29903 | 0.001041 | 6.1e-05 | -0.2492 | 0.1042 | 0.0507
freq_sweep | 10 | 1 | 4 | 1 | 120 | 1 | 0.999 | 59 | 120.019 | 0.99996 | 99.9009 | 0.1199 | -0.0053388 | 0.12001 | 119.72 | 1 | 0.11982 | -0.29903 | 4.1e-05 | -7.7e-05 | -0.2492 | 0.0041 | -0.0644
freq_sweep | 10 | 1 | 5 | 1 | 120 | 1 | 0.999 | 59 | 120.019 | 0.99996 | 99.9009 | 0.1199 | -0.0053388 | 0.12001 | 119.72 | 1.001 | 0.11995 | -0.29903 | 0.001041 | 5.4e-05 | -0.2492 | 0.1042 | 0.0449
freq_sweep | 10 | 1 | 6 | 1 | 120 | 1 | 0.999 | 59 | 120.019 | 0.99996 | 99.9009 | 0.1199 | -0.0053388 | 0.12001 | 119.72 | 0.998 | 0.11961 | -0.29903 | -0.001959 | -0.000285 | -0.2492 | -0.1959 | -0.2379
freq_sweep | 10 | 1 | 7 | 2 | 120 | 1 | 0.999 | 59 | 119.9812 | 1.0002 | 99.9036 | 0.11989 | -0.0052666 | 0.12 | 119.774 | 2.979 | 0.35682 | -0.20719 | 1.9788 | 0.23693 | -0.1727 | -0.71697 | -0.79025
freq_sweep | 10 | 1 | 8 | 2 | 120 | 1 | 0.999 | 59 | 119.9812 | 1.0002 | 99.9036 | 0.11989 | -0.0052666 | 0.12 | 119.774 | 1 | 0.11981 | -0.20719 | -0.000171 | -7.7e-05 | -0.1727 | -0.0171 | -0.0643
freq_sweep | 10 | 1 | 9 | 2 | 120 | 1 | 0.999 | 59 | 119.9812 | 1.0002 | 99.9036 | 0.11989 | -0.0052666 | 0.12 | 119.774 | 1.001 | 0.1199 | -0.20719 | 0.000829 | 1.2e-05 | -0.1727 | 0.0829 | 0.01
freq_sweep | 10 | 1 | 10 | 2 | 120 | 1 | 0.999 | 59 | 119.9812 | 1.0002 | 99.9036 | 0.11989 | -0.0052666 | 0.12 | 119.774 | 1 | 0.11986 | -0.20719 | -0.000171 | -2.3e-05 | -0.1727 | -0.0171 | -0.0192
freq_sweep | 10 | 1 | 11 | 2 | 120 | 1 | 0.999 | 59 | 119.9812 | 1.0002 | 99.9036 | 0.11989 | -0.0052666 | 0.12 | 119.774 | 1 | 0.11977 | -0.20719 | -0.000171 | -0.000116 | -0.1727 | -0.0171 | -0.0968
freq_sweep | 10 | 1 | 12 | 2 | 120 | 1 | 0.999 | 59 | 119.9812 | 1.0002 | 99.9036 | 0.11989 | -0.0052666 | 0.12 | 119.774 | 1 | 0.11982 | -0.20719 | -0.000171 | -6.6e-05 | -0.1727 | -0.0171 | -0.0551
freq_sweep | 10 | 1 | 13 | 3 | 120 | 1 | 0.999 | 59 | 119.9914 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.768 | 1 | 0.11978 | -0.22343 | 8e-05 | -8.5e-05 | -0.1862 | 0.008 | -0.0705
freq_sweep | 10 | 1 | 14 | 3 | 120 | 1 | 0.999 | 59 | 119.9914 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.768 | 0.999 | 0.11975 | -0.22343 | -0.00092 | -0.000109 | -0.1862 | -0.092 | -0.0905
freq_sweep | 10 | 1 | 15 | 3 | 120 | 1 | 0.999 | 59 | 119.9914 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.768 | 1.001 | 0.11992 | -0.22343 | 0.00108 | 5.6e-05 | -0.1862 | 0.108 | 0.0471
freq_sweep | 10 | 1 | 16 | 3 | 120 | 1 | 0.999 | 59 | 119.9914 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.768 | 0.999 | 0.11973 | -0.22343 | -0.00092 | -0.000131 | -0.1862 | -0.092 | -0.1089
freq_sweep | 10 | 1 | 17 | 3 | 120 | 1 | 0.999 | 59 | 119.9914 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.768 | 1 | 0.1198 | -0.22343 | 8e-05 | -6.1e-05 | -0.1862 | 0.008 | -0.0505
freq_sweep | 10 | 1 | 18 | 3 | 120 | 1 | 0.999 | 59 | 119.9914 | 0.99992 | 99.9006 | 0.11986 | -0.0053474 | 0.11998 | 119.768 | 1 | 0.1198 | -0.22343 | 8e-05 | -6e-05 | -0.1862 | 0.008 | -0.0497
freq_sweep | 10 | 2 | 1 | 1 | 120 | 1 | 0.999 | 59 | 120.0134 | 0.99997 | 99.9008 | 0.11989 | -0.0053416 | 0.12001 | 119.71 | 2.968 | 0.35542 | -0.30344 | 1.968 | 0.23553 | -0.2528 | -1.0639 | -1.1834
freq_sweep | 10 | 2 | 2 | 1 | 120 | 1 | 0.999 | 59 | 120.0134 | 0.99997 | 99.9008 | 0.11989 | -0.0053416 | 0.12001 | 119.71 | 1.001 | 0.11991 | -0.30344 | 0.001028 | 1.9e-05 | -0.2528 | 0.1028 | 0.0158
freq_sweep | 10 | 2 | 3 | 1 | 120 | 1 | 0.999 | 59 | 120.0134 | 0.99997 | 99.9008 | 0.11989 | -0.0053416 | 0.12001 | 119.71 | 1.001 | 0.11999 | -0.30344 | 0.001028 | 9.7e-05 | -0.2528 | 0.1028 | 0.0808
freq_sweep | 10 | 2 | 4 | 1 | 120 | 1 | 0.999 | 59 | 120.0134 | 0.99997 | 99.9008 | 0.11989 | -0.0053416 | 0.12001 | 119.71 | 1 | 0.11983 | -0.30344 | 2.8e-05 | -5.8e-05 | -0.2528 | 0.0028 | -0.0485
freq_sweep | 10 | 2 | 5 | 1 | 120 | 1 | 0.999 | 59 | 120.0134 | 0.99997 | 99.9008 | 0.11989 | -0.0053416 | 0.12001 | 119.71 | 1.001 | 0.1199 | -0.30344 | 0.001028 | 8e-06 | -0.2528 | 0.1028 | 0.0066
freq_sweep | 10 | 2 | 6 | 1 | 120 | 1 | 0.999 | 59 | 120.0134 | 0.99997 | 99.9008 | 0.11989 | -0.0053416 | 0.12001 | 119.71 | 0.999 | 0.11964 | -0.30344 | -0.000972 | -0.000253 | -0.2528 | -0.0972 | -0.2111
freq_sweep | 10 | 2 | 7 | 2 | 120 | 1 | 0.999 | 59 | 119.9803 | 1.0002 | 99.9038 | 0.11988 | -0.0052618 | 0.12 | 119.773 | 2.978 | 0.35673 | -0.20725 | 1.9778 | 0.23685 | -0.1727 | -0.74895 | -0.81219
freq_sweep | 10 | 2 | 8 | 2 | 120 | 1 | 0.999 | 59 | 119.9803 | 1.0002 | 99.9038 | 0.11988 | -0.0052618 | 0.12 | 119.773 | 1 | 0.11984 | -0.20725 | -0.000157 | -4.1e-05 | -0.1727 | -0.0157 | -0.0339
freq_sweep | 10 | 2 | 9 | 2 | 120 | 1 | 0.999 | 59 | 119.9803 | 1.0002 | 99.9038 | 0.11988 | -0.0052618 | 0.12 | 119.773 | 1.001 | 0.11988 | -0.20725 | 0.000843 | -5e-06 | -0.1727 | 0.0843 | -0.0039
freq_sweep | 10 | 2 | 10 | 2 | 120 | 1 | 0.999 | 59 | 119.9803 | 1.0002 | 99.9038 | 0.11988 | -0.0052618 | 0.12 | 119.773 | 1 | 0.11983 | -0.20725 | -0.000157 | -5.5e-05 | -0.1727 | -0.0157 | -0.0456
freq_sweep | 10 | 2 | 11 | 2 | 120 | 1 | 0.999 | 59 | 119.9803 | 1.0002 | 99.9038 | 0.11988 | -0.0052618 | 0.12 | 119.773 | 1 | 0.11979 | -0.20725 | -0.000157 | -8.9e-05 | -0.1727 | -0.0157 | -0.074
freq_sweep | 10 | 2 | 12 | 2 | 120 | 1 | 0.999 | 59 | 119.9803 | 1.0002 | 99.9038 | 0.11988 | -0.0052618 | 0.12 | 119.773 | 1 | 0.11981 | -0.20725 | -0.000157 | -7.6e-05 | -0.1727 | -0.0157 | -0.0631
freq_sweep | 10 | 2 | 13 | 3 | 120 | 1 | 0.999 | 59 | 119.9888 | 0.99992 | 99.9005 | 0.11986 | -0.0053512 | 0.11998 | 119.777 | 1 | 0.11981 | -0.21181 | 8.3e-05 | -5.2e-05 | -0.1765 | 0.0083 | -0.0438
freq_sweep | 10 | 2 | 14 | 3 | 120 | 1 | 0.999 | 59 | 119.9888 | 0.99992 | 99.9005 | 0.11986 | -0.0053512 | 0.11998 | 119.777 | 0.999 | 0.11975 | -0.21181 | -0.000917 | -0.000112 | -0.1765 | -0.0918 | -0.0938
freq_sweep | 10 | 2 | 15 | 3 | 120 | 1 | 0.999 | 59 | 119.9888 | 0.99992 | 99.9005 | 0.11986 | -0.0053512 | 0.11998 | 119.777 | 1 | 0.11988 | -0.21181 | 8.3e-05 | 2.5e-05 | -0.1765 | 0.0083 | 0.0205
freq_sweep | 10 | 2 | 16 | 3 | 120 | 1 | 0.999 | 59 | 119.9888 | 0.99992 | 99.9005 | 0.11986 | -0.0053512 | 0.11998 | 119.777 | 0.999 | 0.11976 | -0.21181 | -0.000917 | -9.4e-05 | -0.1765 | -0.0918 | -0.0788
freq_sweep | 10 | 2 | 17 | 3 | 120 | 1 | 0.999 | 59 | 119.9888 | 0.99992 | 99.9005 | 0.11986 | -0.0053512 | 0.11998 | 119.777 | 1 | 0.1198 | -0.21181 | 8.3e-05 | -6.2e-05 | -0.1765 | 0.0083 | -0.0521
freq_sweep | 10 | 2 | 18 | 3 | 120 | 1 | 0.999 | 59 | 119.9888 | 0.99992 | 99.9005 | 0.11986 | -0.0053512 | 0.11998 | 119.777 | 1 | 0.11979 | -0.21181 | 8.3e-05 | -6.8e-05 | -0.1765 | 0.0083 | -0.0571
freq_sweep | 10 | 3 | 1 | 1 | 120 | 1 | 0.999 | 59 | 120.0019 | 0.99999 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.73 | 2.968 | 0.35547 | -0.27185 | 1.968 | 0.23559 | -0.2265 | -1.0657 | -1.1598
freq_sweep | 10 | 3 | 2 | 1 | 120 | 1 | 0.999 | 59 | 120.0019 | 0.99999 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.73 | 1.001 | 0.11994 | -0.27185 | 0.00101 | 5.5e-05 | -0.2265 | 0.101 | 0.0455
freq_sweep | 10 | 3 | 3 | 1 | 120 | 1 | 0.999 | 59 | 120.0019 | 0.99999 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.73 | 1.001 | 0.11995 | -0.27185 | 0.00101 | 7.4e-05 | -0.2265 | 0.101 | 0.0614
freq_sweep | 10 | 3 | 4 | 1 | 120 | 1 | 0.999 | 59 | 120.0019 | 0.99999 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.73 | 1.001 | 0.11986 | -0.27185 | 0.00101 | -1.6e-05 | -0.2265 | 0.101 | -0.0137
freq_sweep | 10 | 3 | 5 | 1 | 120 | 1 | 0.999 | 59 | 120.0019 | 0.99999 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.73 | 1.001 | 0.11993 | -0.27185 | 0.00101 | 5.3e-05 | -0.2265 | 0.101 | 0.0439
freq_sweep | 10 | 3 | 6 | 1 | 120 | 1 | 0.999 | 59 | 120.0019 | 0.99999 | 99.9006 | 0.11988 | -0.0053444 | 0.12 | 119.73 | 0.998 | 0.1196 | -0.27185 | -0.00199 | -0.000285 | -0.2265 | -0.199 | -0.2381
freq_sweep | 10 | 3 | 7 | 2 | 120 | 1 | 0.999 | 59 | 119.9823 | 1.0002 | 99.9038 | 0.11989 | -0.0052612 | 0.12 | 119.757 | 2.978 | 0.35676 | -0.22531 | 1.9778 | 0.23688 | -0.1878 | -0.74909 | -0.80463
freq_sweep | 10 | 3 | 8 | 2 | 120 | 1 | 0.999 | 59 | 119.9823 | 1.0002 | 99.9038 | 0.11989 | -0.0052612 | 0.12 | 119.757 | 1 | 0.1198 | -0.22531 | -0.000159 | -8.9e-05 | -0.1878 | -0.0159 | -0.0742
freq_sweep | 10 | 3 | 9 | 2 | 120 | 1 | 0.999 | 59 | 119.9823 | 1.0002 | 99.9038 | 0.11989 | -0.0052612 | 0.12 | 119.757 | 1.001 | 0.11992 | -0.22531 | 0.000841 | 3.6e-05 | -0.1878 | 0.0841 | 0.0301
freq_sweep | 10 | 3 | 10 | 2 | 120 | 1 | 0.999 | 59 | 119.9823 | 1.0002 | 99.9038 | 0.11989 | -0.0052612 | 0.12 | 119.757 | 1 | 0.11984 | -0.22531 | -0.000159 | -4.1e-05 | -0.1878 | -0.0159 | -0.0342
freq_sweep | 10 | 3 | 11 | 2 | 120 | 1 | 0.999 | 59 | 119.9823 | 1.0002 | 99.9038 | 0.11989 | -0.0052612 | 0.12 | 119.757 | 1 | 0.11977 | -0.22531 | -0.000159 | -0.000117 | -0.1878 | -0.0159 | -0.0976
freq_sweep | 10 | 3 | 12 | 2 | 120 | 1 | 0.999 | 59 | 119.9823 | 1.0002 | 99.9038 | 0.11989 | -0.0052612 | 0.12 | 119.757 | 1 | 0.11983 | -0.22531 | -0.000159 | -6.1e-05 | -0.1878 | -0.0159 | -0.0509
freq_sweep | 10 | 3 | 13 | 3 | 120 | 1 | 0.999 | 59 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053476 | 0.11998 | 119.776 | 1 | 0.11978 | -0.21078 | 9e-05 | -7.9e-05 | -0.1757 | 0.009 | -0.0656
freq_sweep | 10 | 3 | 14 | 3 | 120 | 1 | 0.999 | 59 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053476 | 0.11998 | 119.776 | 0.999 | 0.11975 | -0.21078 | -0.00091 | -0.000109 | -0.1757 | -0.091 | -0.0907
freq_sweep | 10 | 3 | 15 | 3 | 120 | 1 | 0.999 | 59 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053476 | 0.11998 | 119.776 | 1 | 0.11988 | -0.21078 | 9e-05 | 2.3e-05 | -0.1757 | 0.009 | 0.0195
freq_sweep | 10 | 3 | 16 | 3 | 120 | 1 | 0.999 | 59 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053476 | 0.11998 | 119.776 | 0.999 | 0.11973 | -0.21078 | -0.00091 | -0.000127 | -0.1757 | -0.091 | -0.1057
freq_sweep | 10 | 3 | 17 | 3 | 120 | 1 | 0.999 | 59 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053476 | 0.11998 | 119.776 | 1 | 0.11981 | -0.21078 | 9e-05 | -4.5e-05 | -0.1757 | 0.009 | -0.0373
freq_sweep | 10 | 3 | 18 | 3 | 120 | 1 | 0.999 | 59 | 119.9868 | 0.99991 | 99.9006 | 0.11986 | -0.0053476 | 0.11998 | 119.776 | 1 | 0.11978 | -0.21078 | 9e-05 | -7.3e-05 | -0.1757 | 0.009 | -0.0606
freq_sweep | 11 | 1 | 1 | 1 | 120 | 1 | 0.999 | 60 | 120.0182 | 0.99998 | 99.9007 | 0.1199 | -0.005339 | 0.12002 | 119.765 | 2.968 | 0.3554 | -0.2532 | 1.968 | 0.2355 | -0.211 | -1.065 | -1.1938
freq_sweep | 11 | 1 | 2 | 1 | 120 | 1 | 0.999 | 60 | 120.0182 | 0.99998 | 99.9007 | 0.1199 | -0.005339 | 0.12002 | 119.765 | 1.001 | 0.11991 | -0.2532 | 0.001017 | 1.5e-05 | -0.211 | 0.1017 | 0.0125
freq_sweep | 11 | 1 | 3 | 1 | 120 | 1 | 0.999 | 60 | 120.0182 | 0.99998 | 99.9007 | 0.1199 | -0.005339 | 0.12002 | 119.765 | 1.001 | 0.11998 | -0.2532 | 0.001017 | 7.9e-05 | -0.211 | 0.1017 | 0.0659
freq_sweep | 11 | 1 | 4 | 1 | 120 | 1 | 0.999 | 60 | 120.0182 | 0.99998 | 99.9007 | 0.1199 | -0.005339 | 0.12002 | 119.765 | 1 | 0.11984 | -0.2532 | 1.7e-05 | -6.1e-05 | -0.211 | 0.0017 | -0.0509
freq_sweep | 11 | 1 | 5 | 1 | 120 | 1 | 0.999 | 60 | 120.0182 | 0.99998 | 99.9007 | 0.1199 | -0.005339 | 0.12002 | 119.765 | 1.001 | 0.11994 | -0.2532 | 0.001017 | 4.8e-05 | -0.211 | 0.1017 | 0.04
freq_sweep | 11 | 1 | 6 | 1 | 120 | 1 | 0.999 | 60 | 120.0182 | 0.99998 | 99.9007 | 0.1199 | -0.005339 | 0.12002 | 119.765 | 0.999 | 0.11966 | -0.2532 | -0.000983 | -0.000237 | -0.211 | -0.0983 | -0.1977
freq_sweep | 11 | 1 | 7 | 2 | 120 | 1 | 0.999 | 60 | 119.9814 | 1.0002 | 99.9038 | 0.11989 | -0.0052618 | 0.12 | 119.766 | 2.978 | 0.35675 | -0.21543 | 1.9778 | 0.23687 | -0.1796 | -0.7499 | -0.808
freq_sweep | 11 | 1 | 8 | 2 | 120 | 1 | 0.999 | 60 | 119.9814 | 1.0002 | 99.9038 | 0.11989 | -0.0052618 | 0.12 | 119.766 | 1 | 0.11983 | -0.21543 | -0.000167 | -5.3e-05 | -0.1796 | -0.0167 | -0.0442
freq_sweep | 11 | 1 | 9 | 2 | 120 | 1 | 0.999 | 60 | 119.9814 | 1.0002 | 99.9038 | 0.11989 | -0.0052618 | 0.12 | 119.766 | 1.001 | 0.11989 | -0.21543 | 0.000833 | 7e-06 | -0.1796 | 0.0833 | 0.0058
freq_sweep | 11 | 1 | 10 | 2 | 120 | 1 | 0.999 | 60 | 119.9814 | 1.0002 | 99.9038 | 0.11989 | -0.0052618 | 0.12 | 119.766 | 1 | 0.11985 | -0.21543 | -0.000167 | -3.8e-05 | -0.1796 | -0.0167 | -0.0317
freq_sweep | 11 | 1 | 11 | 2 | 120 | 1 | 0.999 | 60 | 119.9814 | 1.0002 | 99.9038 | 0.11989 | -0.0052618 | 0.12 | 119.766 | 1 | 0.11981 | -0.21543 | -0.000167 | -7.9e-05 | -0.1796 | -0.0167 | -0.0659
freq_sweep | 11 | 1 | 12 | 2 | 120 | 1 | 0.999 | 60 | 119.9814 | 1.0002 | 99.9038 | 0.11989 | -0.0052618 | 0.12 | 119.766 | 1 | 0.11983 | -0.21543 | -0.000167 | -6.1e-05 | -0.1796 | -0.0167 | -0.0509
freq_sweep | 11 | 1 | 13 | 3 | 120 | 1 | 0.999 | 60 | 119.9913 | 0.99992 | 99.9006 | 0.11986 | -0.0053486 | 0.11998 | 119.768 | 1 | 0.11981 | -0.22331 | 7.9e-05 | -5.3e-05 | -0.1861 | 0.0079 | -0.0446
freq_sweep | 11 | 1 | 14 | 3 | 120 | 1 | 0.999 | 60 | 119.9913 | 0.99992 | 99.9006 | 0.11986 | -0.0053486 | 0.11998 | 119.768 | 0.999 | 0.11973 | -0.22331 | -0.000921 | -0.000135 | -0.1861 | -0.0921 | -0.113
freq_sweep | 11 | 1 | 15 | 3 | 120 | 1 | 0.999 | 60 | 119.9913 | 0.99992 | 99.9006 | 0.11986 | -0.0053486 | 0.11998 | 119.768 | 1 | 0.11989 | -0.22331 | 7.9e-05 | 2.5e-05 | -0.1861 | 0.0079 | 0.0204
freq_sweep | 11 | 1 | 16 | 3 | 120 | 1 | 0.999 | 60 | 119.9913 | 0.99992 | 99.9006 | 0.11986 | -0.0053486 | 0.11998 | 119.768 | 0.999 | 0.11976 | -0.22331 | -0.000921 | -0.000106 | -0.1861 | -0.0921 | -0.0888
freq_sweep | 11 | 1 | 17 | 3 | 120 | 1 | 0.999 | 60 | 119.9913 | 0.99992 | 99.9006 | 0.11986 | -0.0053486 | 0.11998 | 119.768 | 1 | 0.11979 | -0.22331 | 7.9e-05 | -7.4e-05 | -0.1861 | 0.0079 | -0.0621
freq_sweep | 11 | 1 | 18 | 3 | 120 | 1 | 0.999 | 60 | 119.9913 | 0.99992 | 99.9006 | 0.11986 | -0.0053486 | 0.11998 | 119.768 | 1 | 0.11977 | -0.22331 | 7.9e-05 | -9.2e-05 | -0.1861 | 0.0079 | -0.0772
freq_sweep | 11 | 2 | 1 | 1 | 120 | 1 | 0.999 | 60 | 120.0044 | 0.99998 | 99.9006 | 0.11988 | -0.0053474 | 0.12 | 119.754 | 2.968 | 0.3553 | -0.25044 | 1.968 | 0.23541 | -0.2087 | -1.0651 | -1.2108
freq_sweep | 11 | 2 | 2 | 1 | 120 | 1 | 0.999 | 60 | 120.0044 | 0.99998 | 99.9006 | 0.11988 | -0.0053474 | 0.12 | 119.754 | 1.001 | 0.11996 | -0.25044 | 0.001016 | 7.3e-05 | -0.2087 | 0.1016 | 0.0607
freq_sweep | 11 | 2 | 3 | 1 | 120 | 1 | 0.999 | 60 | 120.0044 | 0.99998 | 99.9006 | 0.11988 | -0.0053474 | 0.12 | 119.754 | 1.001 | 0.11995 | -0.25044 | 0.001016 | 6.7e-05 | -0.2087 | 0.1016 | 0.0557
freq_sweep | 11 | 2 | 4 | 1 | 120 | 1 | 0.999 | 60 | 120.0044 | 0.99998 | 99.9006 | 0.11988 | -0.0053474 | 0.12 | 119.754 | 1 | 0.1198 | -0.25044 | 1.6e-05 | -8.4e-05 | -0.2087 | 0.0016 | -0.0702
freq_sweep | 11 | 2 | 5 | 1 | 120 | 1 | 0.999 | 60 | 120.0044 | 0.99998 | 99.9006 | 0.11988 | -0.0053474 | 0.12 | 119.754 | 1.001 | 0.11993 | -0.25044 | 0.001016 | 4.8e-05 | -0.2087 | 0.1016 | 0.0399
freq_sweep | 11 | 2 | 6 | 1 | 120 | 1 | 0.999 | 60 | 120.0044 | 0.99998 | 99.9006 | 0.11988 | -0.0053474 | 0.12 | 119.754 | 0.999 | 0.11963 | -0.25044 | -0.000984 | -0.000249 | -0.2087 | -0.0984 | -0.2079
freq_sweep | 11 | 2 | 7 | 2 | 120 | 1 | 0.999 | 60 | 119.9838 | 1.0002 | 99.9037 | 0.11989 | -0.0052634 | 0.12 | 119.758 | 2.979 | 0.35683 | -0.22583 | 1.9788 | 0.23694 | -0.1882 | -0.71585 | -0.78778
freq_sweep | 11 | 2 | 8 | 2 | 120 | 1 | 0.999 | 60 | 119.9838 | 1.0002 | 99.9037 | 0.11989 | -0.0052634 | 0.12 | 119.758 | 1 | 0.11981 | -0.22583 | -0.00016 | -7.2e-05 | -0.1882 | -0.016 | -0.0604
freq_sweep | 11 | 2 | 9 | 2 | 120 | 1 | 0.999 | 60 | 119.9838 | 1.0002 | 99.9037 | 0.11989 | -0.0052634 | 0.12 | 119.758 | 1.001 | 0.11989 | -0.22583 | 0.00084 | 1e-06 | -0.1882 | 0.084 | 0.0005
freq_sweep | 11 | 2 | 10 | 2 | 120 | 1 | 0.999 | 60 | 119.9838 | 1.0002 | 99.9037 | 0.11989 | -0.0052634 | 0.12 | 119.758 | 1 | 0.11987 | -0.22583 | -0.00016 | -1.9e-05 | -0.1882 | -0.016 | -0.0162
freq_sweep | 11 | 2 | 11 | 2 | 120 | 1 | 0.999 | 60 | 119.9838 | 1.0002 | 99.9037 | 0.11989 | -0.0052634 | 0.12 | 119.758 | 1 | 0.11978 | -0.22583 | -0.00016 | -0.000105 | -0.1882 | -0.016 | -0.088
freq_sweep | 11 | 2 | 12 | 2 | 120 | 1 | 0.999 | 60 | 119.9838 | 1.0002 | 99.9037 | 0.11989 | -0.0052634 | 0.12 | 119.758 | 1 | 0.11979 | -0.22583 | -0.00016 | -9.4e-05 | -0.1882 | -0.016 | -0.0788
freq_sweep | 11 | 2 | 13 | 3 | 120 | 1 | 0.999 | 60 | 119.9882 | 0.99992 | 99.9006 | 0.11986 | -0.0053465 | 0.11998 | 119.78 | 1 | 0.11977 | -0.20817 | 7.9e-05 | -8.9e-05 | -0.1735 | 0.0079 | -0.0747
freq_sweep | 11 | 2 | 14 | 3 | 120 | 1 | 0.999 | 60 | 119.9882 | 0.99992 | 99.9006 | 0.11986 | -0.0053465 | 0.11998 | 119.78 | 0.999 | 0.11972 | -0.20817 | -0.000921 | -0.000137 | -0.1735 | -0.0921 | -0.1147
freq_sweep | 11 | 2 | 15 | 3 | 120 | 1 | 0.999 | 60 | 119.9882 | 0.99992 | 99.9006 | 0.11986 | -0.0053465 | 0.11998 | 119.78 | 1.001 | 0.1199 | -0.20817 | 0.001079 | 4.5e-05 | -0.1735 | 0.1079 | 0.0371
freq_sweep | 11 | 2 | 16 | 3 | 120 | 1 | 0.999 | 60 | 119.9882 | 0.99992 | 99.9006 | 0.11986 | -0.0053465 | 0.11998 | 119.78 | 0.999 | 0.11974 | -0.20817 | -0.000921 | -0.000122 | -0.1735 | -0.0921 | -0.1022
freq_sweep | 11 | 2 | 17 | 3 | 120 | 1 | 0.999 | 60 | 119.9882 | 0.99992 | 99.9006 | 0.11986 | -0.0053465 | 0.11998 | 119.78 | 1 | 0.11979 | -0.20817 | 7.9e-05 | -7.4e-05 | -0.1735 | 0.0079 | -0.0621
freq_sweep | 11 | 2 | 18 | 3 | 120 | 1 | 0.999 | 60 | 119.9882 | 0.99992 | 99.9006 | 0.11986 | -0.0053465 | 0.11998 | 119.78 | 1 | 0.11978 | -0.20817 | 7.9e-05 | -7.7e-05 | -0.1735 | 0.0079 | -0.0646
freq_sweep | 11 | 3 | 1 | 1 | 120 | 1 | 0.999 | 60 | 119.9991 | 0.99997 | 99.9005 | 0.11988 | -0.005347 | 0.12 | 119.725 | 2.968 | 0.35541 | -0.27411 | 1.968 | 0.23554 | -0.2284 | -1.0636 | -1.1717
freq_sweep | 11 | 3 | 2 | 1 | 120 | 1 | 0.999 | 60 | 119.9991 | 0.99997 | 99.9005 | 0.11988 | -0.005347 | 0.12 | 119.725 | 1.001 | 0.11992 | -0.27411 | 0.001031 | 4.3e-05 | -0.2284 | 0.1031 | 0.0359
freq_sweep | 11 | 3 | 3 | 1 | 120 | 1 | 0.999 | 60 | 119.9991 | 0.99997 | 99.9005 | 0.11988 | -0.005347 | 0.12 | 119.725 | 1.001 | 0.11993 | -0.27411 | 0.001031 | 5.1e-05 | -0.2284 | 0.1031 | 0.0426
freq_sweep | 11 | 3 | 4 | 1 | 120 | 1 | 0.999 | 60 | 119.9991 | 0.99997 | 99.9005 | 0.11988 | -0.005347 | 0.12 | 119.725 | 1 | 0.11984 | -0.27411 | 3.1e-05 | -3.8e-05 | -0.2284 | 0.0031 | -0.0316
freq_sweep | 11 | 3 | 5 | 1 | 120 | 1 | 0.999 | 60 | 119.9991 | 0.99997 | 99.9005 | 0.11988 | -0.005347 | 0.12 | 119.725 | 1.001 | 0.11991 | -0.27411 | 0.001031 | 3.5e-05 | -0.2284 | 0.1031 | 0.0293
freq_sweep | 11 | 3 | 6 | 1 | 120 | 1 | 0.999 | 60 | 119.9991 | 0.99997 | 99.9005 | 0.11988 | -0.005347 | 0.12 | 119.725 | 0.998 | 0.11958 | -0.27411 | -0.001969 | -0.000291 | -0.2284 | -0.1969 | -0.2427
freq_sweep | 11 | 3 | 7 | 2 | 120 | 1 | 0.999 | 60 | 119.9825 | 1.0002 | 99.9036 | 0.11989 | -0.0052665 | 0.12 | 119.756 | 2.978 | 0.35673 | -0.22648 | 1.9778 | 0.23684 | -0.1888 | -0.74925 | -0.81472
freq_sweep | 11 | 3 | 8 | 2 | 120 | 1 | 0.999 | 60 | 119.9825 | 1.0002 | 99.9036 | 0.11989 | -0.0052665 | 0.12 | 119.756 | 1 | 0.11983 | -0.22648 | -0.00016 | -5.5e-05 | -0.1888 | -0.016 | -0.0459
freq_sweep | 11 | 3 | 9 | 2 | 120 | 1 | 0.999 | 60 | 119.9825 | 1.0002 | 99.9036 | 0.11989 | -0.0052665 | 0.12 | 119.756 | 1.001 | 0.11989 | -0.22648 | 0.00084 | 8e-06 | -0.1888 | 0.084 | 0.0066
freq_sweep | 11 | 3 | 10 | 2 | 120 | 1 | 0.999 | 60 | 119.9825 | 1.0002 | 99.9036 | 0.11989 | -0.0052665 | 0.12 | 119.756 | 1 | 0.11982 | -0.22648 | -0.00016 | -6.8e-05 | -0.1888 | -0.016 | -0.0568
freq_sweep | 11 | 3 | 11 | 2 | 120 | 1 | 0.999 | 60 | 119.9825 | 1.0002 | 99.9036 | 0.11989 | -0.0052665 | 0.12 | 119.756 | 1 | 0.1198 | -0.22648 | -0.00016 | -8.7e-05 | -0.1888 | -0.016 | -0.0726
freq_sweep | 11 | 3 | 12 | 2 | 120 | 1 | 0.999 | 60 | 119.9825 | 1.0002 | 99.9036 | 0.11989 | -0.0052665 | 0.12 | 119.756 | 1 | 0.11983 | -0.22648 | -0.00016 | -6e-05 | -0.1888 | -0.016 | -0.0501
freq_sweep | 11 | 3 | 13 | 3 | 120 | 1 | 0.999 | 60 | 119.9896 | 0.99992 | 99.9007 | 0.11986 | -0.0053472 | 0.11998 | 119.774 | 1 | 0.1198 | -0.21562 | 8.2e-05 | -6.5e-05 | -0.1797 | 0.0082 | -0.0539
freq_sweep | 11 | 3 | 14 | 3 | 120 | 1 | 0.999 | 60 | 119.9896 | 0.99992 | 99.9007 | 0.11986 | -0.0053472 | 0.11998 | 119.774 | 0.999 | 0.11975 | -0.21562 | -0.000918 | -0.000116 | -0.1797 | -0.0918 | -0.0965
freq_sweep | 11 | 3 | 15 | 3 | 120 | 1 | 0.999 | 60 | 119.9896 | 0.99992 | 99.9007 | 0.11986 | -0.0053472 | 0.11998 | 119.774 | 1 | 0.11987 | -0.21562 | 8.2e-05 | 1.1e-05 | -0.1797 | 0.0082 | 0.0095
freq_sweep | 11 | 3 | 16 | 3 | 120 | 1 | 0.999 | 60 | 119.9896 | 0.99992 | 99.9007 | 0.11986 | -0.0053472 | 0.11998 | 119.774 | 0.999 | 0.11974 | -0.21562 | -0.000918 | -0.000117 | -0.1797 | -0.0918 | -0.0973
freq_sweep | 11 | 3 | 17 | 3 | 120 | 1 | 0.999 | 60 | 119.9896 | 0.99992 | 99.9007 | 0.11986 | -0.0053472 | 0.11998 | 119.774 | 1 | 0.11981 | -0.21562 | 8.2e-05 | -5.4e-05 | -0.1797 | 0.0082 | -0.0448
freq_sweep | 11 | 3 | 18 | 3 | 120 | 1 | 0.999 | 60 | 119.9896 | 0.99992 | 99.9007 | 0.11986 | -0.0053472 | 0.11998 | 119.774 | 1 | 0.11978 | -0.21562 | 8.2e-05 | -8.5e-05 | -0.1797 | 0.0082 | -0.0706

## PF_Sweep_Investigate
sweep | point | sample | channel | phase | source_vln | source_i | source_pf | source_freq | ref_vln | ref_i | ref_pf | ref_kW | ref_kVAR | ref_kVA | dut_vln | dut_i | dut_kW | err_vln | err_i | err_kW | pct_err_vln | pct_err_i | pct_err_kW
pf_sweep | 1 | 1 | 1 | 1 | 120 | 1 | -0.2 | 60 | 120.0073 | 0.99998 | -19.9831 | 0.023981 | 0.11758 | 0.12001 | 119.745 | 2.97 | 0.058702 | -0.26234 | 1.97 | 0.034721 | -0.2186 | -0.99851 | -18.4043
pf_sweep | 1 | 1 | 2 | 1 | 120 | 1 | -0.2 | 60 | 120.0073 | 0.99998 | -19.9831 | 0.023981 | 0.11758 | 0.12001 | 119.745 | 1.001 | 0.020427 | -0.26234 | 0.001015 | -0.003554 | -0.2186 | 0.1015 | -14.8195
pf_sweep | 1 | 1 | 3 | 1 | 120 | 1 | -0.2 | 60 | 120.0073 | 0.99998 | -19.9831 | 0.023981 | 0.11758 | 0.12001 | 119.745 | 1.001 | 0.020444 | -0.26234 | 0.001015 | -0.003537 | -0.2186 | 0.1015 | -14.7486
pf_sweep | 1 | 1 | 4 | 1 | 120 | 1 | -0.2 | 60 | 120.0073 | 0.99998 | -19.9831 | 0.023981 | 0.11758 | 0.12001 | 119.745 | 1.001 | 0.02044 | -0.26234 | 0.001015 | -0.003541 | -0.2186 | 0.1015 | -14.7653
pf_sweep | 1 | 1 | 5 | 1 | 120 | 1 | -0.2 | 60 | 120.0073 | 0.99998 | -19.9831 | 0.023981 | 0.11758 | 0.12001 | 119.745 | 1.001 | 0.020529 | -0.26234 | 0.001015 | -0.003452 | -0.2186 | 0.1015 | -14.3941
pf_sweep | 1 | 1 | 6 | 1 | 120 | 1 | -0.2 | 60 | 120.0073 | 0.99998 | -19.9831 | 0.023981 | 0.11758 | 0.12001 | 119.745 | 0.999 | 0.020945 | -0.26234 | -0.000985 | -0.003036 | -0.2186 | -0.0985 | -12.6594
pf_sweep | 1 | 1 | 7 | 2 | 120 | 1 | -0.2 | 60 | 119.9906 | 1.0002 | -19.9496 | 0.023942 | 0.1176 | 0.12001 | 119.769 | 2.98 | 0.058551 | -0.22156 | 1.9798 | 0.034609 | -0.1846 | -0.68321 | -18.4809
pf_sweep | 1 | 1 | 8 | 2 | 120 | 1 | -0.2 | 60 | 119.9906 | 1.0002 | -19.9496 | 0.023942 | 0.1176 | 0.12001 | 119.769 | 1 | 0.020365 | -0.22156 | -0.000167 | -0.003577 | -0.1846 | -0.0167 | -14.9389
pf_sweep | 1 | 1 | 9 | 2 | 120 | 1 | -0.2 | 60 | 119.9906 | 1.0002 | -19.9496 | 0.023942 | 0.1176 | 0.12001 | 119.769 | 1.001 | 0.020517 | -0.22156 | 0.000833 | -0.003425 | -0.1846 | 0.0833 | -14.3041
pf_sweep | 1 | 1 | 10 | 2 | 120 | 1 | -0.2 | 60 | 119.9906 | 1.0002 | -19.9496 | 0.023942 | 0.1176 | 0.12001 | 119.769 | 1 | 0.020467 | -0.22156 | -0.000167 | -0.003475 | -0.1846 | -0.0167 | -14.5129
pf_sweep | 1 | 1 | 11 | 2 | 120 | 1 | -0.2 | 60 | 119.9906 | 1.0002 | -19.9496 | 0.023942 | 0.1176 | 0.12001 | 119.769 | 1 | 0.020729 | -0.22156 | -0.000167 | -0.003213 | -0.1846 | -0.0167 | -13.4186
pf_sweep | 1 | 1 | 12 | 2 | 120 | 1 | -0.2 | 60 | 119.9906 | 1.0002 | -19.9496 | 0.023942 | 0.1176 | 0.12001 | 119.769 | 1 | 0.020355 | -0.22156 | -0.000167 | -0.003587 | -0.1846 | -0.0167 | -14.9807
pf_sweep | 1 | 1 | 13 | 3 | 120 | 1 | -0.2 | 60 | 119.9926 | 1 | -19.9864 | 0.023982 | 0.11757 | 0.11999 | 119.784 | 1 | 0.020626 | -0.20864 | 4e-06 | -0.003356 | -0.1739 | 0.0004 | -13.9941
pf_sweep | 1 | 1 | 14 | 3 | 120 | 1 | -0.2 | 60 | 119.9926 | 1 | -19.9864 | 0.023982 | 0.11757 | 0.11999 | 119.784 | 0.999 | 0.020828 | -0.20864 | -0.000996 | -0.003154 | -0.1739 | -0.0996 | -13.1518
pf_sweep | 1 | 1 | 15 | 3 | 120 | 1 | -0.2 | 60 | 119.9926 | 1 | -19.9864 | 0.023982 | 0.11757 | 0.11999 | 119.784 | 1.001 | 0.020681 | -0.20864 | 0.001004 | -0.003301 | -0.1739 | 0.1004 | -13.7648
pf_sweep | 1 | 1 | 16 | 3 | 120 | 1 | -0.2 | 60 | 119.9926 | 1 | -19.9864 | 0.023982 | 0.11757 | 0.11999 | 119.784 | 0.999 | 0.020902 | -0.20864 | -0.000996 | -0.00308 | -0.1739 | -0.0996 | -12.8433
pf_sweep | 1 | 1 | 17 | 3 | 120 | 1 | -0.2 | 60 | 119.9926 | 1 | -19.9864 | 0.023982 | 0.11757 | 0.11999 | 119.784 | 1 | 0.020726 | -0.20864 | 4e-06 | -0.003256 | -0.1739 | 0.0004 | -13.5772
pf_sweep | 1 | 1 | 18 | 3 | 120 | 1 | -0.2 | 60 | 119.9926 | 1 | -19.9864 | 0.023982 | 0.11757 | 0.11999 | 119.784 | 1 | 0.020781 | -0.20864 | 4e-06 | -0.003201 | -0.1739 | 0.0004 | -13.3478
pf_sweep | 1 | 2 | 1 | 1 | 120 | 1 | -0.2 | 60 | 120.0168 | 0.99998 | -19.9799 | 0.023979 | 0.11759 | 0.12001 | 119.75 | 2.97 | 0.05863 | -0.26675 | 1.97 | 0.034651 | -0.2223 | -0.99829 | -18.4973
pf_sweep | 1 | 2 | 2 | 1 | 120 | 1 | -0.2 | 60 | 120.0168 | 0.99998 | -19.9799 | 0.023979 | 0.11759 | 0.12001 | 119.75 | 1.001 | 0.020481 | -0.26675 | 0.001017 | -0.003498 | -0.2223 | 0.1017 | -14.5869
pf_sweep | 1 | 2 | 3 | 1 | 120 | 1 | -0.2 | 60 | 120.0168 | 0.99998 | -19.9799 | 0.023979 | 0.11759 | 0.12001 | 119.75 | 1.001 | 0.02042 | -0.26675 | 0.001017 | -0.003559 | -0.2223 | 0.1017 | -14.8413
pf_sweep | 1 | 2 | 4 | 1 | 120 | 1 | -0.2 | 60 | 120.0168 | 0.99998 | -19.9799 | 0.023979 | 0.11759 | 0.12001 | 119.75 | 1.001 | 0.020412 | -0.26675 | 0.001017 | -0.003567 | -0.2223 | 0.1017 | -14.8747
pf_sweep | 1 | 2 | 5 | 1 | 120 | 1 | -0.2 | 60 | 120.0168 | 0.99998 | -19.9799 | 0.023979 | 0.11759 | 0.12001 | 119.75 | 1.001 | 0.020544 | -0.26675 | 0.001017 | -0.003435 | -0.2223 | 0.1017 | -14.3242
pf_sweep | 1 | 2 | 6 | 1 | 120 | 1 | -0.2 | 60 | 120.0168 | 0.99998 | -19.9799 | 0.023979 | 0.11759 | 0.12001 | 119.75 | 0.999 | 0.020931 | -0.26675 | -0.000983 | -0.003048 | -0.2223 | -0.0983 | -12.7103
pf_sweep | 1 | 2 | 7 | 2 | 120 | 1 | -0.2 | 60 | 119.9886 | 1.0002 | -19.9485 | 0.02394 | 0.1176 | 0.12001 | 119.77 | 2.979 | 0.058631 | -0.21863 | 1.9788 | 0.034691 | -0.1822 | -0.71737 | -18.3644
pf_sweep | 1 | 2 | 8 | 2 | 120 | 1 | -0.2 | 60 | 119.9886 | 1.0002 | -19.9485 | 0.02394 | 0.1176 | 0.12001 | 119.77 | 1 | 0.020353 | -0.21863 | -0.000175 | -0.003587 | -0.1822 | -0.0175 | -14.9838
pf_sweep | 1 | 2 | 9 | 2 | 120 | 1 | -0.2 | 60 | 119.9886 | 1.0002 | -19.9485 | 0.02394 | 0.1176 | 0.12001 | 119.77 | 1.001 | 0.020518 | -0.21863 | 0.000825 | -0.003422 | -0.1822 | 0.0825 | -14.2946
pf_sweep | 1 | 2 | 10 | 2 | 120 | 1 | -0.2 | 60 | 119.9886 | 1.0002 | -19.9485 | 0.02394 | 0.1176 | 0.12001 | 119.77 | 1 | 0.020498 | -0.21863 | -0.000175 | -0.003442 | -0.1822 | -0.0175 | -14.3781
pf_sweep | 1 | 2 | 11 | 2 | 120 | 1 | -0.2 | 60 | 119.9886 | 1.0002 | -19.9485 | 0.02394 | 0.1176 | 0.12001 | 119.77 | 1 | 0.020727 | -0.21863 | -0.000175 | -0.003213 | -0.1822 | -0.0175 | -13.4216
pf_sweep | 1 | 2 | 12 | 2 | 120 | 1 | -0.2 | 60 | 119.9886 | 1.0002 | -19.9485 | 0.02394 | 0.1176 | 0.12001 | 119.77 | 1 | 0.020346 | -0.21863 | -0.000175 | -0.003594 | -0.1822 | -0.0175 | -15.013
pf_sweep | 1 | 2 | 13 | 3 | 120 | 1 | -0.2 | 60 | 119.9925 | 1 | -19.9893 | 0.023986 | 0.11757 | 0.11999 | 119.773 | 1 | 0.020633 | -0.21953 | 0 | -0.003353 | -0.183 | 0 | -13.9776
pf_sweep | 1 | 2 | 14 | 3 | 120 | 1 | -0.2 | 60 | 119.9925 | 1 | -19.9893 | 0.023986 | 0.11757 | 0.11999 | 119.773 | 0.999 | 0.02081 | -0.21953 | -0.001 | -0.003176 | -0.183 | -0.1 | -13.2397
pf_sweep | 1 | 2 | 15 | 3 | 120 | 1 | -0.2 | 60 | 119.9925 | 1 | -19.9893 | 0.023986 | 0.11757 | 0.11999 | 119.773 | 1.001 | 0.020732 | -0.21953 | 0.001 | -0.003254 | -0.183 | 0.1 | -13.5649
pf_sweep | 1 | 2 | 16 | 3 | 120 | 1 | -0.2 | 60 | 119.9925 | 1 | -19.9893 | 0.023986 | 0.11757 | 0.11999 | 119.773 | 0.999 | 0.0209 | -0.21953 | -0.001 | -0.003086 | -0.183 | -0.1 | -12.8645
pf_sweep | 1 | 2 | 17 | 3 | 120 | 1 | -0.2 | 60 | 119.9925 | 1 | -19.9893 | 0.023986 | 0.11757 | 0.11999 | 119.773 | 1 | 0.020706 | -0.21953 | 0 | -0.00328 | -0.183 | 0 | -13.6733
pf_sweep | 1 | 2 | 18 | 3 | 120 | 1 | -0.2 | 60 | 119.9925 | 1 | -19.9893 | 0.023986 | 0.11757 | 0.11999 | 119.773 | 0.999 | 0.020822 | -0.21953 | -0.001 | -0.003164 | -0.183 | -0.1 | -13.1897
pf_sweep | 1 | 3 | 1 | 1 | 120 | 1 | -0.2 | 60 | 120.0246 | 1 | -19.9836 | 0.023985 | 0.1176 | 0.12003 | 119.765 | 2.97 | 0.058676 | -0.25963 | 1.97 | 0.034691 | -0.2163 | -1.0009 | -18.4562
pf_sweep | 1 | 3 | 2 | 1 | 120 | 1 | -0.2 | 60 | 120.0246 | 1 | -19.9836 | 0.023985 | 0.1176 | 0.12003 | 119.765 | 1.001 | 0.020474 | -0.25963 | 0.000991 | -0.003511 | -0.2163 | 0.0991 | -14.64
pf_sweep | 1 | 3 | 3 | 1 | 120 | 1 | -0.2 | 60 | 120.0246 | 1 | -19.9836 | 0.023985 | 0.1176 | 0.12003 | 119.765 | 1.001 | 0.020458 | -0.25963 | 0.000991 | -0.003527 | -0.2163 | 0.0991 | -14.7067
pf_sweep | 1 | 3 | 4 | 1 | 120 | 1 | -0.2 | 60 | 120.0246 | 1 | -19.9836 | 0.023985 | 0.1176 | 0.12003 | 119.765 | 1.001 | 0.020421 | -0.25963 | 0.000991 | -0.003564 | -0.2163 | 0.0991 | -14.861
pf_sweep | 1 | 3 | 5 | 1 | 120 | 1 | -0.2 | 60 | 120.0246 | 1 | -19.9836 | 0.023985 | 0.1176 | 0.12003 | 119.765 | 1.001 | 0.020546 | -0.25963 | 0.000991 | -0.003439 | -0.2163 | 0.0991 | -14.3399
pf_sweep | 1 | 3 | 6 | 1 | 120 | 1 | -0.2 | 60 | 120.0246 | 1 | -19.9836 | 0.023985 | 0.1176 | 0.12003 | 119.765 | 0.998 | 0.02093 | -0.25963 | -0.002009 | -0.003055 | -0.2163 | -0.2009 | -12.7389
pf_sweep | 1 | 3 | 7 | 2 | 120 | 1 | -0.2 | 60 | 119.99 | 1.0002 | -19.9473 | 0.023939 | 0.1176 | 0.12001 | 119.769 | 2.98 | 0.058649 | -0.22095 | 1.9798 | 0.03471 | -0.1841 | -0.68464 | -18.3358
pf_sweep | 1 | 3 | 8 | 2 | 120 | 1 | -0.2 | 60 | 119.99 | 1.0002 | -19.9473 | 0.023939 | 0.1176 | 0.12001 | 119.769 | 1 | 0.020363 | -0.22095 | -0.000181 | -0.003576 | -0.1841 | -0.0181 | -14.9383
pf_sweep | 1 | 3 | 9 | 2 | 120 | 1 | -0.2 | 60 | 119.99 | 1.0002 | -19.9473 | 0.023939 | 0.1176 | 0.12001 | 119.769 | 1.001 | 0.020525 | -0.22095 | 0.000819 | -0.003414 | -0.1841 | 0.0819 | -14.2616
pf_sweep | 1 | 3 | 10 | 2 | 120 | 1 | -0.2 | 60 | 119.99 | 1.0002 | -19.9473 | 0.023939 | 0.1176 | 0.12001 | 119.769 | 1 | 0.020486 | -0.22095 | -0.000181 | -0.003453 | -0.1841 | -0.0181 | -14.4245
pf_sweep | 1 | 3 | 11 | 2 | 120 | 1 | -0.2 | 60 | 119.99 | 1.0002 | -19.9473 | 0.023939 | 0.1176 | 0.12001 | 119.769 | 1 | 0.020707 | -0.22095 | -0.000181 | -0.003232 | -0.1841 | -0.0181 | -13.5013
pf_sweep | 1 | 3 | 12 | 2 | 120 | 1 | -0.2 | 60 | 119.99 | 1.0002 | -19.9473 | 0.023939 | 0.1176 | 0.12001 | 119.769 | 1 | 0.020354 | -0.22095 | -0.000181 | -0.003585 | -0.1841 | -0.0181 | -14.9759
pf_sweep | 1 | 3 | 13 | 3 | 120 | 1 | -0.2 | 60 | 119.9873 | 0.99999 | -19.9868 | 0.023981 | 0.11757 | 0.11999 | 119.782 | 1 | 0.020614 | -0.20533 | 7e-06 | -0.003367 | -0.1711 | 0.0007 | -14.0418
pf_sweep | 1 | 3 | 14 | 3 | 120 | 1 | -0.2 | 60 | 119.9873 | 0.99999 | -19.9868 | 0.023981 | 0.11757 | 0.11999 | 119.782 | 0.999 | 0.020811 | -0.20533 | -0.000993 | -0.00317 | -0.1711 | -0.0993 | -13.2203
pf_sweep | 1 | 3 | 15 | 3 | 120 | 1 | -0.2 | 60 | 119.9873 | 0.99999 | -19.9868 | 0.023981 | 0.11757 | 0.11999 | 119.782 | 1.001 | 0.020682 | -0.20533 | 0.001007 | -0.003299 | -0.1711 | 0.1007 | -13.7582
pf_sweep | 1 | 3 | 16 | 3 | 120 | 1 | -0.2 | 60 | 119.9873 | 0.99999 | -19.9868 | 0.023981 | 0.11757 | 0.11999 | 119.782 | 0.999 | 0.020881 | -0.20533 | -0.000993 | -0.0031 | -0.1711 | -0.0993 | -12.9284
pf_sweep | 1 | 3 | 17 | 3 | 120 | 1 | -0.2 | 60 | 119.9873 | 0.99999 | -19.9868 | 0.023981 | 0.11757 | 0.11999 | 119.782 | 1 | 0.020737 | -0.20533 | 7e-06 | -0.003244 | -0.1711 | 0.0007 | -13.5289
pf_sweep | 1 | 3 | 18 | 3 | 120 | 1 | -0.2 | 60 | 119.9873 | 0.99999 | -19.9868 | 0.023981 | 0.11757 | 0.11999 | 119.782 | 1 | 0.020782 | -0.20533 | 7e-06 | -0.003199 | -0.1711 | 0.0007 | -13.3412
pf_sweep | 2 | 1 | 1 | 1 | 120 | 1 | -0.178 | 60 | 120.0072 | 1 | -17.79 | 0.02135 | 0.1181 | 0.12001 | 119.686 | 2.97 | 0.050762 | -0.32124 | 1.97 | 0.029412 | -0.2677 | -1.0027 | -20.7458
pf_sweep | 2 | 1 | 2 | 1 | 120 | 1 | -0.178 | 60 | 120.0072 | 1 | -17.79 | 0.02135 | 0.1181 | 0.12001 | 119.686 | 1.001 | 0.017821 | -0.32124 | 0.000973 | -0.003529 | -0.2677 | 0.0973 | -16.5287
pf_sweep | 2 | 1 | 3 | 1 | 120 | 1 | -0.178 | 60 | 120.0072 | 1 | -17.79 | 0.02135 | 0.1181 | 0.12001 | 119.686 | 1.001 | 0.017802 | -0.32124 | 0.000973 | -0.003548 | -0.2677 | 0.0973 | -16.6177
pf_sweep | 2 | 1 | 4 | 1 | 120 | 1 | -0.178 | 60 | 120.0072 | 1 | -17.79 | 0.02135 | 0.1181 | 0.12001 | 119.686 | 1.001 | 0.017746 | -0.32124 | 0.000973 | -0.003604 | -0.2677 | 0.0973 | -16.88
pf_sweep | 2 | 1 | 5 | 1 | 120 | 1 | -0.178 | 60 | 120.0072 | 1 | -17.79 | 0.02135 | 0.1181 | 0.12001 | 119.686 | 1.001 | 0.01785 | -0.32124 | 0.000973 | -0.0035 | -0.2677 | 0.0973 | -16.3929
pf_sweep | 2 | 1 | 6 | 1 | 120 | 1 | -0.178 | 60 | 120.0072 | 1 | -17.79 | 0.02135 | 0.1181 | 0.12001 | 119.686 | 0.999 | 0.018273 | -0.32124 | -0.001027 | -0.003077 | -0.2677 | -0.1027 | -14.4116
pf_sweep | 2 | 1 | 7 | 2 | 120 | 1 | -0.178 | 60 | 119.9914 | 1.0002 | -17.746 | 0.021297 | 0.11811 | 0.12001 | 119.768 | 2.979 | 0.050625 | -0.22339 | 1.9788 | 0.029328 | -0.1862 | -0.71661 | -20.7642
pf_sweep | 2 | 1 | 8 | 2 | 120 | 1 | -0.178 | 60 | 119.9914 | 1.0002 | -17.746 | 0.021297 | 0.11811 | 0.12001 | 119.768 | 1.001 | 0.017704 | -0.22339 | 0.000833 | -0.003593 | -0.1862 | 0.0833 | -16.8717
pf_sweep | 2 | 1 | 9 | 2 | 120 | 1 | -0.178 | 60 | 119.9914 | 1.0002 | -17.746 | 0.021297 | 0.11811 | 0.12001 | 119.768 | 1.001 | 0.017854 | -0.22339 | 0.000833 | -0.003443 | -0.1862 | 0.0833 | -16.1674
pf_sweep | 2 | 1 | 10 | 2 | 120 | 1 | -0.178 | 60 | 119.9914 | 1.0002 | -17.746 | 0.021297 | 0.11811 | 0.12001 | 119.768 | 1 | 0.017811 | -0.22339 | -0.000167 | -0.003486 | -0.1862 | -0.0167 | -16.3693
pf_sweep | 2 | 1 | 11 | 2 | 120 | 1 | -0.178 | 60 | 119.9914 | 1.0002 | -17.746 | 0.021297 | 0.11811 | 0.12001 | 119.768 | 1 | 0.018072 | -0.22339 | -0.000167 | -0.003225 | -0.1862 | -0.0167 | -15.1438
pf_sweep | 2 | 1 | 12 | 2 | 120 | 1 | -0.178 | 60 | 119.9914 | 1.0002 | -17.746 | 0.021297 | 0.11811 | 0.12001 | 119.768 | 1 | 0.01771 | -0.22339 | -0.000167 | -0.003587 | -0.1862 | -0.0167 | -16.8435
pf_sweep | 2 | 1 | 13 | 3 | 120 | 1 | -0.178 | 60 | 119.989 | 1 | -17.788 | 0.021344 | 0.11808 | 0.11999 | 119.781 | 1 | 0.017976 | -0.20804 | -2.3e-05 | -0.003368 | -0.1734 | -0.0023 | -15.7799
pf_sweep | 2 | 1 | 14 | 3 | 120 | 1 | -0.178 | 60 | 119.989 | 1 | -17.788 | 0.021344 | 0.11808 | 0.11999 | 119.781 | 0.999 | 0.018178 | -0.20804 | -0.001023 | -0.003166 | -0.1734 | -0.1023 | -14.8335
pf_sweep | 2 | 1 | 15 | 3 | 120 | 1 | -0.178 | 60 | 119.989 | 1 | -17.788 | 0.021344 | 0.11808 | 0.11999 | 119.781 | 1.001 | 0.018037 | -0.20804 | 0.000977 | -0.003307 | -0.1734 | 0.0977 | -15.4941
pf_sweep | 2 | 1 | 16 | 3 | 120 | 1 | -0.178 | 60 | 119.989 | 1 | -17.788 | 0.021344 | 0.11808 | 0.11999 | 119.781 | 0.999 | 0.018256 | -0.20804 | -0.001023 | -0.003088 | -0.1734 | -0.1023 | -14.4681
pf_sweep | 2 | 1 | 17 | 3 | 120 | 1 | -0.178 | 60 | 119.989 | 1 | -17.788 | 0.021344 | 0.11808 | 0.11999 | 119.781 | 1 | 0.018084 | -0.20804 | -2.3e-05 | -0.00326 | -0.1734 | -0.0023 | -15.2739
pf_sweep | 2 | 1 | 18 | 3 | 120 | 1 | -0.178 | 60 | 119.989 | 1 | -17.788 | 0.021344 | 0.11808 | 0.11999 | 119.781 | 1 | 0.018134 | -0.20804 | -2.3e-05 | -0.00321 | -0.1734 | -0.0023 | -15.0397
pf_sweep | 2 | 2 | 1 | 1 | 120 | 1 | -0.178 | 60 | 120.0029 | 1 | -17.784 | 0.021342 | 0.11809 | 0.12 | 119.752 | 2.97 | 0.050845 | -0.25094 | 1.97 | 0.029503 | -0.2091 | -1.0016 | -20.5856
pf_sweep | 2 | 2 | 2 | 1 | 120 | 1 | -0.178 | 60 | 120.0029 | 1 | -17.784 | 0.021342 | 0.11809 | 0.12 | 119.752 | 1.001 | 0.017813 | -0.25094 | 0.000984 | -0.003529 | -0.2091 | 0.0984 | -16.534
pf_sweep | 2 | 2 | 3 | 1 | 120 | 1 | -0.178 | 60 | 120.0029 | 1 | -17.784 | 0.021342 | 0.11809 | 0.12 | 119.752 | 1.001 | 0.017773 | -0.25094 | 0.000984 | -0.003569 | -0.2091 | 0.0984 | -16.7215
pf_sweep | 2 | 2 | 4 | 1 | 120 | 1 | -0.178 | 60 | 120.0029 | 1 | -17.784 | 0.021342 | 0.11809 | 0.12 | 119.752 | 1 | 0.017788 | -0.25094 | -1.6e-05 | -0.003554 | -0.2091 | -0.0016 | -16.6512
pf_sweep | 2 | 2 | 5 | 1 | 120 | 1 | -0.178 | 60 | 120.0029 | 1 | -17.784 | 0.021342 | 0.11809 | 0.12 | 119.752 | 1.001 | 0.017881 | -0.25094 | 0.000984 | -0.003461 | -0.2091 | 0.0984 | -16.2154
pf_sweep | 2 | 2 | 6 | 1 | 120 | 1 | -0.178 | 60 | 120.0029 | 1 | -17.784 | 0.021342 | 0.11809 | 0.12 | 119.752 | 0.999 | 0.018273 | -0.25094 | -0.001016 | -0.003069 | -0.2091 | -0.1016 | -14.3786
pf_sweep | 2 | 2 | 7 | 2 | 120 | 1 | -0.178 | 60 | 119.9902 | 1.0002 | -17.7519 | 0.021304 | 0.1181 | 0.12001 | 119.783 | 2.98 | 0.050642 | -0.20722 | 1.9798 | 0.029338 | -0.1727 | -0.68356 | -20.7634
pf_sweep | 2 | 2 | 8 | 2 | 120 | 1 | -0.178 | 60 | 119.9902 | 1.0002 | -17.7519 | 0.021304 | 0.1181 | 0.12001 | 119.783 | 1 | 0.017675 | -0.20722 | -0.00017 | -0.003629 | -0.1727 | -0.017 | -17.0349
pf_sweep | 2 | 2 | 9 | 2 | 120 | 1 | -0.178 | 60 | 119.9902 | 1.0002 | -17.7519 | 0.021304 | 0.1181 | 0.12001 | 119.783 | 1 | 0.017875 | -0.20722 | -0.00017 | -0.003429 | -0.1727 | -0.017 | -16.0961
pf_sweep | 2 | 2 | 10 | 2 | 120 | 1 | -0.178 | 60 | 119.9902 | 1.0002 | -17.7519 | 0.021304 | 0.1181 | 0.12001 | 119.783 | 1 | 0.017809 | -0.20722 | -0.00017 | -0.003495 | -0.1727 | -0.017 | -16.4059
pf_sweep | 2 | 2 | 11 | 2 | 120 | 1 | -0.178 | 60 | 119.9902 | 1.0002 | -17.7519 | 0.021304 | 0.1181 | 0.12001 | 119.783 | 1 | 0.018053 | -0.20722 | -0.00017 | -0.003251 | -0.1727 | -0.017 | -15.2606
pf_sweep | 2 | 2 | 12 | 2 | 120 | 1 | -0.178 | 60 | 119.9902 | 1.0002 | -17.7519 | 0.021304 | 0.1181 | 0.12001 | 119.783 | 1 | 0.017703 | -0.20722 | -0.00017 | -0.003601 | -0.1727 | -0.017 | -16.9035
pf_sweep | 2 | 2 | 13 | 3 | 120 | 1 | -0.178 | 60 | 119.9947 | 1 | -17.7881 | 0.021345 | 0.11808 | 0.12 | 119.771 | 1 | 0.017949 | -0.22372 | -2.3e-05 | -0.003396 | -0.1864 | -0.0023 | -15.9113
pf_sweep | 2 | 2 | 14 | 3 | 120 | 1 | -0.178 | 60 | 119.9947 | 1 | -17.7881 | 0.021345 | 0.11808 | 0.12 | 119.771 | 0.999 | 0.018161 | -0.22372 | -0.001023 | -0.003184 | -0.1864 | -0.1023 | -14.9181
pf_sweep | 2 | 2 | 15 | 3 | 120 | 1 | -0.178 | 60 | 119.9947 | 1 | -17.7881 | 0.021345 | 0.11808 | 0.12 | 119.771 | 1.001 | 0.018055 | -0.22372 | 0.000977 | -0.00329 | -0.1864 | 0.0977 | -15.4147
pf_sweep | 2 | 2 | 16 | 3 | 120 | 1 | -0.178 | 60 | 119.9947 | 1 | -17.7881 | 0.021345 | 0.11808 | 0.12 | 119.771 | 0.999 | 0.018219 | -0.22372 | -0.001023 | -0.003126 | -0.1864 | -0.1023 | -14.6464
pf_sweep | 2 | 2 | 17 | 3 | 120 | 1 | -0.178 | 60 | 119.9947 | 1 | -17.7881 | 0.021345 | 0.11808 | 0.12 | 119.771 | 1 | 0.018047 | -0.22372 | -2.3e-05 | -0.003298 | -0.1864 | -0.0023 | -15.4522
pf_sweep | 2 | 2 | 18 | 3 | 120 | 1 | -0.178 | 60 | 119.9947 | 1 | -17.7881 | 0.021345 | 0.11808 | 0.12 | 119.771 | 1 | 0.018172 | -0.22372 | -2.3e-05 | -0.003173 | -0.1864 | -0.0023 | -14.8666
pf_sweep | 2 | 3 | 1 | 1 | 120 | 1 | -0.178 | 60 | 119.9755 | 1 | -17.7899 | 0.021344 | 0.11806 | 0.11998 | 119.704 | 2.97 | 0.050843 | -0.27153 | 1.97 | 0.029499 | -0.2263 | -1.0003 | -20.5961
pf_sweep | 2 | 3 | 2 | 1 | 120 | 1 | -0.178 | 60 | 119.9755 | 1 | -17.7899 | 0.021344 | 0.11806 | 0.11998 | 119.704 | 1.001 | 0.017807 | -0.27153 | 0.000997 | -0.003537 | -0.2263 | 0.0997 | -16.5699
pf_sweep | 2 | 3 | 3 | 1 | 120 | 1 | -0.178 | 60 | 119.9755 | 1 | -17.7899 | 0.021344 | 0.11806 | 0.11998 | 119.704 | 1.001 | 0.017769 | -0.27153 | 0.000997 | -0.003575 | -0.2263 | 0.0997 | -16.7479
pf_sweep | 2 | 3 | 4 | 1 | 120 | 1 | -0.178 | 60 | 119.9755 | 1 | -17.7899 | 0.021344 | 0.11806 | 0.11998 | 119.704 | 1 | 0.017769 | -0.27153 | -3e-06 | -0.003575 | -0.2263 | -0.0003 | -16.7479
pf_sweep | 2 | 3 | 5 | 1 | 120 | 1 | -0.178 | 60 | 119.9755 | 1 | -17.7899 | 0.021344 | 0.11806 | 0.11998 | 119.704 | 1.001 | 0.017855 | -0.27153 | 0.000997 | -0.003489 | -0.2263 | 0.0997 | -16.345
pf_sweep | 2 | 3 | 6 | 1 | 120 | 1 | -0.178 | 60 | 119.9755 | 1 | -17.7899 | 0.021344 | 0.11806 | 0.11998 | 119.704 | 0.998 | 0.018269 | -0.27153 | -0.002003 | -0.003075 | -0.2263 | -0.2003 | -14.4053
pf_sweep | 2 | 3 | 7 | 2 | 120 | 1 | -0.178 | 60 | 119.9905 | 1.0002 | -17.7479 | 0.021299 | 0.11811 | 0.12001 | 119.768 | 2.98 | 0.050643 | -0.2225 | 1.9798 | 0.029344 | -0.1854 | -0.68328 | -20.7439
pf_sweep | 2 | 3 | 8 | 2 | 120 | 1 | -0.178 | 60 | 119.9905 | 1.0002 | -17.7479 | 0.021299 | 0.11811 | 0.12001 | 119.768 | 1 | 0.017705 | -0.2225 | -0.000167 | -0.003594 | -0.1854 | -0.0167 | -16.8752
pf_sweep | 2 | 3 | 9 | 2 | 120 | 1 | -0.178 | 60 | 119.9905 | 1.0002 | -17.7479 | 0.021299 | 0.11811 | 0.12001 | 119.768 | 1.001 | 0.017864 | -0.2225 | 0.000833 | -0.003435 | -0.1854 | 0.0833 | -16.1287
pf_sweep | 2 | 3 | 10 | 2 | 120 | 1 | -0.178 | 60 | 119.9905 | 1.0002 | -17.7479 | 0.021299 | 0.11811 | 0.12001 | 119.768 | 1 | 0.017819 | -0.2225 | -0.000167 | -0.00348 | -0.1854 | -0.0167 | -16.34
pf_sweep | 2 | 3 | 11 | 2 | 120 | 1 | -0.178 | 60 | 119.9905 | 1.0002 | -17.7479 | 0.021299 | 0.11811 | 0.12001 | 119.768 | 1 | 0.018074 | -0.2225 | -0.000167 | -0.003225 | -0.1854 | -0.0167 | -15.1427
pf_sweep | 2 | 3 | 12 | 2 | 120 | 1 | -0.178 | 60 | 119.9905 | 1.0002 | -17.7479 | 0.021299 | 0.11811 | 0.12001 | 119.768 | 1 | 0.017721 | -0.2225 | -0.000167 | -0.003578 | -0.1854 | -0.0167 | -16.8001
pf_sweep | 2 | 3 | 13 | 3 | 120 | 1 | -0.178 | 60 | 119.9893 | 1 | -17.7905 | 0.021347 | 0.11808 | 0.11999 | 119.781 | 1 | 0.017939 | -0.20833 | -1.1e-05 | -0.003408 | -0.1736 | -0.0011 | -15.9646
pf_sweep | 2 | 3 | 14 | 3 | 120 | 1 | -0.178 | 60 | 119.9893 | 1 | -17.7905 | 0.021347 | 0.11808 | 0.11999 | 119.781 | 0.999 | 0.018178 | -0.20833 | -0.001011 | -0.003169 | -0.1736 | -0.1011 | -14.845
pf_sweep | 2 | 3 | 15 | 3 | 120 | 1 | -0.178 | 60 | 119.9893 | 1 | -17.7905 | 0.021347 | 0.11808 | 0.11999 | 119.781 | 1.001 | 0.018046 | -0.20833 | 0.000989 | -0.003301 | -0.1736 | 0.0989 | -15.4633
pf_sweep | 2 | 3 | 16 | 3 | 120 | 1 | -0.178 | 60 | 119.9893 | 1 | -17.7905 | 0.021347 | 0.11808 | 0.11999 | 119.781 | 0.999 | 0.018219 | -0.20833 | -0.001011 | -0.003128 | -0.1736 | -0.1011 | -14.6529
pf_sweep | 2 | 3 | 17 | 3 | 120 | 1 | -0.178 | 60 | 119.9893 | 1 | -17.7905 | 0.021347 | 0.11808 | 0.11999 | 119.781 | 1 | 0.018094 | -0.20833 | -1.1e-05 | -0.003253 | -0.1736 | -0.0011 | -15.2385
pf_sweep | 2 | 3 | 18 | 3 | 120 | 1 | -0.178 | 60 | 119.9893 | 1 | -17.7905 | 0.021347 | 0.11808 | 0.11999 | 119.781 | 1 | 0.018143 | -0.20833 | -1.1e-05 | -0.003204 | -0.1736 | -0.0011 | -15.0089
pf_sweep | 3 | 1 | 1 | 1 | 120 | 1 | -0.156 | 60 | 120.006 | 1 | -15.5881 | 0.018707 | 0.11854 | 0.12001 | 119.744 | 2.97 | 0.042968 | -0.26204 | 1.97 | 0.024261 | -0.2184 | -0.9999 | -23.4356
pf_sweep | 3 | 1 | 2 | 1 | 120 | 1 | -0.156 | 60 | 120.006 | 1 | -15.5881 | 0.018707 | 0.11854 | 0.12001 | 119.744 | 1.001 | 0.015134 | -0.26204 | 0.001001 | -0.003573 | -0.2184 | 0.1001 | -19.0985
pf_sweep | 3 | 1 | 3 | 1 | 120 | 1 | -0.156 | 60 | 120.006 | 1 | -15.5881 | 0.018707 | 0.11854 | 0.12001 | 119.744 | 1.001 | 0.015154 | -0.26204 | 0.001001 | -0.003553 | -0.2184 | 0.1001 | -18.9916
pf_sweep | 3 | 1 | 4 | 1 | 120 | 1 | -0.156 | 60 | 120.006 | 1 | -15.5881 | 0.018707 | 0.11854 | 0.12001 | 119.744 | 1.001 | 0.0151 | -0.26204 | 0.001001 | -0.003607 | -0.2184 | 0.1001 | -19.2802
pf_sweep | 3 | 1 | 5 | 1 | 120 | 1 | -0.156 | 60 | 120.006 | 1 | -15.5881 | 0.018707 | 0.11854 | 0.12001 | 119.744 | 1.001 | 0.015197 | -0.26204 | 0.001001 | -0.00351 | -0.2184 | 0.1001 | -18.7617
pf_sweep | 3 | 1 | 6 | 1 | 120 | 1 | -0.156 | 60 | 120.006 | 1 | -15.5881 | 0.018707 | 0.11854 | 0.12001 | 119.744 | 0.998 | 0.015643 | -0.26204 | -0.001999 | -0.003064 | -0.2184 | -0.1999 | -16.3775
pf_sweep | 3 | 1 | 7 | 2 | 120 | 1 | -0.156 | 60 | 119.9918 | 1.0002 | -15.5482 | 0.01866 | 0.11855 | 0.12001 | 119.783 | 2.979 | 0.042714 | -0.20881 | 1.9788 | 0.024054 | -0.174 | -0.71705 | -23.697
pf_sweep | 3 | 1 | 8 | 2 | 120 | 1 | -0.156 | 60 | 119.9918 | 1.0002 | -15.5482 | 0.01866 | 0.11855 | 0.12001 | 119.783 | 1 | 0.015036 | -0.20881 | -0.000172 | -0.003624 | -0.174 | -0.0172 | -19.4205
pf_sweep | 3 | 1 | 9 | 2 | 120 | 1 | -0.156 | 60 | 119.9918 | 1.0002 | -15.5482 | 0.01866 | 0.11855 | 0.12001 | 119.783 | 1.001 | 0.015191 | -0.20881 | 0.000828 | -0.003469 | -0.174 | 0.0828 | -18.5898
pf_sweep | 3 | 1 | 10 | 2 | 120 | 1 | -0.156 | 60 | 119.9918 | 1.0002 | -15.5482 | 0.01866 | 0.11855 | 0.12001 | 119.783 | 1 | 0.015143 | -0.20881 | -0.000172 | -0.003517 | -0.174 | -0.0172 | -18.847
pf_sweep | 3 | 1 | 11 | 2 | 120 | 1 | -0.156 | 60 | 119.9918 | 1.0002 | -15.5482 | 0.01866 | 0.11855 | 0.12001 | 119.783 | 1 | 0.015418 | -0.20881 | -0.000172 | -0.003242 | -0.174 | -0.0172 | -17.3733
pf_sweep | 3 | 1 | 12 | 2 | 120 | 1 | -0.156 | 60 | 119.9918 | 1.0002 | -15.5482 | 0.01866 | 0.11855 | 0.12001 | 119.783 | 1 | 0.015047 | -0.20881 | -0.000172 | -0.003613 | -0.174 | -0.0172 | -19.3615
pf_sweep | 3 | 1 | 13 | 3 | 120 | 1 | -0.156 | 60 | 119.9904 | 0.99999 | -15.587 | 0.018703 | 0.11852 | 0.11999 | 119.774 | 1 | 0.01533 | -0.21636 | 5e-06 | -0.003373 | -0.1803 | 0.0005 | -18.0338
pf_sweep | 3 | 1 | 14 | 3 | 120 | 1 | -0.156 | 60 | 119.9904 | 0.99999 | -15.587 | 0.018703 | 0.11852 | 0.11999 | 119.774 | 0.999 | 0.015547 | -0.21636 | -0.000995 | -0.003156 | -0.1803 | -0.0995 | -16.8735
pf_sweep | 3 | 1 | 15 | 3 | 120 | 1 | -0.156 | 60 | 119.9904 | 0.99999 | -15.587 | 0.018703 | 0.11852 | 0.11999 | 119.774 | 1.001 | 0.015372 | -0.21636 | 0.001005 | -0.003331 | -0.1803 | 0.1005 | -17.8092
pf_sweep | 3 | 1 | 16 | 3 | 120 | 1 | -0.156 | 60 | 119.9904 | 0.99999 | -15.587 | 0.018703 | 0.11852 | 0.11999 | 119.774 | 0.999 | 0.015594 | -0.21636 | -0.000995 | -0.003109 | -0.1803 | -0.0995 | -16.6222
pf_sweep | 3 | 1 | 17 | 3 | 120 | 1 | -0.156 | 60 | 119.9904 | 0.99999 | -15.587 | 0.018703 | 0.11852 | 0.11999 | 119.774 | 1 | 0.015433 | -0.21636 | 5e-06 | -0.00327 | -0.1803 | 0.0005 | -17.4831
pf_sweep | 3 | 1 | 18 | 3 | 120 | 1 | -0.156 | 60 | 119.9904 | 0.99999 | -15.587 | 0.018703 | 0.11852 | 0.11999 | 119.774 | 1 | 0.015486 | -0.21636 | 5e-06 | -0.003217 | -0.1803 | 0.0005 | -17.1997
pf_sweep | 3 | 2 | 1 | 1 | 120 | 1 | -0.156 | 60 | 120.0058 | 1 | -15.5848 | 0.018703 | 0.11854 | 0.12001 | 119.747 | 2.97 | 0.042942 | -0.25881 | 1.97 | 0.024239 | -0.2157 | -1.0027 | -23.4676
pf_sweep | 3 | 2 | 2 | 1 | 120 | 1 | -0.156 | 60 | 120.0058 | 1 | -15.5848 | 0.018703 | 0.11854 | 0.12001 | 119.747 | 1.001 | 0.015154 | -0.25881 | 0.000972 | -0.003549 | -0.2157 | 0.0972 | -18.9764
pf_sweep | 3 | 2 | 3 | 1 | 120 | 1 | -0.156 | 60 | 120.0058 | 1 | -15.5848 | 0.018703 | 0.11854 | 0.12001 | 119.747 | 1.001 | 0.01513 | -0.25881 | 0.000972 | -0.003573 | -0.2157 | 0.0972 | -19.1048
pf_sweep | 3 | 2 | 4 | 1 | 120 | 1 | -0.156 | 60 | 120.0058 | 1 | -15.5848 | 0.018703 | 0.11854 | 0.12001 | 119.747 | 1 | 0.01509 | -0.25881 | -2.8e-05 | -0.003613 | -0.2157 | -0.0028 | -19.3186
pf_sweep | 3 | 2 | 5 | 1 | 120 | 1 | -0.156 | 60 | 120.0058 | 1 | -15.5848 | 0.018703 | 0.11854 | 0.12001 | 119.747 | 1.001 | 0.015246 | -0.25881 | 0.000972 | -0.003457 | -0.2157 | 0.0972 | -18.4845
pf_sweep | 3 | 2 | 6 | 1 | 120 | 1 | -0.156 | 60 | 120.0058 | 1 | -15.5848 | 0.018703 | 0.11854 | 0.12001 | 119.747 | 0.998 | 0.015591 | -0.25881 | -0.002028 | -0.003112 | -0.2157 | -0.2028 | -16.6399
pf_sweep | 3 | 2 | 7 | 2 | 120 | 1 | -0.156 | 60 | 119.9879 | 1.0002 | -15.5498 | 0.018661 | 0.11855 | 0.12001 | 119.767 | 2.979 | 0.042836 | -0.22091 | 1.9788 | 0.024175 | -0.1841 | -0.71584 | -23.4832
pf_sweep | 3 | 2 | 8 | 2 | 120 | 1 | -0.156 | 60 | 119.9879 | 1.0002 | -15.5498 | 0.018661 | 0.11855 | 0.12001 | 119.767 | 1.001 | 0.015063 | -0.22091 | 0.00084 | -0.003598 | -0.1841 | 0.084 | -19.2801
pf_sweep | 3 | 2 | 9 | 2 | 120 | 1 | -0.156 | 60 | 119.9879 | 1.0002 | -15.5498 | 0.018661 | 0.11855 | 0.12001 | 119.767 | 1 | 0.015194 | -0.22091 | -0.00016 | -0.003467 | -0.1841 | -0.0159 | -18.5781
pf_sweep | 3 | 2 | 10 | 2 | 120 | 1 | -0.156 | 60 | 119.9879 | 1.0002 | -15.5498 | 0.018661 | 0.11855 | 0.12001 | 119.767 | 1 | 0.015172 | -0.22091 | -0.00016 | -0.003489 | -0.1841 | -0.0159 | -18.696
pf_sweep | 3 | 2 | 11 | 2 | 120 | 1 | -0.156 | 60 | 119.9879 | 1.0002 | -15.5498 | 0.018661 | 0.11855 | 0.12001 | 119.767 | 1 | 0.015407 | -0.22091 | -0.00016 | -0.003254 | -0.1841 | -0.0159 | -17.4367
pf_sweep | 3 | 2 | 12 | 2 | 120 | 1 | -0.156 | 60 | 119.9879 | 1.0002 | -15.5498 | 0.018661 | 0.11855 | 0.12001 | 119.767 | 1 | 0.01504 | -0.22091 | -0.00016 | -0.003621 | -0.1841 | -0.0159 | -19.4033
pf_sweep | 3 | 2 | 13 | 3 | 120 | 1 | -0.156 | 60 | 119.9953 | 1 | -15.5892 | 0.018707 | 0.11853 | 0.12 | 119.785 | 1 | 0.015318 | -0.21032 | -1.1e-05 | -0.003389 | -0.1753 | -0.0011 | -18.1141
pf_sweep | 3 | 2 | 14 | 3 | 120 | 1 | -0.156 | 60 | 119.9953 | 1 | -15.5892 | 0.018707 | 0.11853 | 0.12 | 119.785 | 0.999 | 0.015491 | -0.21032 | -0.001011 | -0.003216 | -0.1753 | -0.1011 | -17.1893
pf_sweep | 3 | 2 | 15 | 3 | 120 | 1 | -0.156 | 60 | 119.9953 | 1 | -15.5892 | 0.018707 | 0.11853 | 0.12 | 119.785 | 1.001 | 0.015391 | -0.21032 | 0.000989 | -0.003316 | -0.1753 | 0.0989 | -17.7239
pf_sweep | 3 | 2 | 16 | 3 | 120 | 1 | -0.156 | 60 | 119.9953 | 1 | -15.5892 | 0.018707 | 0.11853 | 0.12 | 119.785 | 0.999 | 0.015573 | -0.21032 | -0.001011 | -0.003134 | -0.1753 | -0.1011 | -16.7509
pf_sweep | 3 | 2 | 17 | 3 | 120 | 1 | -0.156 | 60 | 119.9953 | 1 | -15.5892 | 0.018707 | 0.11853 | 0.12 | 119.785 | 1 | 0.015405 | -0.21032 | -1.1e-05 | -0.003302 | -0.1753 | -0.0011 | -17.649
pf_sweep | 3 | 2 | 18 | 3 | 120 | 1 | -0.156 | 60 | 119.9953 | 1 | -15.5892 | 0.018707 | 0.11853 | 0.12 | 119.785 | 1 | 0.015504 | -0.21032 | -1.1e-05 | -0.003203 | -0.1753 | -0.0011 | -17.1198
pf_sweep | 3 | 3 | 1 | 1 | 120 | 1 | -0.156 | 60 | 120.0266 | 1 | -15.5846 | 0.018706 | 0.11856 | 0.12003 | 119.75 | 2.97 | 0.042942 | -0.27657 | 1.97 | 0.024236 | -0.2304 | -1.0007 | -23.4785
pf_sweep | 3 | 3 | 2 | 1 | 120 | 1 | -0.156 | 60 | 120.0266 | 1 | -15.5846 | 0.018706 | 0.11856 | 0.12003 | 119.75 | 1.001 | 0.015105 | -0.27657 | 0.000993 | -0.003601 | -0.2304 | 0.0993 | -19.2498
pf_sweep | 3 | 3 | 3 | 1 | 120 | 1 | -0.156 | 60 | 120.0266 | 1 | -15.5846 | 0.018706 | 0.11856 | 0.12003 | 119.75 | 1.001 | 0.015114 | -0.27657 | 0.000993 | -0.003592 | -0.2304 | 0.0993 | -19.2017
pf_sweep | 3 | 3 | 4 | 1 | 120 | 1 | -0.156 | 60 | 120.0266 | 1 | -15.5846 | 0.018706 | 0.11856 | 0.12003 | 119.75 | 1.001 | 0.01509 | -0.27657 | 0.000993 | -0.003616 | -0.2304 | 0.0993 | -19.33
pf_sweep | 3 | 3 | 5 | 1 | 120 | 1 | -0.156 | 60 | 120.0266 | 1 | -15.5846 | 0.018706 | 0.11856 | 0.12003 | 119.75 | 1.001 | 0.015207 | -0.27657 | 0.000993 | -0.003499 | -0.2304 | 0.0993 | -18.7046
pf_sweep | 3 | 3 | 6 | 1 | 120 | 1 | -0.156 | 60 | 120.0266 | 1 | -15.5846 | 0.018706 | 0.11856 | 0.12003 | 119.75 | 0.998 | 0.015613 | -0.27657 | -0.002007 | -0.003093 | -0.2304 | -0.2007 | -16.5341
pf_sweep | 3 | 3 | 7 | 2 | 120 | 1 | -0.156 | 60 | 119.9904 | 1.0002 | -15.5485 | 0.01866 | 0.11855 | 0.12001 | 119.774 | 2.979 | 0.04269 | -0.21639 | 1.9788 | 0.02403 | -0.1803 | -0.71666 | -23.7399
pf_sweep | 3 | 3 | 8 | 2 | 120 | 1 | -0.156 | 60 | 119.9904 | 1.0002 | -15.5485 | 0.01866 | 0.11855 | 0.12001 | 119.774 | 1 | 0.015075 | -0.21639 | -0.000168 | -0.003585 | -0.1803 | -0.0168 | -19.2115
pf_sweep | 3 | 3 | 9 | 2 | 120 | 1 | -0.156 | 60 | 119.9904 | 1.0002 | -15.5485 | 0.01866 | 0.11855 | 0.12001 | 119.774 | 1.001 | 0.015231 | -0.21639 | 0.000832 | -0.003429 | -0.1803 | 0.0832 | -18.3755
pf_sweep | 3 | 3 | 10 | 2 | 120 | 1 | -0.156 | 60 | 119.9904 | 1.0002 | -15.5485 | 0.01866 | 0.11855 | 0.12001 | 119.774 | 1 | 0.015153 | -0.21639 | -0.000168 | -0.003507 | -0.1803 | -0.0168 | -18.7935
pf_sweep | 3 | 3 | 11 | 2 | 120 | 1 | -0.156 | 60 | 119.9904 | 1.0002 | -15.5485 | 0.01866 | 0.11855 | 0.12001 | 119.774 | 1 | 0.015428 | -0.21639 | -0.000168 | -0.003232 | -0.1803 | -0.0168 | -17.3197
pf_sweep | 3 | 3 | 12 | 2 | 120 | 1 | -0.156 | 60 | 119.9904 | 1.0002 | -15.5485 | 0.01866 | 0.11855 | 0.12001 | 119.774 | 1 | 0.015058 | -0.21639 | -0.000168 | -0.003602 | -0.1803 | -0.0168 | -19.3026
pf_sweep | 3 | 3 | 13 | 3 | 120 | 1 | -0.156 | 60 | 119.9893 | 1 | -15.5897 | 0.018706 | 0.11852 | 0.11999 | 119.771 | 1 | 0.01532 | -0.2183 | -6e-06 | -0.003386 | -0.1819 | -0.0006 | -18.1014
pf_sweep | 3 | 3 | 14 | 3 | 120 | 1 | -0.156 | 60 | 119.9893 | 1 | -15.5897 | 0.018706 | 0.11852 | 0.11999 | 119.771 | 0.999 | 0.015537 | -0.2183 | -0.001006 | -0.003169 | -0.1819 | -0.1006 | -16.9414
pf_sweep | 3 | 3 | 15 | 3 | 120 | 1 | -0.156 | 60 | 119.9893 | 1 | -15.5897 | 0.018706 | 0.11852 | 0.11999 | 119.771 | 1.001 | 0.01539 | -0.2183 | 0.000994 | -0.003316 | -0.1819 | 0.0994 | -17.7272
pf_sweep | 3 | 3 | 16 | 3 | 120 | 1 | -0.156 | 60 | 119.9893 | 1 | -15.5897 | 0.018706 | 0.11852 | 0.11999 | 119.771 | 0.999 | 0.015594 | -0.2183 | -0.001006 | -0.003112 | -0.1819 | -0.1006 | -16.6366
pf_sweep | 3 | 3 | 17 | 3 | 120 | 1 | -0.156 | 60 | 119.9893 | 1 | -15.5897 | 0.018706 | 0.11852 | 0.11999 | 119.771 | 1 | 0.015432 | -0.2183 | -6e-06 | -0.003274 | -0.1819 | -0.0006 | -17.5027
pf_sweep | 3 | 3 | 18 | 3 | 120 | 1 | -0.156 | 60 | 119.9893 | 1 | -15.5897 | 0.018706 | 0.11852 | 0.11999 | 119.771 | 1 | 0.015484 | -0.2183 | -6e-06 | -0.003222 | -0.1819 | -0.0006 | -17.2247
pf_sweep | 4 | 1 | 1 | 1 | 120 | 1 | -0.133 | 60 | 119.9908 | 1 | -13.29 | 0.015947 | 0.11893 | 0.11999 | 119.748 | 0 | 0 | -0.24278 | -1 | -0.015947 | -0.2023 | -100 | -100
pf_sweep | 4 | 1 | 2 | 1 | 120 | 1 | -0.133 | 60 | 119.9908 | 1 | -13.29 | 0.015947 | 0.11893 | 0.11999 | 119.748 | 1.001 | 0.012388 | -0.24278 | 0.000996 | -0.003559 | -0.2023 | 0.0996 | -22.3167
pf_sweep | 4 | 1 | 3 | 1 | 120 | 1 | -0.133 | 60 | 119.9908 | 1 | -13.29 | 0.015947 | 0.11893 | 0.11999 | 119.748 | 1.001 | 0.012348 | -0.24278 | 0.000996 | -0.003599 | -0.2023 | 0.0996 | -22.5675
pf_sweep | 4 | 1 | 4 | 1 | 120 | 1 | -0.133 | 60 | 119.9908 | 1 | -13.29 | 0.015947 | 0.11893 | 0.11999 | 119.748 | 1.001 | 0.01232 | -0.24278 | 0.000996 | -0.003627 | -0.2023 | 0.0996 | -22.7431
pf_sweep | 4 | 1 | 5 | 1 | 120 | 1 | -0.133 | 60 | 119.9908 | 1 | -13.29 | 0.015947 | 0.11893 | 0.11999 | 119.748 | 1.001 | 0.012456 | -0.24278 | 0.000996 | -0.003491 | -0.2023 | 0.0996 | -21.8903
pf_sweep | 4 | 1 | 6 | 1 | 120 | 1 | -0.133 | 60 | 119.9908 | 1 | -13.29 | 0.015947 | 0.11893 | 0.11999 | 119.748 | 0.999 | 0.012904 | -0.24278 | -0.001004 | -0.003043 | -0.2023 | -0.1004 | -19.0809
pf_sweep | 4 | 1 | 7 | 2 | 120 | 1 | -0.133 | 60 | 119.992 | 1.0001 | -13.2513 | 0.015903 | 0.11895 | 0.12001 | 119.776 | 0 | 0 | -0.21601 | -1.0002 | -0.015903 | -0.18 | -100 | -100
pf_sweep | 4 | 1 | 8 | 2 | 120 | 1 | -0.133 | 60 | 119.992 | 1.0001 | -13.2513 | 0.015903 | 0.11895 | 0.12001 | 119.776 | 1 | 0.012248 | -0.21601 | -0.00015 | -0.003655 | -0.18 | -0.015 | -22.9827
pf_sweep | 4 | 1 | 9 | 2 | 120 | 1 | -0.133 | 60 | 119.992 | 1.0001 | -13.2513 | 0.015903 | 0.11895 | 0.12001 | 119.776 | 1.001 | 0.012423 | -0.21601 | 0.00085 | -0.00348 | -0.18 | 0.085 | -21.8823
pf_sweep | 4 | 1 | 10 | 2 | 120 | 1 | -0.133 | 60 | 119.992 | 1.0001 | -13.2513 | 0.015903 | 0.11895 | 0.12001 | 119.776 | 1 | 0.01236 | -0.21601 | -0.00015 | -0.003543 | -0.18 | -0.015 | -22.2785
pf_sweep | 4 | 1 | 11 | 2 | 120 | 1 | -0.133 | 60 | 119.992 | 1.0001 | -13.2513 | 0.015903 | 0.11895 | 0.12001 | 119.776 | 1 | 0.012635 | -0.21601 | -0.00015 | -0.003268 | -0.18 | -0.015 | -20.5492
pf_sweep | 4 | 1 | 12 | 2 | 120 | 1 | -0.133 | 60 | 119.992 | 1.0001 | -13.2513 | 0.015903 | 0.11895 | 0.12001 | 119.776 | 1 | 0.012268 | -0.21601 | -0.00015 | -0.003635 | -0.18 | -0.015 | -22.857
pf_sweep | 4 | 1 | 13 | 3 | 120 | 1 | -0.133 | 60 | 119.9917 | 1 | -13.2878 | 0.015944 | 0.11893 | 0.11999 | 119.783 | 1 | 0.012527 | -0.20866 | -5e-06 | -0.003417 | -0.1739 | -0.0005 | -21.4332
pf_sweep | 4 | 1 | 14 | 3 | 120 | 1 | -0.133 | 60 | 119.9917 | 1 | -13.2878 | 0.015944 | 0.11893 | 0.11999 | 119.783 | 0.999 | 0.012734 | -0.20866 | -0.001005 | -0.00321 | -0.1739 | -0.1005 | -20.1349
pf_sweep | 4 | 1 | 15 | 3 | 120 | 1 | -0.133 | 60 | 119.9917 | 1 | -13.2878 | 0.015944 | 0.11893 | 0.11999 | 119.783 | 1.001 | 0.01259 | -0.20866 | 0.000995 | -0.003354 | -0.1739 | 0.0995 | -21.0381
pf_sweep | 4 | 1 | 16 | 3 | 120 | 1 | -0.133 | 60 | 119.9917 | 1 | -13.2878 | 0.015944 | 0.11893 | 0.11999 | 119.783 | 0.999 | 0.012824 | -0.20866 | -0.001005 | -0.00312 | -0.1739 | -0.1005 | -19.5705
pf_sweep | 4 | 1 | 17 | 3 | 120 | 1 | -0.133 | 60 | 119.9917 | 1 | -13.2878 | 0.015944 | 0.11893 | 0.11999 | 119.783 | 1 | 0.012627 | -0.20866 | -5e-06 | -0.003317 | -0.1739 | -0.0005 | -20.806
pf_sweep | 4 | 1 | 18 | 3 | 120 | 1 | -0.133 | 60 | 119.9917 | 1 | -13.2878 | 0.015944 | 0.11893 | 0.11999 | 119.783 | 1 | 0.012701 | -0.20866 | -5e-06 | -0.003243 | -0.1739 | -0.0005 | -20.3419
pf_sweep | 4 | 2 | 1 | 1 | 120 | 1 | -0.133 | 60 | 120.0114 | 1 | -13.2837 | 0.015942 | 0.11895 | 0.12001 | 119.754 | 0 | 0 | -0.25744 | -1 | -0.015942 | -0.2145 | -100 | -100
pf_sweep | 4 | 2 | 2 | 1 | 120 | 1 | -0.133 | 60 | 120.0114 | 1 | -13.2837 | 0.015942 | 0.11895 | 0.12001 | 119.754 | 1.001 | 0.01239 | -0.25744 | 0.000997 | -0.003552 | -0.2145 | 0.0997 | -22.2808
pf_sweep | 4 | 2 | 3 | 1 | 120 | 1 | -0.133 | 60 | 120.0114 | 1 | -13.2837 | 0.015942 | 0.11895 | 0.12001 | 119.754 | 1.001 | 0.012337 | -0.25744 | 0.000997 | -0.003605 | -0.2145 | 0.0997 | -22.6133
pf_sweep | 4 | 2 | 4 | 1 | 120 | 1 | -0.133 | 60 | 120.0114 | 1 | -13.2837 | 0.015942 | 0.11895 | 0.12001 | 119.754 | 1.001 | 0.01234 | -0.25744 | 0.000997 | -0.003602 | -0.2145 | 0.0997 | -22.5944
pf_sweep | 4 | 2 | 5 | 1 | 120 | 1 | -0.133 | 60 | 120.0114 | 1 | -13.2837 | 0.015942 | 0.11895 | 0.12001 | 119.754 | 1.001 | 0.012429 | -0.25744 | 0.000997 | -0.003513 | -0.2145 | 0.0997 | -22.0362
pf_sweep | 4 | 2 | 6 | 1 | 120 | 1 | -0.133 | 60 | 120.0114 | 1 | -13.2837 | 0.015942 | 0.11895 | 0.12001 | 119.754 | 0.999 | 0.012816 | -0.25744 | -0.001003 | -0.003126 | -0.2145 | -0.1003 | -19.6086
pf_sweep | 4 | 2 | 7 | 2 | 120 | 1 | -0.133 | 60 | 119.9881 | 1.0002 | -13.2487 | 0.015899 | 0.11895 | 0.12001 | 119.772 | 0 | 0 | -0.21613 | -1.0002 | -0.015899 | -0.1801 | -100 | -100
pf_sweep | 4 | 2 | 8 | 2 | 120 | 1 | -0.133 | 60 | 119.9881 | 1.0002 | -13.2487 | 0.015899 | 0.11895 | 0.12001 | 119.772 | 1 | 0.01224 | -0.21613 | -0.000154 | -0.003659 | -0.1801 | -0.0154 | -23.0158
pf_sweep | 4 | 2 | 9 | 2 | 120 | 1 | -0.133 | 60 | 119.9881 | 1.0002 | -13.2487 | 0.015899 | 0.11895 | 0.12001 | 119.772 | 1.001 | 0.012434 | -0.21613 | 0.000846 | -0.003465 | -0.1801 | 0.0846 | -21.7956
pf_sweep | 4 | 2 | 10 | 2 | 120 | 1 | -0.133 | 60 | 119.9881 | 1.0002 | -13.2487 | 0.015899 | 0.11895 | 0.12001 | 119.772 | 1 | 0.012398 | -0.21613 | -0.000154 | -0.003501 | -0.1801 | -0.0154 | -22.022
pf_sweep | 4 | 2 | 11 | 2 | 120 | 1 | -0.133 | 60 | 119.9881 | 1.0002 | -13.2487 | 0.015899 | 0.11895 | 0.12001 | 119.772 | 1 | 0.012617 | -0.21613 | -0.000154 | -0.003282 | -0.1801 | -0.0154 | -20.6446
pf_sweep | 4 | 2 | 12 | 2 | 120 | 1 | -0.133 | 60 | 119.9881 | 1.0002 | -13.2487 | 0.015899 | 0.11895 | 0.12001 | 119.772 | 1 | 0.01227 | -0.21613 | -0.000154 | -0.003629 | -0.1801 | -0.0154 | -22.8271
pf_sweep | 4 | 2 | 13 | 3 | 120 | 1 | -0.133 | 60 | 119.993 | 0.99999 | -13.289 | 0.015946 | 0.11893 | 0.11999 | 119.774 | 1 | 0.012526 | -0.21902 | 1.5e-05 | -0.00342 | -0.1825 | 0.0015 | -21.4456
pf_sweep | 4 | 2 | 14 | 3 | 120 | 1 | -0.133 | 60 | 119.993 | 0.99999 | -13.289 | 0.015946 | 0.11893 | 0.11999 | 119.774 | 0.999 | 0.012762 | -0.21902 | -0.000985 | -0.003184 | -0.1825 | -0.0985 | -19.9656
pf_sweep | 4 | 2 | 15 | 3 | 120 | 1 | -0.133 | 60 | 119.993 | 0.99999 | -13.289 | 0.015946 | 0.11893 | 0.11999 | 119.774 | 1.001 | 0.012618 | -0.21902 | 0.001015 | -0.003328 | -0.1825 | 0.1015 | -20.8686
pf_sweep | 4 | 2 | 16 | 3 | 120 | 1 | -0.133 | 60 | 119.993 | 0.99999 | -13.289 | 0.015946 | 0.11893 | 0.11999 | 119.774 | 0.999 | 0.012784 | -0.21902 | -0.000985 | -0.003162 | -0.1825 | -0.0985 | -19.8276
pf_sweep | 4 | 2 | 17 | 3 | 120 | 1 | -0.133 | 60 | 119.993 | 0.99999 | -13.289 | 0.015946 | 0.11893 | 0.11999 | 119.774 | 1 | 0.012676 | -0.21902 | 1.5e-05 | -0.00327 | -0.1825 | 0.0015 | -20.5049
pf_sweep | 4 | 2 | 18 | 3 | 120 | 1 | -0.133 | 60 | 119.993 | 0.99999 | -13.289 | 0.015946 | 0.11893 | 0.11999 | 119.774 | 1 | 0.012719 | -0.21902 | 1.5e-05 | -0.003227 | -0.1825 | 0.0015 | -20.2352
pf_sweep | 4 | 3 | 1 | 1 | 120 | 1 | -0.133 | 60 | 120.0115 | 0.99999 | -13.2891 | 0.015948 | 0.11895 | 0.12001 | 119.741 | 0 | 0 | -0.27048 | -0.99999 | -0.015948 | -0.2254 | -100 | -100
pf_sweep | 4 | 3 | 2 | 1 | 120 | 1 | -0.133 | 60 | 120.0115 | 0.99999 | -13.2891 | 0.015948 | 0.11895 | 0.12001 | 119.741 | 1.001 | 0.012349 | -0.27048 | 0.001012 | -0.003599 | -0.2254 | 0.1012 | -22.5685
pf_sweep | 4 | 3 | 3 | 1 | 120 | 1 | -0.133 | 60 | 120.0115 | 0.99999 | -13.2891 | 0.015948 | 0.11895 | 0.12001 | 119.741 | 1.001 | 0.01236 | -0.27048 | 0.001012 | -0.003588 | -0.2254 | 0.1012 | -22.4995
pf_sweep | 4 | 3 | 4 | 1 | 120 | 1 | -0.133 | 60 | 120.0115 | 0.99999 | -13.2891 | 0.015948 | 0.11895 | 0.12001 | 119.741 | 1.001 | 0.012322 | -0.27048 | 0.001012 | -0.003626 | -0.2254 | 0.1012 | -22.7378
pf_sweep | 4 | 3 | 5 | 1 | 120 | 1 | -0.133 | 60 | 120.0115 | 0.99999 | -13.2891 | 0.015948 | 0.11895 | 0.12001 | 119.741 | 1.001 | 0.012418 | -0.27048 | 0.001012 | -0.00353 | -0.2254 | 0.1012 | -22.1358
pf_sweep | 4 | 3 | 6 | 1 | 120 | 1 | -0.133 | 60 | 120.0115 | 0.99999 | -13.2891 | 0.015948 | 0.11895 | 0.12001 | 119.741 | 0.998 | 0.012887 | -0.27048 | -0.001988 | -0.003061 | -0.2254 | -0.1988 | -19.1951
pf_sweep | 4 | 3 | 7 | 2 | 120 | 1 | -0.133 | 60 | 119.9911 | 1.0002 | -13.248 | 0.015899 | 0.11895 | 0.12001 | 119.78 | 0 | 0 | -0.21105 | -1.0002 | -0.015899 | -0.1759 | -100 | -100
pf_sweep | 4 | 3 | 8 | 2 | 120 | 1 | -0.133 | 60 | 119.9911 | 1.0002 | -13.248 | 0.015899 | 0.11895 | 0.12001 | 119.78 | 1 | 0.01225 | -0.21105 | -0.000161 | -0.003649 | -0.1759 | -0.0161 | -22.951
pf_sweep | 4 | 3 | 9 | 2 | 120 | 1 | -0.133 | 60 | 119.9911 | 1.0002 | -13.248 | 0.015899 | 0.11895 | 0.12001 | 119.78 | 1.001 | 0.012414 | -0.21105 | 0.000839 | -0.003485 | -0.1759 | 0.0839 | -21.9195
pf_sweep | 4 | 3 | 10 | 2 | 120 | 1 | -0.133 | 60 | 119.9911 | 1.0002 | -13.248 | 0.015899 | 0.11895 | 0.12001 | 119.78 | 1 | 0.01238 | -0.21105 | -0.000161 | -0.003519 | -0.1759 | -0.0161 | -22.1333
pf_sweep | 4 | 3 | 11 | 2 | 120 | 1 | -0.133 | 60 | 119.9911 | 1.0002 | -13.248 | 0.015899 | 0.11895 | 0.12001 | 119.78 | 1 | 0.012656 | -0.21105 | -0.000161 | -0.003243 | -0.1759 | -0.0161 | -20.3974
pf_sweep | 4 | 3 | 12 | 2 | 120 | 1 | -0.133 | 60 | 119.9911 | 1.0002 | -13.248 | 0.015899 | 0.11895 | 0.12001 | 119.78 | 1 | 0.012249 | -0.21105 | -0.000161 | -0.00365 | -0.1759 | -0.0161 | -22.9573
pf_sweep | 4 | 3 | 13 | 3 | 120 | 1 | -0.133 | 60 | 119.9897 | 1 | -13.2857 | 0.015941 | 0.11893 | 0.11999 | 119.77 | 1 | 0.012556 | -0.21966 | -4e-06 | -0.003385 | -0.1831 | -0.0004 | -21.2369
pf_sweep | 4 | 3 | 14 | 3 | 120 | 1 | -0.133 | 60 | 119.9897 | 1 | -13.2857 | 0.015941 | 0.11893 | 0.11999 | 119.77 | 0.999 | 0.012742 | -0.21966 | -0.001004 | -0.003199 | -0.1831 | -0.1004 | -20.0701
pf_sweep | 4 | 3 | 15 | 3 | 120 | 1 | -0.133 | 60 | 119.9897 | 1 | -13.2857 | 0.015941 | 0.11893 | 0.11999 | 119.77 | 1.001 | 0.01261 | -0.21966 | 0.000996 | -0.003331 | -0.1831 | 0.0996 | -20.8981
pf_sweep | 4 | 3 | 16 | 3 | 120 | 1 | -0.133 | 60 | 119.9897 | 1 | -13.2857 | 0.015941 | 0.11893 | 0.11999 | 119.77 | 0.999 | 0.012833 | -0.21966 | -0.001004 | -0.003108 | -0.1831 | -0.1004 | -19.4993
pf_sweep | 4 | 3 | 17 | 3 | 120 | 1 | -0.133 | 60 | 119.9897 | 1 | -13.2857 | 0.015941 | 0.11893 | 0.11999 | 119.77 | 1 | 0.012645 | -0.21966 | -4e-06 | -0.003296 | -0.1831 | -0.0004 | -20.6786
pf_sweep | 4 | 3 | 18 | 3 | 120 | 1 | -0.133 | 60 | 119.9897 | 1 | -13.2857 | 0.015941 | 0.11893 | 0.11999 | 119.77 | 1 | 0.012711 | -0.21966 | -4e-06 | -0.00323 | -0.1831 | -0.0004 | -20.2646
pf_sweep | 5 | 1 | 1 | 1 | 120 | 1 | -0.111 | 60 | 120.0048 | 1 | -11.0888 | 0.013308 | 0.11927 | 0.12001 | 119.729 | 0 | 0 | -0.27584 | -1 | -0.013308 | -0.2299 | -100 | -100
pf_sweep | 5 | 1 | 2 | 1 | 120 | 1 | -0.111 | 60 | 120.0048 | 1 | -11.0888 | 0.013308 | 0.11927 | 0.12001 | 119.729 | 0 | 0 | -0.27584 | -1 | -0.013308 | -0.2299 | -100 | -100
pf_sweep | 5 | 1 | 3 | 1 | 120 | 1 | -0.111 | 60 | 120.0048 | 1 | -11.0888 | 0.013308 | 0.11927 | 0.12001 | 119.729 | 0 | 0 | -0.27584 | -1 | -0.013308 | -0.2299 | -100 | -100
pf_sweep | 5 | 1 | 4 | 1 | 120 | 1 | -0.111 | 60 | 120.0048 | 1 | -11.0888 | 0.013308 | 0.11927 | 0.12001 | 119.729 | 0 | 0 | -0.27584 | -1 | -0.013308 | -0.2299 | -100 | -100
pf_sweep | 5 | 1 | 5 | 1 | 120 | 1 | -0.111 | 60 | 120.0048 | 1 | -11.0888 | 0.013308 | 0.11927 | 0.12001 | 119.729 | 0 | 0 | -0.27584 | -1 | -0.013308 | -0.2299 | -100 | -100
pf_sweep | 5 | 1 | 6 | 1 | 120 | 1 | -0.111 | 60 | 120.0048 | 1 | -11.0888 | 0.013308 | 0.11927 | 0.12001 | 119.729 | 0 | 0 | -0.27584 | -1 | -0.013308 | -0.2299 | -100 | -100
pf_sweep | 5 | 1 | 7 | 2 | 120 | 1 | -0.111 | 60 | 119.9903 | 1.0002 | -11.0519 | 0.013263 | 0.11927 | 0.12001 | 119.782 | 0 | 0 | -0.2083 | -1.0002 | -0.013263 | -0.1736 | -100 | -100
pf_sweep | 5 | 1 | 8 | 2 | 120 | 1 | -0.111 | 60 | 119.9903 | 1.0002 | -11.0519 | 0.013263 | 0.11927 | 0.12001 | 119.782 | 0 | 0 | -0.2083 | -1.0002 | -0.013263 | -0.1736 | -100 | -100
pf_sweep | 5 | 1 | 9 | 2 | 120 | 1 | -0.111 | 60 | 119.9903 | 1.0002 | -11.0519 | 0.013263 | 0.11927 | 0.12001 | 119.782 | 0 | 0 | -0.2083 | -1.0002 | -0.013263 | -0.1736 | -100 | -100
pf_sweep | 5 | 1 | 10 | 2 | 120 | 1 | -0.111 | 60 | 119.9903 | 1.0002 | -11.0519 | 0.013263 | 0.11927 | 0.12001 | 119.782 | 0 | 0 | -0.2083 | -1.0002 | -0.013263 | -0.1736 | -100 | -100
pf_sweep | 5 | 1 | 11 | 2 | 120 | 1 | -0.111 | 60 | 119.9903 | 1.0002 | -11.0519 | 0.013263 | 0.11927 | 0.12001 | 119.782 | 0 | 0 | -0.2083 | -1.0002 | -0.013263 | -0.1736 | -100 | -100
pf_sweep | 5 | 1 | 12 | 2 | 120 | 1 | -0.111 | 60 | 119.9903 | 1.0002 | -11.0519 | 0.013263 | 0.11927 | 0.12001 | 119.782 | 0 | 0 | -0.2083 | -1.0002 | -0.013263 | -0.1736 | -100 | -100
pf_sweep | 5 | 1 | 13 | 3 | 120 | 1 | -0.111 | 60 | 119.9912 | 1 | -11.0863 | 0.013303 | 0.11925 | 0.11999 | 119.778 | 0 | 0 | -0.21323 | -1 | -0.013303 | -0.1777 | -100 | -100
pf_sweep | 5 | 1 | 14 | 3 | 120 | 1 | -0.111 | 60 | 119.9912 | 1 | -11.0863 | 0.013303 | 0.11925 | 0.11999 | 119.778 | 0 | 0 | -0.21323 | -1 | -0.013303 | -0.1777 | -100 | -100
pf_sweep | 5 | 1 | 15 | 3 | 120 | 1 | -0.111 | 60 | 119.9912 | 1 | -11.0863 | 0.013303 | 0.11925 | 0.11999 | 119.778 | 0 | 0 | -0.21323 | -1 | -0.013303 | -0.1777 | -100 | -100
pf_sweep | 5 | 1 | 16 | 3 | 120 | 1 | -0.111 | 60 | 119.9912 | 1 | -11.0863 | 0.013303 | 0.11925 | 0.11999 | 119.778 | 0 | 0 | -0.21323 | -1 | -0.013303 | -0.1777 | -100 | -100
pf_sweep | 5 | 1 | 17 | 3 | 120 | 1 | -0.111 | 60 | 119.9912 | 1 | -11.0863 | 0.013303 | 0.11925 | 0.11999 | 119.778 | 0 | 0 | -0.21323 | -1 | -0.013303 | -0.1777 | -100 | -100
pf_sweep | 5 | 1 | 18 | 3 | 120 | 1 | -0.111 | 60 | 119.9912 | 1 | -11.0863 | 0.013303 | 0.11925 | 0.11999 | 119.778 | 0 | 0 | -0.21323 | -1 | -0.013303 | -0.1777 | -100 | -100
pf_sweep | 5 | 2 | 1 | 1 | 120 | 1 | -0.111 | 60 | 120.0154 | 1 | -11.0832 | 0.013302 | 0.11928 | 0.12002 | 119.742 | 0 | 0 | -0.2734 | -1 | -0.013302 | -0.2278 | -100 | -100
pf_sweep | 5 | 2 | 2 | 1 | 120 | 1 | -0.111 | 60 | 120.0154 | 1 | -11.0832 | 0.013302 | 0.11928 | 0.12002 | 119.742 | 0 | 0 | -0.2734 | -1 | -0.013302 | -0.2278 | -100 | -100
pf_sweep | 5 | 2 | 3 | 1 | 120 | 1 | -0.111 | 60 | 120.0154 | 1 | -11.0832 | 0.013302 | 0.11928 | 0.12002 | 119.742 | 0 | 0 | -0.2734 | -1 | -0.013302 | -0.2278 | -100 | -100
pf_sweep | 5 | 2 | 4 | 1 | 120 | 1 | -0.111 | 60 | 120.0154 | 1 | -11.0832 | 0.013302 | 0.11928 | 0.12002 | 119.742 | 0 | 0 | -0.2734 | -1 | -0.013302 | -0.2278 | -100 | -100
pf_sweep | 5 | 2 | 5 | 1 | 120 | 1 | -0.111 | 60 | 120.0154 | 1 | -11.0832 | 0.013302 | 0.11928 | 0.12002 | 119.742 | 0 | 0 | -0.2734 | -1 | -0.013302 | -0.2278 | -100 | -100
pf_sweep | 5 | 2 | 6 | 1 | 120 | 1 | -0.111 | 60 | 120.0154 | 1 | -11.0832 | 0.013302 | 0.11928 | 0.12002 | 119.742 | 0 | 0 | -0.2734 | -1 | -0.013302 | -0.2278 | -100 | -100
pf_sweep | 5 | 2 | 7 | 2 | 120 | 1 | -0.111 | 60 | 119.988 | 1.0002 | -11.0505 | 0.013261 | 0.11927 | 0.12001 | 119.775 | 0 | 0 | -0.21305 | -1.0002 | -0.013261 | -0.1776 | -100 | -100
pf_sweep | 5 | 2 | 8 | 2 | 120 | 1 | -0.111 | 60 | 119.988 | 1.0002 | -11.0505 | 0.013261 | 0.11927 | 0.12001 | 119.775 | 0 | 0 | -0.21305 | -1.0002 | -0.013261 | -0.1776 | -100 | -100
pf_sweep | 5 | 2 | 9 | 2 | 120 | 1 | -0.111 | 60 | 119.988 | 1.0002 | -11.0505 | 0.013261 | 0.11927 | 0.12001 | 119.775 | 0 | 0 | -0.21305 | -1.0002 | -0.013261 | -0.1776 | -100 | -100
pf_sweep | 5 | 2 | 10 | 2 | 120 | 1 | -0.111 | 60 | 119.988 | 1.0002 | -11.0505 | 0.013261 | 0.11927 | 0.12001 | 119.775 | 0 | 0 | -0.21305 | -1.0002 | -0.013261 | -0.1776 | -100 | -100
pf_sweep | 5 | 2 | 11 | 2 | 120 | 1 | -0.111 | 60 | 119.988 | 1.0002 | -11.0505 | 0.013261 | 0.11927 | 0.12001 | 119.775 | 0 | 0 | -0.21305 | -1.0002 | -0.013261 | -0.1776 | -100 | -100
pf_sweep | 5 | 2 | 12 | 2 | 120 | 1 | -0.111 | 60 | 119.988 | 1.0002 | -11.0505 | 0.013261 | 0.11927 | 0.12001 | 119.775 | 0 | 0 | -0.21305 | -1.0002 | -0.013261 | -0.1776 | -100 | -100
pf_sweep | 5 | 2 | 13 | 3 | 120 | 1 | -0.111 | 60 | 119.9899 | 1 | -11.0889 | 0.013306 | 0.11925 | 0.11999 | 119.785 | 0 | 0 | -0.20492 | -1 | -0.013306 | -0.1708 | -100 | -100
pf_sweep | 5 | 2 | 14 | 3 | 120 | 1 | -0.111 | 60 | 119.9899 | 1 | -11.0889 | 0.013306 | 0.11925 | 0.11999 | 119.785 | 0 | 0 | -0.20492 | -1 | -0.013306 | -0.1708 | -100 | -100
pf_sweep | 5 | 2 | 15 | 3 | 120 | 1 | -0.111 | 60 | 119.9899 | 1 | -11.0889 | 0.013306 | 0.11925 | 0.11999 | 119.785 | 0 | 0 | -0.20492 | -1 | -0.013306 | -0.1708 | -100 | -100
pf_sweep | 5 | 2 | 16 | 3 | 120 | 1 | -0.111 | 60 | 119.9899 | 1 | -11.0889 | 0.013306 | 0.11925 | 0.11999 | 119.785 | 0 | 0 | -0.20492 | -1 | -0.013306 | -0.1708 | -100 | -100
pf_sweep | 5 | 2 | 17 | 3 | 120 | 1 | -0.111 | 60 | 119.9899 | 1 | -11.0889 | 0.013306 | 0.11925 | 0.11999 | 119.785 | 0 | 0 | -0.20492 | -1 | -0.013306 | -0.1708 | -100 | -100
pf_sweep | 5 | 2 | 18 | 3 | 120 | 1 | -0.111 | 60 | 119.9899 | 1 | -11.0889 | 0.013306 | 0.11925 | 0.11999 | 119.785 | 0 | 0 | -0.20492 | -1 | -0.013306 | -0.1708 | -100 | -100
pf_sweep | 5 | 3 | 1 | 1 | 120 | 1 | -0.111 | 60 | 120.025 | 0.99999 | -11.0858 | 0.013306 | 0.11928 | 0.12002 | 119.759 | 0 | 0 | -0.26599 | -0.99999 | -0.013306 | -0.2216 | -100 | -100
pf_sweep | 5 | 3 | 2 | 1 | 120 | 1 | -0.111 | 60 | 120.025 | 0.99999 | -11.0858 | 0.013306 | 0.11928 | 0.12002 | 119.759 | 0 | 0 | -0.26599 | -0.99999 | -0.013306 | -0.2216 | -100 | -100
pf_sweep | 5 | 3 | 3 | 1 | 120 | 1 | -0.111 | 60 | 120.025 | 0.99999 | -11.0858 | 0.013306 | 0.11928 | 0.12002 | 119.759 | 0 | 0 | -0.26599 | -0.99999 | -0.013306 | -0.2216 | -100 | -100
pf_sweep | 5 | 3 | 4 | 1 | 120 | 1 | -0.111 | 60 | 120.025 | 0.99999 | -11.0858 | 0.013306 | 0.11928 | 0.12002 | 119.759 | 0 | 0 | -0.26599 | -0.99999 | -0.013306 | -0.2216 | -100 | -100
pf_sweep | 5 | 3 | 5 | 1 | 120 | 1 | -0.111 | 60 | 120.025 | 0.99999 | -11.0858 | 0.013306 | 0.11928 | 0.12002 | 119.759 | 0 | 0 | -0.26599 | -0.99999 | -0.013306 | -0.2216 | -100 | -100
pf_sweep | 5 | 3 | 6 | 1 | 120 | 1 | -0.111 | 60 | 120.025 | 0.99999 | -11.0858 | 0.013306 | 0.11928 | 0.12002 | 119.759 | 0 | 0 | -0.26599 | -0.99999 | -0.013306 | -0.2216 | -100 | -100
pf_sweep | 5 | 3 | 7 | 2 | 120 | 1 | -0.111 | 60 | 119.9923 | 1.0002 | -11.0478 | 0.013259 | 0.11928 | 0.12001 | 119.77 | 0 | 0 | -0.22234 | -1.0002 | -0.013259 | -0.1853 | -100 | -100
pf_sweep | 5 | 3 | 8 | 2 | 120 | 1 | -0.111 | 60 | 119.9923 | 1.0002 | -11.0478 | 0.013259 | 0.11928 | 0.12001 | 119.77 | 0 | 0 | -0.22234 | -1.0002 | -0.013259 | -0.1853 | -100 | -100
pf_sweep | 5 | 3 | 9 | 2 | 120 | 1 | -0.111 | 60 | 119.9923 | 1.0002 | -11.0478 | 0.013259 | 0.11928 | 0.12001 | 119.77 | 0 | 0 | -0.22234 | -1.0002 | -0.013259 | -0.1853 | -100 | -100
pf_sweep | 5 | 3 | 10 | 2 | 120 | 1 | -0.111 | 60 | 119.9923 | 1.0002 | -11.0478 | 0.013259 | 0.11928 | 0.12001 | 119.77 | 0 | 0 | -0.22234 | -1.0002 | -0.013259 | -0.1853 | -100 | -100
pf_sweep | 5 | 3 | 11 | 2 | 120 | 1 | -0.111 | 60 | 119.9923 | 1.0002 | -11.0478 | 0.013259 | 0.11928 | 0.12001 | 119.77 | 0 | 0 | -0.22234 | -1.0002 | -0.013259 | -0.1853 | -100 | -100
pf_sweep | 5 | 3 | 12 | 2 | 120 | 1 | -0.111 | 60 | 119.9923 | 1.0002 | -11.0478 | 0.013259 | 0.11928 | 0.12001 | 119.77 | 0 | 0 | -0.22234 | -1.0002 | -0.013259 | -0.1853 | -100 | -100
pf_sweep | 5 | 3 | 13 | 3 | 120 | 1 | -0.111 | 60 | 119.9905 | 1 | -11.0852 | 0.013301 | 0.11925 | 0.11999 | 119.77 | 0 | 0 | -0.22055 | -1 | -0.013301 | -0.1838 | -100 | -100
pf_sweep | 5 | 3 | 14 | 3 | 120 | 1 | -0.111 | 60 | 119.9905 | 1 | -11.0852 | 0.013301 | 0.11925 | 0.11999 | 119.77 | 0 | 0 | -0.22055 | -1 | -0.013301 | -0.1838 | -100 | -100
pf_sweep | 5 | 3 | 15 | 3 | 120 | 1 | -0.111 | 60 | 119.9905 | 1 | -11.0852 | 0.013301 | 0.11925 | 0.11999 | 119.77 | 0 | 0 | -0.22055 | -1 | -0.013301 | -0.1838 | -100 | -100
pf_sweep | 5 | 3 | 16 | 3 | 120 | 1 | -0.111 | 60 | 119.9905 | 1 | -11.0852 | 0.013301 | 0.11925 | 0.11999 | 119.77 | 0 | 0 | -0.22055 | -1 | -0.013301 | -0.1838 | -100 | -100
pf_sweep | 5 | 3 | 17 | 3 | 120 | 1 | -0.111 | 60 | 119.9905 | 1 | -11.0852 | 0.013301 | 0.11925 | 0.11999 | 119.77 | 0 | 0 | -0.22055 | -1 | -0.013301 | -0.1838 | -100 | -100
pf_sweep | 5 | 3 | 18 | 3 | 120 | 1 | -0.111 | 60 | 119.9905 | 1 | -11.0852 | 0.013301 | 0.11925 | 0.11999 | 119.77 | 0 | 0 | -0.22055 | -1 | -0.013301 | -0.1838 | -100 | -100
pf_sweep | 6 | 1 | 1 | 1 | 120 | 1 | -0.089 | 60 | 120.0026 | 0.99999 | -8.8866 | 0.010664 | 0.11953 | 0.12 | 119.733 | 0 | 0 | -0.26956 | -0.99999 | -0.010664 | -0.2246 | -100 | -100
pf_sweep | 6 | 1 | 2 | 1 | 120 | 1 | -0.089 | 60 | 120.0026 | 0.99999 | -8.8866 | 0.010664 | 0.11953 | 0.12 | 119.733 | 0 | 0 | -0.26956 | -0.99999 | -0.010664 | -0.2246 | -100 | -100
pf_sweep | 6 | 1 | 3 | 1 | 120 | 1 | -0.089 | 60 | 120.0026 | 0.99999 | -8.8866 | 0.010664 | 0.11953 | 0.12 | 119.733 | 0 | 0 | -0.26956 | -0.99999 | -0.010664 | -0.2246 | -100 | -100
pf_sweep | 6 | 1 | 4 | 1 | 120 | 1 | -0.089 | 60 | 120.0026 | 0.99999 | -8.8866 | 0.010664 | 0.11953 | 0.12 | 119.733 | 0 | 0 | -0.26956 | -0.99999 | -0.010664 | -0.2246 | -100 | -100
pf_sweep | 6 | 1 | 5 | 1 | 120 | 1 | -0.089 | 60 | 120.0026 | 0.99999 | -8.8866 | 0.010664 | 0.11953 | 0.12 | 119.733 | 0 | 0 | -0.26956 | -0.99999 | -0.010664 | -0.2246 | -100 | -100
pf_sweep | 6 | 1 | 6 | 1 | 120 | 1 | -0.089 | 60 | 120.0026 | 0.99999 | -8.8866 | 0.010664 | 0.11953 | 0.12 | 119.733 | 0 | 0 | -0.26956 | -0.99999 | -0.010664 | -0.2246 | -100 | -100
pf_sweep | 6 | 1 | 7 | 2 | 120 | 1 | -0.089 | 60 | 119.9901 | 1.0002 | -8.8508 | 0.010622 | 0.11954 | 0.12001 | 119.778 | 0 | 0 | -0.21213 | -1.0002 | -0.010622 | -0.1768 | -100 | -100
pf_sweep | 6 | 1 | 8 | 2 | 120 | 1 | -0.089 | 60 | 119.9901 | 1.0002 | -8.8508 | 0.010622 | 0.11954 | 0.12001 | 119.778 | 0 | 0 | -0.21213 | -1.0002 | -0.010622 | -0.1768 | -100 | -100
pf_sweep | 6 | 1 | 9 | 2 | 120 | 1 | -0.089 | 60 | 119.9901 | 1.0002 | -8.8508 | 0.010622 | 0.11954 | 0.12001 | 119.778 | 0 | 0 | -0.21213 | -1.0002 | -0.010622 | -0.1768 | -100 | -100
pf_sweep | 6 | 1 | 10 | 2 | 120 | 1 | -0.089 | 60 | 119.9901 | 1.0002 | -8.8508 | 0.010622 | 0.11954 | 0.12001 | 119.778 | 0 | 0 | -0.21213 | -1.0002 | -0.010622 | -0.1768 | -100 | -100
pf_sweep | 6 | 1 | 11 | 2 | 120 | 1 | -0.089 | 60 | 119.9901 | 1.0002 | -8.8508 | 0.010622 | 0.11954 | 0.12001 | 119.778 | 0 | 0 | -0.21213 | -1.0002 | -0.010622 | -0.1768 | -100 | -100
pf_sweep | 6 | 1 | 12 | 2 | 120 | 1 | -0.089 | 60 | 119.9901 | 1.0002 | -8.8508 | 0.010622 | 0.11954 | 0.12001 | 119.778 | 0 | 0 | -0.21213 | -1.0002 | -0.010622 | -0.1768 | -100 | -100
pf_sweep | 6 | 1 | 13 | 3 | 120 | 1 | -0.089 | 60 | 119.9954 | 1 | -8.8894 | 0.010667 | 0.11952 | 0.12 | 119.784 | 0 | 0 | -0.2114 | -1 | -0.010667 | -0.1762 | -100 | -100
pf_sweep | 6 | 1 | 14 | 3 | 120 | 1 | -0.089 | 60 | 119.9954 | 1 | -8.8894 | 0.010667 | 0.11952 | 0.12 | 119.784 | 0 | 0 | -0.2114 | -1 | -0.010667 | -0.1762 | -100 | -100
pf_sweep | 6 | 1 | 15 | 3 | 120 | 1 | -0.089 | 60 | 119.9954 | 1 | -8.8894 | 0.010667 | 0.11952 | 0.12 | 119.784 | 0 | 0 | -0.2114 | -1 | -0.010667 | -0.1762 | -100 | -100
pf_sweep | 6 | 1 | 16 | 3 | 120 | 1 | -0.089 | 60 | 119.9954 | 1 | -8.8894 | 0.010667 | 0.11952 | 0.12 | 119.784 | 0 | 0 | -0.2114 | -1 | -0.010667 | -0.1762 | -100 | -100
pf_sweep | 6 | 1 | 17 | 3 | 120 | 1 | -0.089 | 60 | 119.9954 | 1 | -8.8894 | 0.010667 | 0.11952 | 0.12 | 119.784 | 0 | 0 | -0.2114 | -1 | -0.010667 | -0.1762 | -100 | -100
pf_sweep | 6 | 1 | 18 | 3 | 120 | 1 | -0.089 | 60 | 119.9954 | 1 | -8.8894 | 0.010667 | 0.11952 | 0.12 | 119.784 | 0 | 0 | -0.2114 | -1 | -0.010667 | -0.1762 | -100 | -100
pf_sweep | 6 | 2 | 1 | 1 | 120 | 1 | -0.089 | 60 | 120.0141 | 1 | -8.8852 | 0.010664 | 0.11954 | 0.12001 | 119.773 | 0 | 0 | -0.24114 | -1 | -0.010664 | -0.2009 | -100 | -100
pf_sweep | 6 | 2 | 2 | 1 | 120 | 1 | -0.089 | 60 | 120.0141 | 1 | -8.8852 | 0.010664 | 0.11954 | 0.12001 | 119.773 | 0 | 0 | -0.24114 | -1 | -0.010664 | -0.2009 | -100 | -100
pf_sweep | 6 | 2 | 3 | 1 | 120 | 1 | -0.089 | 60 | 120.0141 | 1 | -8.8852 | 0.010664 | 0.11954 | 0.12001 | 119.773 | 0 | 0 | -0.24114 | -1 | -0.010664 | -0.2009 | -100 | -100
pf_sweep | 6 | 2 | 4 | 1 | 120 | 1 | -0.089 | 60 | 120.0141 | 1 | -8.8852 | 0.010664 | 0.11954 | 0.12001 | 119.773 | 0 | 0 | -0.24114 | -1 | -0.010664 | -0.2009 | -100 | -100
pf_sweep | 6 | 2 | 5 | 1 | 120 | 1 | -0.089 | 60 | 120.0141 | 1 | -8.8852 | 0.010664 | 0.11954 | 0.12001 | 119.773 | 0 | 0 | -0.24114 | -1 | -0.010664 | -0.2009 | -100 | -100
pf_sweep | 6 | 2 | 6 | 1 | 120 | 1 | -0.089 | 60 | 120.0141 | 1 | -8.8852 | 0.010664 | 0.11954 | 0.12001 | 119.773 | 0 | 0 | -0.24114 | -1 | -0.010664 | -0.2009 | -100 | -100
pf_sweep | 6 | 2 | 7 | 2 | 120 | 1 | -0.089 | 60 | 119.9873 | 1.0002 | -8.8494 | 0.01062 | 0.11954 | 0.12001 | 119.764 | 0 | 0 | -0.22333 | -1.0002 | -0.01062 | -0.1861 | -100 | -100
pf_sweep | 6 | 2 | 8 | 2 | 120 | 1 | -0.089 | 60 | 119.9873 | 1.0002 | -8.8494 | 0.01062 | 0.11954 | 0.12001 | 119.764 | 0 | 0 | -0.22333 | -1.0002 | -0.01062 | -0.1861 | -100 | -100
pf_sweep | 6 | 2 | 9 | 2 | 120 | 1 | -0.089 | 60 | 119.9873 | 1.0002 | -8.8494 | 0.01062 | 0.11954 | 0.12001 | 119.764 | 0 | 0 | -0.22333 | -1.0002 | -0.01062 | -0.1861 | -100 | -100
pf_sweep | 6 | 2 | 10 | 2 | 120 | 1 | -0.089 | 60 | 119.9873 | 1.0002 | -8.8494 | 0.01062 | 0.11954 | 0.12001 | 119.764 | 0 | 0 | -0.22333 | -1.0002 | -0.01062 | -0.1861 | -100 | -100
pf_sweep | 6 | 2 | 11 | 2 | 120 | 1 | -0.089 | 60 | 119.9873 | 1.0002 | -8.8494 | 0.01062 | 0.11954 | 0.12001 | 119.764 | 0 | 0 | -0.22333 | -1.0002 | -0.01062 | -0.1861 | -100 | -100
pf_sweep | 6 | 2 | 12 | 2 | 120 | 1 | -0.089 | 60 | 119.9873 | 1.0002 | -8.8494 | 0.01062 | 0.11954 | 0.12001 | 119.764 | 0 | 0 | -0.22333 | -1.0002 | -0.01062 | -0.1861 | -100 | -100
pf_sweep | 6 | 2 | 13 | 3 | 120 | 1 | -0.089 | 60 | 119.9905 | 1 | -8.8926 | 0.01067 | 0.11952 | 0.11999 | 119.783 | 0 | 0 | -0.20746 | -1 | -0.01067 | -0.1729 | -100 | -100
pf_sweep | 6 | 2 | 14 | 3 | 120 | 1 | -0.089 | 60 | 119.9905 | 1 | -8.8926 | 0.01067 | 0.11952 | 0.11999 | 119.783 | 0 | 0 | -0.20746 | -1 | -0.01067 | -0.1729 | -100 | -100
pf_sweep | 6 | 2 | 15 | 3 | 120 | 1 | -0.089 | 60 | 119.9905 | 1 | -8.8926 | 0.01067 | 0.11952 | 0.11999 | 119.783 | 0 | 0 | -0.20746 | -1 | -0.01067 | -0.1729 | -100 | -100
pf_sweep | 6 | 2 | 16 | 3 | 120 | 1 | -0.089 | 60 | 119.9905 | 1 | -8.8926 | 0.01067 | 0.11952 | 0.11999 | 119.783 | 0 | 0 | -0.20746 | -1 | -0.01067 | -0.1729 | -100 | -100
pf_sweep | 6 | 2 | 17 | 3 | 120 | 1 | -0.089 | 60 | 119.9905 | 1 | -8.8926 | 0.01067 | 0.11952 | 0.11999 | 119.783 | 0 | 0 | -0.20746 | -1 | -0.01067 | -0.1729 | -100 | -100
pf_sweep | 6 | 2 | 18 | 3 | 120 | 1 | -0.089 | 60 | 119.9905 | 1 | -8.8926 | 0.01067 | 0.11952 | 0.11999 | 119.783 | 0 | 0 | -0.20746 | -1 | -0.01067 | -0.1729 | -100 | -100
pf_sweep | 6 | 3 | 1 | 1 | 120 | 1 | -0.089 | 60 | 120.0093 | 1 | -8.8884 | 0.010667 | 0.11953 | 0.12001 | 119.743 | 0 | 0 | -0.26633 | -1 | -0.010667 | -0.2219 | -100 | -100
pf_sweep | 6 | 3 | 2 | 1 | 120 | 1 | -0.089 | 60 | 120.0093 | 1 | -8.8884 | 0.010667 | 0.11953 | 0.12001 | 119.743 | 0 | 0 | -0.26633 | -1 | -0.010667 | -0.2219 | -100 | -100
pf_sweep | 6 | 3 | 3 | 1 | 120 | 1 | -0.089 | 60 | 120.0093 | 1 | -8.8884 | 0.010667 | 0.11953 | 0.12001 | 119.743 | 0 | 0 | -0.26633 | -1 | -0.010667 | -0.2219 | -100 | -100
pf_sweep | 6 | 3 | 4 | 1 | 120 | 1 | -0.089 | 60 | 120.0093 | 1 | -8.8884 | 0.010667 | 0.11953 | 0.12001 | 119.743 | 0 | 0 | -0.26633 | -1 | -0.010667 | -0.2219 | -100 | -100
pf_sweep | 6 | 3 | 5 | 1 | 120 | 1 | -0.089 | 60 | 120.0093 | 1 | -8.8884 | 0.010667 | 0.11953 | 0.12001 | 119.743 | 0 | 0 | -0.26633 | -1 | -0.010667 | -0.2219 | -100 | -100
pf_sweep | 6 | 3 | 6 | 1 | 120 | 1 | -0.089 | 60 | 120.0093 | 1 | -8.8884 | 0.010667 | 0.11953 | 0.12001 | 119.743 | 0 | 0 | -0.26633 | -1 | -0.010667 | -0.2219 | -100 | -100
pf_sweep | 6 | 3 | 7 | 2 | 120 | 1 | -0.089 | 60 | 119.9929 | 1.0002 | -8.8486 | 0.010619 | 0.11954 | 0.12001 | 119.782 | 0 | 0 | -0.21092 | -1.0002 | -0.010619 | -0.1758 | -100 | -100
pf_sweep | 6 | 3 | 8 | 2 | 120 | 1 | -0.089 | 60 | 119.9929 | 1.0002 | -8.8486 | 0.010619 | 0.11954 | 0.12001 | 119.782 | 0 | 0 | -0.21092 | -1.0002 | -0.010619 | -0.1758 | -100 | -100
pf_sweep | 6 | 3 | 9 | 2 | 120 | 1 | -0.089 | 60 | 119.9929 | 1.0002 | -8.8486 | 0.010619 | 0.11954 | 0.12001 | 119.782 | 0 | 0 | -0.21092 | -1.0002 | -0.010619 | -0.1758 | -100 | -100
pf_sweep | 6 | 3 | 10 | 2 | 120 | 1 | -0.089 | 60 | 119.9929 | 1.0002 | -8.8486 | 0.010619 | 0.11954 | 0.12001 | 119.782 | 0 | 0 | -0.21092 | -1.0002 | -0.010619 | -0.1758 | -100 | -100
pf_sweep | 6 | 3 | 11 | 2 | 120 | 1 | -0.089 | 60 | 119.9929 | 1.0002 | -8.8486 | 0.010619 | 0.11954 | 0.12001 | 119.782 | 0 | 0 | -0.21092 | -1.0002 | -0.010619 | -0.1758 | -100 | -100
pf_sweep | 6 | 3 | 12 | 2 | 120 | 1 | -0.089 | 60 | 119.9929 | 1.0002 | -8.8486 | 0.010619 | 0.11954 | 0.12001 | 119.782 | 0 | 0 | -0.21092 | -1.0002 | -0.010619 | -0.1758 | -100 | -100
pf_sweep | 6 | 3 | 13 | 3 | 120 | 1 | -0.089 | 60 | 119.9901 | 0.99999 | -8.8883 | 0.010665 | 0.11951 | 0.11999 | 119.771 | 0 | 0 | -0.21913 | -0.99999 | -0.010665 | -0.1826 | -100 | -100
pf_sweep | 6 | 3 | 14 | 3 | 120 | 1 | -0.089 | 60 | 119.9901 | 0.99999 | -8.8883 | 0.010665 | 0.11951 | 0.11999 | 119.771 | 0 | 0 | -0.21913 | -0.99999 | -0.010665 | -0.1826 | -100 | -100
pf_sweep | 6 | 3 | 15 | 3 | 120 | 1 | -0.089 | 60 | 119.9901 | 0.99999 | -8.8883 | 0.010665 | 0.11951 | 0.11999 | 119.771 | 0 | 0 | -0.21913 | -0.99999 | -0.010665 | -0.1826 | -100 | -100
pf_sweep | 6 | 3 | 16 | 3 | 120 | 1 | -0.089 | 60 | 119.9901 | 0.99999 | -8.8883 | 0.010665 | 0.11951 | 0.11999 | 119.771 | 0 | 0 | -0.21913 | -0.99999 | -0.010665 | -0.1826 | -100 | -100
pf_sweep | 6 | 3 | 17 | 3 | 120 | 1 | -0.089 | 60 | 119.9901 | 0.99999 | -8.8883 | 0.010665 | 0.11951 | 0.11999 | 119.771 | 0 | 0 | -0.21913 | -0.99999 | -0.010665 | -0.1826 | -100 | -100
pf_sweep | 6 | 3 | 18 | 3 | 120 | 1 | -0.089 | 60 | 119.9901 | 0.99999 | -8.8883 | 0.010665 | 0.11951 | 0.11999 | 119.771 | 0 | 0 | -0.21913 | -0.99999 | -0.010665 | -0.1826 | -100 | -100
pf_sweep | 7 | 1 | 1 | 1 | 120 | 1 | -0.067 | 60 | 120.0168 | 1 | -6.6841 | 0.008022 | 0.11975 | 0.12002 | 119.743 | 0 | 0 | -0.27381 | -1 | -0.008022 | -0.2281 | -100 | -100
pf_sweep | 7 | 1 | 2 | 1 | 120 | 1 | -0.067 | 60 | 120.0168 | 1 | -6.6841 | 0.008022 | 0.11975 | 0.12002 | 119.743 | 0 | 0 | -0.27381 | -1 | -0.008022 | -0.2281 | -100 | -100
pf_sweep | 7 | 1 | 3 | 1 | 120 | 1 | -0.067 | 60 | 120.0168 | 1 | -6.6841 | 0.008022 | 0.11975 | 0.12002 | 119.743 | 0 | 0 | -0.27381 | -1 | -0.008022 | -0.2281 | -100 | -100
pf_sweep | 7 | 1 | 4 | 1 | 120 | 1 | -0.067 | 60 | 120.0168 | 1 | -6.6841 | 0.008022 | 0.11975 | 0.12002 | 119.743 | 0 | 0 | -0.27381 | -1 | -0.008022 | -0.2281 | -100 | -100
pf_sweep | 7 | 1 | 5 | 1 | 120 | 1 | -0.067 | 60 | 120.0168 | 1 | -6.6841 | 0.008022 | 0.11975 | 0.12002 | 119.743 | 0 | 0 | -0.27381 | -1 | -0.008022 | -0.2281 | -100 | -100
pf_sweep | 7 | 1 | 6 | 1 | 120 | 1 | -0.067 | 60 | 120.0168 | 1 | -6.6841 | 0.008022 | 0.11975 | 0.12002 | 119.743 | 0 | 0 | -0.27381 | -1 | -0.008022 | -0.2281 | -100 | -100
pf_sweep | 7 | 1 | 7 | 2 | 120 | 1 | -0.067 | 60 | 119.9868 | 1.0002 | -6.6485 | 0.0079786 | 0.11974 | 0.12001 | 119.77 | 0 | 0 | -0.21678 | -1.0002 | -0.007979 | -0.1807 | -100 | -100
pf_sweep | 7 | 1 | 8 | 2 | 120 | 1 | -0.067 | 60 | 119.9868 | 1.0002 | -6.6485 | 0.0079786 | 0.11974 | 0.12001 | 119.77 | 0 | 0 | -0.21678 | -1.0002 | -0.007979 | -0.1807 | -100 | -100
pf_sweep | 7 | 1 | 9 | 2 | 120 | 1 | -0.067 | 60 | 119.9868 | 1.0002 | -6.6485 | 0.0079786 | 0.11974 | 0.12001 | 119.77 | 0 | 0 | -0.21678 | -1.0002 | -0.007979 | -0.1807 | -100 | -100
pf_sweep | 7 | 1 | 10 | 2 | 120 | 1 | -0.067 | 60 | 119.9868 | 1.0002 | -6.6485 | 0.0079786 | 0.11974 | 0.12001 | 119.77 | 0 | 0 | -0.21678 | -1.0002 | -0.007979 | -0.1807 | -100 | -100
pf_sweep | 7 | 1 | 11 | 2 | 120 | 1 | -0.067 | 60 | 119.9868 | 1.0002 | -6.6485 | 0.0079786 | 0.11974 | 0.12001 | 119.77 | 0 | 0 | -0.21678 | -1.0002 | -0.007979 | -0.1807 | -100 | -100
pf_sweep | 7 | 1 | 12 | 2 | 120 | 1 | -0.067 | 60 | 119.9868 | 1.0002 | -6.6485 | 0.0079786 | 0.11974 | 0.12001 | 119.77 | 0 | 0 | -0.21678 | -1.0002 | -0.007979 | -0.1807 | -100 | -100
pf_sweep | 7 | 1 | 13 | 3 | 120 | 1 | -0.067 | 60 | 119.9907 | 0.99998 | -6.6894 | 0.0080265 | 0.11972 | 0.11999 | 119.787 | 0 | 0 | -0.2037 | -0.99998 | -0.008027 | -0.1698 | -100 | -100
pf_sweep | 7 | 1 | 14 | 3 | 120 | 1 | -0.067 | 60 | 119.9907 | 0.99998 | -6.6894 | 0.0080265 | 0.11972 | 0.11999 | 119.787 | 0 | 0 | -0.2037 | -0.99998 | -0.008027 | -0.1698 | -100 | -100
pf_sweep | 7 | 1 | 15 | 3 | 120 | 1 | -0.067 | 60 | 119.9907 | 0.99998 | -6.6894 | 0.0080265 | 0.11972 | 0.11999 | 119.787 | 0 | 0 | -0.2037 | -0.99998 | -0.008027 | -0.1698 | -100 | -100
pf_sweep | 7 | 1 | 16 | 3 | 120 | 1 | -0.067 | 60 | 119.9907 | 0.99998 | -6.6894 | 0.0080265 | 0.11972 | 0.11999 | 119.787 | 0 | 0 | -0.2037 | -0.99998 | -0.008027 | -0.1698 | -100 | -100
pf_sweep | 7 | 1 | 17 | 3 | 120 | 1 | -0.067 | 60 | 119.9907 | 0.99998 | -6.6894 | 0.0080265 | 0.11972 | 0.11999 | 119.787 | 0 | 0 | -0.2037 | -0.99998 | -0.008027 | -0.1698 | -100 | -100
pf_sweep | 7 | 1 | 18 | 3 | 120 | 1 | -0.067 | 60 | 119.9907 | 0.99998 | -6.6894 | 0.0080265 | 0.11972 | 0.11999 | 119.787 | 0 | 0 | -0.2037 | -0.99998 | -0.008027 | -0.1698 | -100 | -100
pf_sweep | 7 | 2 | 1 | 1 | 120 | 1 | -0.067 | 60 | 120.0029 | 0.99999 | -6.6885 | 0.0080263 | 0.11973 | 0.12 | 119.724 | 0 | 0 | -0.27894 | -0.99999 | -0.008026 | -0.2324 | -100 | -100
pf_sweep | 7 | 2 | 2 | 1 | 120 | 1 | -0.067 | 60 | 120.0029 | 0.99999 | -6.6885 | 0.0080263 | 0.11973 | 0.12 | 119.724 | 0 | 0 | -0.27894 | -0.99999 | -0.008026 | -0.2324 | -100 | -100
pf_sweep | 7 | 2 | 3 | 1 | 120 | 1 | -0.067 | 60 | 120.0029 | 0.99999 | -6.6885 | 0.0080263 | 0.11973 | 0.12 | 119.724 | 0 | 0 | -0.27894 | -0.99999 | -0.008026 | -0.2324 | -100 | -100
pf_sweep | 7 | 2 | 4 | 1 | 120 | 1 | -0.067 | 60 | 120.0029 | 0.99999 | -6.6885 | 0.0080263 | 0.11973 | 0.12 | 119.724 | 0 | 0 | -0.27894 | -0.99999 | -0.008026 | -0.2324 | -100 | -100
pf_sweep | 7 | 2 | 5 | 1 | 120 | 1 | -0.067 | 60 | 120.0029 | 0.99999 | -6.6885 | 0.0080263 | 0.11973 | 0.12 | 119.724 | 0 | 0 | -0.27894 | -0.99999 | -0.008026 | -0.2324 | -100 | -100
pf_sweep | 7 | 2 | 6 | 1 | 120 | 1 | -0.067 | 60 | 120.0029 | 0.99999 | -6.6885 | 0.0080263 | 0.11973 | 0.12 | 119.724 | 0 | 0 | -0.27894 | -0.99999 | -0.008026 | -0.2324 | -100 | -100
pf_sweep | 7 | 2 | 7 | 2 | 120 | 1 | -0.067 | 60 | 119.9903 | 1.0002 | -6.6475 | 0.0079776 | 0.11974 | 0.12001 | 119.764 | 0 | 0 | -0.22628 | -1.0002 | -0.007978 | -0.1886 | -100 | -100
pf_sweep | 7 | 2 | 8 | 2 | 120 | 1 | -0.067 | 60 | 119.9903 | 1.0002 | -6.6475 | 0.0079776 | 0.11974 | 0.12001 | 119.764 | 0 | 0 | -0.22628 | -1.0002 | -0.007978 | -0.1886 | -100 | -100
pf_sweep | 7 | 2 | 9 | 2 | 120 | 1 | -0.067 | 60 | 119.9903 | 1.0002 | -6.6475 | 0.0079776 | 0.11974 | 0.12001 | 119.764 | 0 | 0 | -0.22628 | -1.0002 | -0.007978 | -0.1886 | -100 | -100
pf_sweep | 7 | 2 | 10 | 2 | 120 | 1 | -0.067 | 60 | 119.9903 | 1.0002 | -6.6475 | 0.0079776 | 0.11974 | 0.12001 | 119.764 | 0 | 0 | -0.22628 | -1.0002 | -0.007978 | -0.1886 | -100 | -100
pf_sweep | 7 | 2 | 11 | 2 | 120 | 1 | -0.067 | 60 | 119.9903 | 1.0002 | -6.6475 | 0.0079776 | 0.11974 | 0.12001 | 119.764 | 0 | 0 | -0.22628 | -1.0002 | -0.007978 | -0.1886 | -100 | -100
pf_sweep | 7 | 2 | 12 | 2 | 120 | 1 | -0.067 | 60 | 119.9903 | 1.0002 | -6.6475 | 0.0079776 | 0.11974 | 0.12001 | 119.764 | 0 | 0 | -0.22628 | -1.0002 | -0.007978 | -0.1886 | -100 | -100
pf_sweep | 7 | 2 | 13 | 3 | 120 | 1 | -0.067 | 60 | 119.9887 | 0.99998 | -6.6876 | 0.0080242 | 0.11972 | 0.11999 | 119.781 | 0 | 0 | -0.20769 | -0.99998 | -0.008024 | -0.1731 | -100 | -100
pf_sweep | 7 | 2 | 14 | 3 | 120 | 1 | -0.067 | 60 | 119.9887 | 0.99998 | -6.6876 | 0.0080242 | 0.11972 | 0.11999 | 119.781 | 0 | 0 | -0.20769 | -0.99998 | -0.008024 | -0.1731 | -100 | -100
pf_sweep | 7 | 2 | 15 | 3 | 120 | 1 | -0.067 | 60 | 119.9887 | 0.99998 | -6.6876 | 0.0080242 | 0.11972 | 0.11999 | 119.781 | 0 | 0 | -0.20769 | -0.99998 | -0.008024 | -0.1731 | -100 | -100
pf_sweep | 7 | 2 | 16 | 3 | 120 | 1 | -0.067 | 60 | 119.9887 | 0.99998 | -6.6876 | 0.0080242 | 0.11972 | 0.11999 | 119.781 | 0 | 0 | -0.20769 | -0.99998 | -0.008024 | -0.1731 | -100 | -100
pf_sweep | 7 | 2 | 17 | 3 | 120 | 1 | -0.067 | 60 | 119.9887 | 0.99998 | -6.6876 | 0.0080242 | 0.11972 | 0.11999 | 119.781 | 0 | 0 | -0.20769 | -0.99998 | -0.008024 | -0.1731 | -100 | -100
pf_sweep | 7 | 2 | 18 | 3 | 120 | 1 | -0.067 | 60 | 119.9887 | 0.99998 | -6.6876 | 0.0080242 | 0.11972 | 0.11999 | 119.781 | 0 | 0 | -0.20769 | -0.99998 | -0.008024 | -0.1731 | -100 | -100
pf_sweep | 7 | 3 | 1 | 1 | 120 | 1 | -0.067 | 60 | 120.0098 | 0.99998 | -6.6835 | 0.0080207 | 0.11974 | 0.12001 | 119.746 | 0 | 0 | -0.26383 | -0.99998 | -0.008021 | -0.2198 | -100 | -100
pf_sweep | 7 | 3 | 2 | 1 | 120 | 1 | -0.067 | 60 | 120.0098 | 0.99998 | -6.6835 | 0.0080207 | 0.11974 | 0.12001 | 119.746 | 0 | 0 | -0.26383 | -0.99998 | -0.008021 | -0.2198 | -100 | -100
pf_sweep | 7 | 3 | 3 | 1 | 120 | 1 | -0.067 | 60 | 120.0098 | 0.99998 | -6.6835 | 0.0080207 | 0.11974 | 0.12001 | 119.746 | 0 | 0 | -0.26383 | -0.99998 | -0.008021 | -0.2198 | -100 | -100
pf_sweep | 7 | 3 | 4 | 1 | 120 | 1 | -0.067 | 60 | 120.0098 | 0.99998 | -6.6835 | 0.0080207 | 0.11974 | 0.12001 | 119.746 | 0 | 0 | -0.26383 | -0.99998 | -0.008021 | -0.2198 | -100 | -100
pf_sweep | 7 | 3 | 5 | 1 | 120 | 1 | -0.067 | 60 | 120.0098 | 0.99998 | -6.6835 | 0.0080207 | 0.11974 | 0.12001 | 119.746 | 0 | 0 | -0.26383 | -0.99998 | -0.008021 | -0.2198 | -100 | -100
pf_sweep | 7 | 3 | 6 | 1 | 120 | 1 | -0.067 | 60 | 120.0098 | 0.99998 | -6.6835 | 0.0080207 | 0.11974 | 0.12001 | 119.746 | 0 | 0 | -0.26383 | -0.99998 | -0.008021 | -0.2198 | -100 | -100
pf_sweep | 7 | 3 | 7 | 2 | 120 | 1 | -0.067 | 60 | 119.9904 | 1.0002 | -6.6502 | 0.0079809 | 0.11975 | 0.12001 | 119.779 | 0 | 0 | -0.21136 | -1.0002 | -0.007981 | -0.1762 | -100 | -100
pf_sweep | 7 | 3 | 8 | 2 | 120 | 1 | -0.067 | 60 | 119.9904 | 1.0002 | -6.6502 | 0.0079809 | 0.11975 | 0.12001 | 119.779 | 0 | 0 | -0.21136 | -1.0002 | -0.007981 | -0.1762 | -100 | -100
pf_sweep | 7 | 3 | 9 | 2 | 120 | 1 | -0.067 | 60 | 119.9904 | 1.0002 | -6.6502 | 0.0079809 | 0.11975 | 0.12001 | 119.779 | 0 | 0 | -0.21136 | -1.0002 | -0.007981 | -0.1762 | -100 | -100
pf_sweep | 7 | 3 | 10 | 2 | 120 | 1 | -0.067 | 60 | 119.9904 | 1.0002 | -6.6502 | 0.0079809 | 0.11975 | 0.12001 | 119.779 | 0 | 0 | -0.21136 | -1.0002 | -0.007981 | -0.1762 | -100 | -100
pf_sweep | 7 | 3 | 11 | 2 | 120 | 1 | -0.067 | 60 | 119.9904 | 1.0002 | -6.6502 | 0.0079809 | 0.11975 | 0.12001 | 119.779 | 0 | 0 | -0.21136 | -1.0002 | -0.007981 | -0.1762 | -100 | -100
pf_sweep | 7 | 3 | 12 | 2 | 120 | 1 | -0.067 | 60 | 119.9904 | 1.0002 | -6.6502 | 0.0079809 | 0.11975 | 0.12001 | 119.779 | 0 | 0 | -0.21136 | -1.0002 | -0.007981 | -0.1762 | -100 | -100
pf_sweep | 7 | 3 | 13 | 3 | 120 | 1 | -0.067 | 60 | 119.994 | 1 | -6.689 | 0.0080264 | 0.11973 | 0.11999 | 119.771 | 0 | 0 | -0.22301 | -1 | -0.008026 | -0.1859 | -100 | -100
pf_sweep | 7 | 3 | 14 | 3 | 120 | 1 | -0.067 | 60 | 119.994 | 1 | -6.689 | 0.0080264 | 0.11973 | 0.11999 | 119.771 | 0 | 0 | -0.22301 | -1 | -0.008026 | -0.1859 | -100 | -100
pf_sweep | 7 | 3 | 15 | 3 | 120 | 1 | -0.067 | 60 | 119.994 | 1 | -6.689 | 0.0080264 | 0.11973 | 0.11999 | 119.771 | 0 | 0 | -0.22301 | -1 | -0.008026 | -0.1859 | -100 | -100
pf_sweep | 7 | 3 | 16 | 3 | 120 | 1 | -0.067 | 60 | 119.994 | 1 | -6.689 | 0.0080264 | 0.11973 | 0.11999 | 119.771 | 0 | 0 | -0.22301 | -1 | -0.008026 | -0.1859 | -100 | -100
pf_sweep | 7 | 3 | 17 | 3 | 120 | 1 | -0.067 | 60 | 119.994 | 1 | -6.689 | 0.0080264 | 0.11973 | 0.11999 | 119.771 | 0 | 0 | -0.22301 | -1 | -0.008026 | -0.1859 | -100 | -100
pf_sweep | 7 | 3 | 18 | 3 | 120 | 1 | -0.067 | 60 | 119.994 | 1 | -6.689 | 0.0080264 | 0.11973 | 0.11999 | 119.771 | 0 | 0 | -0.22301 | -1 | -0.008026 | -0.1859 | -100 | -100
pf_sweep | 8 | 1 | 1 | 1 | 120 | 1 | -0.044 | 60 | 120.0122 | 1 | -4.3871 | 0.0052651 | 0.1199 | 0.12001 | 119.738 | 0 | 0 | -0.27415 | -1 | -0.005265 | -0.2284 | -100 | -100
pf_sweep | 8 | 1 | 2 | 1 | 120 | 1 | -0.044 | 60 | 120.0122 | 1 | -4.3871 | 0.0052651 | 0.1199 | 0.12001 | 119.738 | 0 | 0 | -0.27415 | -1 | -0.005265 | -0.2284 | -100 | -100
pf_sweep | 8 | 1 | 3 | 1 | 120 | 1 | -0.044 | 60 | 120.0122 | 1 | -4.3871 | 0.0052651 | 0.1199 | 0.12001 | 119.738 | 0 | 0 | -0.27415 | -1 | -0.005265 | -0.2284 | -100 | -100
pf_sweep | 8 | 1 | 4 | 1 | 120 | 1 | -0.044 | 60 | 120.0122 | 1 | -4.3871 | 0.0052651 | 0.1199 | 0.12001 | 119.738 | 0 | 0 | -0.27415 | -1 | -0.005265 | -0.2284 | -100 | -100
pf_sweep | 8 | 1 | 5 | 1 | 120 | 1 | -0.044 | 60 | 120.0122 | 1 | -4.3871 | 0.0052651 | 0.1199 | 0.12001 | 119.738 | 0 | 0 | -0.27415 | -1 | -0.005265 | -0.2284 | -100 | -100
pf_sweep | 8 | 1 | 6 | 1 | 120 | 1 | -0.044 | 60 | 120.0122 | 1 | -4.3871 | 0.0052651 | 0.1199 | 0.12001 | 119.738 | 0 | 0 | -0.27415 | -1 | -0.005265 | -0.2284 | -100 | -100
pf_sweep | 8 | 1 | 7 | 2 | 120 | 1 | -0.044 | 60 | 119.9897 | 1.0002 | -4.3535 | 0.0052246 | 0.11989 | 0.12001 | 119.781 | 0 | 0 | -0.20872 | -1.0002 | -0.005225 | -0.174 | -100 | -100
pf_sweep | 8 | 1 | 8 | 2 | 120 | 1 | -0.044 | 60 | 119.9897 | 1.0002 | -4.3535 | 0.0052246 | 0.11989 | 0.12001 | 119.781 | 0 | 0 | -0.20872 | -1.0002 | -0.005225 | -0.174 | -100 | -100
pf_sweep | 8 | 1 | 9 | 2 | 120 | 1 | -0.044 | 60 | 119.9897 | 1.0002 | -4.3535 | 0.0052246 | 0.11989 | 0.12001 | 119.781 | 0 | 0 | -0.20872 | -1.0002 | -0.005225 | -0.174 | -100 | -100
pf_sweep | 8 | 1 | 10 | 2 | 120 | 1 | -0.044 | 60 | 119.9897 | 1.0002 | -4.3535 | 0.0052246 | 0.11989 | 0.12001 | 119.781 | 0 | 0 | -0.20872 | -1.0002 | -0.005225 | -0.174 | -100 | -100
pf_sweep | 8 | 1 | 11 | 2 | 120 | 1 | -0.044 | 60 | 119.9897 | 1.0002 | -4.3535 | 0.0052246 | 0.11989 | 0.12001 | 119.781 | 0 | 0 | -0.20872 | -1.0002 | -0.005225 | -0.174 | -100 | -100
pf_sweep | 8 | 1 | 12 | 2 | 120 | 1 | -0.044 | 60 | 119.9897 | 1.0002 | -4.3535 | 0.0052246 | 0.11989 | 0.12001 | 119.781 | 0 | 0 | -0.20872 | -1.0002 | -0.005225 | -0.174 | -100 | -100
pf_sweep | 8 | 1 | 13 | 3 | 120 | 1 | -0.044 | 60 | 119.9935 | 1 | -4.3871 | 0.0052643 | 0.11988 | 0.12 | 119.776 | 0 | 0 | -0.21751 | -1 | -0.005264 | -0.1813 | -100 | -100
pf_sweep | 8 | 1 | 14 | 3 | 120 | 1 | -0.044 | 60 | 119.9935 | 1 | -4.3871 | 0.0052643 | 0.11988 | 0.12 | 119.776 | 0 | 0 | -0.21751 | -1 | -0.005264 | -0.1813 | -100 | -100
pf_sweep | 8 | 1 | 15 | 3 | 120 | 1 | -0.044 | 60 | 119.9935 | 1 | -4.3871 | 0.0052643 | 0.11988 | 0.12 | 119.776 | 0 | 0 | -0.21751 | -1 | -0.005264 | -0.1813 | -100 | -100
pf_sweep | 8 | 1 | 16 | 3 | 120 | 1 | -0.044 | 60 | 119.9935 | 1 | -4.3871 | 0.0052643 | 0.11988 | 0.12 | 119.776 | 0 | 0 | -0.21751 | -1 | -0.005264 | -0.1813 | -100 | -100
pf_sweep | 8 | 1 | 17 | 3 | 120 | 1 | -0.044 | 60 | 119.9935 | 1 | -4.3871 | 0.0052643 | 0.11988 | 0.12 | 119.776 | 0 | 0 | -0.21751 | -1 | -0.005264 | -0.1813 | -100 | -100
pf_sweep | 8 | 1 | 18 | 3 | 120 | 1 | -0.044 | 60 | 119.9935 | 1 | -4.3871 | 0.0052643 | 0.11988 | 0.12 | 119.776 | 0 | 0 | -0.21751 | -1 | -0.005264 | -0.1813 | -100 | -100
pf_sweep | 8 | 2 | 1 | 1 | 120 | 1 | -0.044 | 60 | 120.0092 | 1 | -4.3857 | 0.0052633 | 0.11989 | 0.12001 | 119.764 | 0 | 0 | -0.24516 | -1 | -0.005263 | -0.2043 | -100 | -100
pf_sweep | 8 | 2 | 2 | 1 | 120 | 1 | -0.044 | 60 | 120.0092 | 1 | -4.3857 | 0.0052633 | 0.11989 | 0.12001 | 119.764 | 0 | 0 | -0.24516 | -1 | -0.005263 | -0.2043 | -100 | -100
pf_sweep | 8 | 2 | 3 | 1 | 120 | 1 | -0.044 | 60 | 120.0092 | 1 | -4.3857 | 0.0052633 | 0.11989 | 0.12001 | 119.764 | 0 | 0 | -0.24516 | -1 | -0.005263 | -0.2043 | -100 | -100
pf_sweep | 8 | 2 | 4 | 1 | 120 | 1 | -0.044 | 60 | 120.0092 | 1 | -4.3857 | 0.0052633 | 0.11989 | 0.12001 | 119.764 | 0 | 0 | -0.24516 | -1 | -0.005263 | -0.2043 | -100 | -100
pf_sweep | 8 | 2 | 5 | 1 | 120 | 1 | -0.044 | 60 | 120.0092 | 1 | -4.3857 | 0.0052633 | 0.11989 | 0.12001 | 119.764 | 0 | 0 | -0.24516 | -1 | -0.005263 | -0.2043 | -100 | -100
pf_sweep | 8 | 2 | 6 | 1 | 120 | 1 | -0.044 | 60 | 120.0092 | 1 | -4.3857 | 0.0052633 | 0.11989 | 0.12001 | 119.764 | 0 | 0 | -0.24516 | -1 | -0.005263 | -0.2043 | -100 | -100
pf_sweep | 8 | 2 | 7 | 2 | 120 | 1 | -0.044 | 60 | 119.9892 | 1.0002 | -4.3507 | 0.0052212 | 0.1199 | 0.12001 | 119.768 | 0 | 0 | -0.22122 | -1.0002 | -0.005221 | -0.1844 | -100 | -100
pf_sweep | 8 | 2 | 8 | 2 | 120 | 1 | -0.044 | 60 | 119.9892 | 1.0002 | -4.3507 | 0.0052212 | 0.1199 | 0.12001 | 119.768 | 0 | 0 | -0.22122 | -1.0002 | -0.005221 | -0.1844 | -100 | -100
pf_sweep | 8 | 2 | 9 | 2 | 120 | 1 | -0.044 | 60 | 119.9892 | 1.0002 | -4.3507 | 0.0052212 | 0.1199 | 0.12001 | 119.768 | 0 | 0 | -0.22122 | -1.0002 | -0.005221 | -0.1844 | -100 | -100
pf_sweep | 8 | 2 | 10 | 2 | 120 | 1 | -0.044 | 60 | 119.9892 | 1.0002 | -4.3507 | 0.0052212 | 0.1199 | 0.12001 | 119.768 | 0 | 0 | -0.22122 | -1.0002 | -0.005221 | -0.1844 | -100 | -100
pf_sweep | 8 | 2 | 11 | 2 | 120 | 1 | -0.044 | 60 | 119.9892 | 1.0002 | -4.3507 | 0.0052212 | 0.1199 | 0.12001 | 119.768 | 0 | 0 | -0.22122 | -1.0002 | -0.005221 | -0.1844 | -100 | -100
pf_sweep | 8 | 2 | 12 | 2 | 120 | 1 | -0.044 | 60 | 119.9892 | 1.0002 | -4.3507 | 0.0052212 | 0.1199 | 0.12001 | 119.768 | 0 | 0 | -0.22122 | -1.0002 | -0.005221 | -0.1844 | -100 | -100
pf_sweep | 8 | 2 | 13 | 3 | 120 | 1 | -0.044 | 60 | 119.9915 | 0.99998 | -4.3925 | 0.0052705 | 0.11987 | 0.11999 | 119.785 | 0 | 0 | -0.20651 | -0.99998 | -0.005271 | -0.1721 | -100 | -100
pf_sweep | 8 | 2 | 14 | 3 | 120 | 1 | -0.044 | 60 | 119.9915 | 0.99998 | -4.3925 | 0.0052705 | 0.11987 | 0.11999 | 119.785 | 0 | 0 | -0.20651 | -0.99998 | -0.005271 | -0.1721 | -100 | -100
pf_sweep | 8 | 2 | 15 | 3 | 120 | 1 | -0.044 | 60 | 119.9915 | 0.99998 | -4.3925 | 0.0052705 | 0.11987 | 0.11999 | 119.785 | 0 | 0 | -0.20651 | -0.99998 | -0.005271 | -0.1721 | -100 | -100
pf_sweep | 8 | 2 | 16 | 3 | 120 | 1 | -0.044 | 60 | 119.9915 | 0.99998 | -4.3925 | 0.0052705 | 0.11987 | 0.11999 | 119.785 | 0 | 0 | -0.20651 | -0.99998 | -0.005271 | -0.1721 | -100 | -100
pf_sweep | 8 | 2 | 17 | 3 | 120 | 1 | -0.044 | 60 | 119.9915 | 0.99998 | -4.3925 | 0.0052705 | 0.11987 | 0.11999 | 119.785 | 0 | 0 | -0.20651 | -0.99998 | -0.005271 | -0.1721 | -100 | -100
pf_sweep | 8 | 2 | 18 | 3 | 120 | 1 | -0.044 | 60 | 119.9915 | 0.99998 | -4.3925 | 0.0052705 | 0.11987 | 0.11999 | 119.785 | 0 | 0 | -0.20651 | -0.99998 | -0.005271 | -0.1721 | -100 | -100
pf_sweep | 8 | 3 | 1 | 1 | 120 | 1 | -0.044 | 60 | 120.0125 | 1 | -4.3856 | 0.0052632 | 0.1199 | 0.12001 | 119.709 | 0 | 0 | -0.30353 | -1 | -0.005263 | -0.2529 | -100 | -100
pf_sweep | 8 | 3 | 2 | 1 | 120 | 1 | -0.044 | 60 | 120.0125 | 1 | -4.3856 | 0.0052632 | 0.1199 | 0.12001 | 119.709 | 0 | 0 | -0.30353 | -1 | -0.005263 | -0.2529 | -100 | -100
pf_sweep | 8 | 3 | 3 | 1 | 120 | 1 | -0.044 | 60 | 120.0125 | 1 | -4.3856 | 0.0052632 | 0.1199 | 0.12001 | 119.709 | 0 | 0 | -0.30353 | -1 | -0.005263 | -0.2529 | -100 | -100
pf_sweep | 8 | 3 | 4 | 1 | 120 | 1 | -0.044 | 60 | 120.0125 | 1 | -4.3856 | 0.0052632 | 0.1199 | 0.12001 | 119.709 | 0 | 0 | -0.30353 | -1 | -0.005263 | -0.2529 | -100 | -100
pf_sweep | 8 | 3 | 5 | 1 | 120 | 1 | -0.044 | 60 | 120.0125 | 1 | -4.3856 | 0.0052632 | 0.1199 | 0.12001 | 119.709 | 0 | 0 | -0.30353 | -1 | -0.005263 | -0.2529 | -100 | -100
pf_sweep | 8 | 3 | 6 | 1 | 120 | 1 | -0.044 | 60 | 120.0125 | 1 | -4.3856 | 0.0052632 | 0.1199 | 0.12001 | 119.709 | 0 | 0 | -0.30353 | -1 | -0.005263 | -0.2529 | -100 | -100
pf_sweep | 8 | 3 | 7 | 2 | 120 | 1 | -0.044 | 60 | 119.988 | 1.0002 | -4.3481 | 0.0052181 | 0.11989 | 0.12001 | 119.78 | 0 | 0 | -0.20796 | -1.0002 | -0.005218 | -0.1733 | -100 | -100
pf_sweep | 8 | 3 | 8 | 2 | 120 | 1 | -0.044 | 60 | 119.988 | 1.0002 | -4.3481 | 0.0052181 | 0.11989 | 0.12001 | 119.78 | 0 | 0 | -0.20796 | -1.0002 | -0.005218 | -0.1733 | -100 | -100
pf_sweep | 8 | 3 | 9 | 2 | 120 | 1 | -0.044 | 60 | 119.988 | 1.0002 | -4.3481 | 0.0052181 | 0.11989 | 0.12001 | 119.78 | 0 | 0 | -0.20796 | -1.0002 | -0.005218 | -0.1733 | -100 | -100
pf_sweep | 8 | 3 | 10 | 2 | 120 | 1 | -0.044 | 60 | 119.988 | 1.0002 | -4.3481 | 0.0052181 | 0.11989 | 0.12001 | 119.78 | 0 | 0 | -0.20796 | -1.0002 | -0.005218 | -0.1733 | -100 | -100
pf_sweep | 8 | 3 | 11 | 2 | 120 | 1 | -0.044 | 60 | 119.988 | 1.0002 | -4.3481 | 0.0052181 | 0.11989 | 0.12001 | 119.78 | 0 | 0 | -0.20796 | -1.0002 | -0.005218 | -0.1733 | -100 | -100
pf_sweep | 8 | 3 | 12 | 2 | 120 | 1 | -0.044 | 60 | 119.988 | 1.0002 | -4.3481 | 0.0052181 | 0.11989 | 0.12001 | 119.78 | 0 | 0 | -0.20796 | -1.0002 | -0.005218 | -0.1733 | -100 | -100
pf_sweep | 8 | 3 | 13 | 3 | 120 | 1 | -0.044 | 60 | 119.9918 | 1 | -4.3926 | 0.0052708 | 0.11988 | 0.11999 | 119.769 | 0 | 0 | -0.2228 | -1 | -0.005271 | -0.1857 | -100 | -100
pf_sweep | 8 | 3 | 14 | 3 | 120 | 1 | -0.044 | 60 | 119.9918 | 1 | -4.3926 | 0.0052708 | 0.11988 | 0.11999 | 119.769 | 0 | 0 | -0.2228 | -1 | -0.005271 | -0.1857 | -100 | -100
pf_sweep | 8 | 3 | 15 | 3 | 120 | 1 | -0.044 | 60 | 119.9918 | 1 | -4.3926 | 0.0052708 | 0.11988 | 0.11999 | 119.769 | 0 | 0 | -0.2228 | -1 | -0.005271 | -0.1857 | -100 | -100
pf_sweep | 8 | 3 | 16 | 3 | 120 | 1 | -0.044 | 60 | 119.9918 | 1 | -4.3926 | 0.0052708 | 0.11988 | 0.11999 | 119.769 | 0 | 0 | -0.2228 | -1 | -0.005271 | -0.1857 | -100 | -100
pf_sweep | 8 | 3 | 17 | 3 | 120 | 1 | -0.044 | 60 | 119.9918 | 1 | -4.3926 | 0.0052708 | 0.11988 | 0.11999 | 119.769 | 0 | 0 | -0.2228 | -1 | -0.005271 | -0.1857 | -100 | -100
pf_sweep | 8 | 3 | 18 | 3 | 120 | 1 | -0.044 | 60 | 119.9918 | 1 | -4.3926 | 0.0052708 | 0.11988 | 0.11999 | 119.769 | 0 | 0 | -0.2228 | -1 | -0.005271 | -0.1857 | -100 | -100
pf_sweep | 9 | 1 | 1 | 1 | 120 | 1 | -0.022 | 60 | 120.0173 | 0.99999 | -2.1837 | 0.0026208 | 0.11999 | 0.12002 | 119.707 | 0 | 0 | -0.31033 | -0.99999 | -0.002621 | -0.2586 | -100 | -100
pf_sweep | 9 | 1 | 2 | 1 | 120 | 1 | -0.022 | 60 | 120.0173 | 0.99999 | -2.1837 | 0.0026208 | 0.11999 | 0.12002 | 119.707 | 0 | 0 | -0.31033 | -0.99999 | -0.002621 | -0.2586 | -100 | -100
pf_sweep | 9 | 1 | 3 | 1 | 120 | 1 | -0.022 | 60 | 120.0173 | 0.99999 | -2.1837 | 0.0026208 | 0.11999 | 0.12002 | 119.707 | 0 | 0 | -0.31033 | -0.99999 | -0.002621 | -0.2586 | -100 | -100
pf_sweep | 9 | 1 | 4 | 1 | 120 | 1 | -0.022 | 60 | 120.0173 | 0.99999 | -2.1837 | 0.0026208 | 0.11999 | 0.12002 | 119.707 | 0 | 0 | -0.31033 | -0.99999 | -0.002621 | -0.2586 | -100 | -100
pf_sweep | 9 | 1 | 5 | 1 | 120 | 1 | -0.022 | 60 | 120.0173 | 0.99999 | -2.1837 | 0.0026208 | 0.11999 | 0.12002 | 119.707 | 0 | 0 | -0.31033 | -0.99999 | -0.002621 | -0.2586 | -100 | -100
pf_sweep | 9 | 1 | 6 | 1 | 120 | 1 | -0.022 | 60 | 120.0173 | 0.99999 | -2.1837 | 0.0026208 | 0.11999 | 0.12002 | 119.707 | 0 | 0 | -0.31033 | -0.99999 | -0.002621 | -0.2586 | -100 | -100
pf_sweep | 9 | 1 | 7 | 2 | 120 | 1 | -0.022 | 60 | 119.987 | 1.0002 | -2.1484 | 0.0025782 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20798 | -1.0002 | -0.002578 | -0.1733 | -100 | -100
pf_sweep | 9 | 1 | 8 | 2 | 120 | 1 | -0.022 | 60 | 119.987 | 1.0002 | -2.1484 | 0.0025782 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20798 | -1.0002 | -0.002578 | -0.1733 | -100 | -100
pf_sweep | 9 | 1 | 9 | 2 | 120 | 1 | -0.022 | 60 | 119.987 | 1.0002 | -2.1484 | 0.0025782 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20798 | -1.0002 | -0.002578 | -0.1733 | -100 | -100
pf_sweep | 9 | 1 | 10 | 2 | 120 | 1 | -0.022 | 60 | 119.987 | 1.0002 | -2.1484 | 0.0025782 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20798 | -1.0002 | -0.002578 | -0.1733 | -100 | -100
pf_sweep | 9 | 1 | 11 | 2 | 120 | 1 | -0.022 | 60 | 119.987 | 1.0002 | -2.1484 | 0.0025782 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20798 | -1.0002 | -0.002578 | -0.1733 | -100 | -100
pf_sweep | 9 | 1 | 12 | 2 | 120 | 1 | -0.022 | 60 | 119.987 | 1.0002 | -2.1484 | 0.0025782 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20798 | -1.0002 | -0.002578 | -0.1733 | -100 | -100
pf_sweep | 9 | 1 | 13 | 3 | 120 | 1 | -0.022 | 60 | 119.9931 | 1 | -2.1896 | 0.0026274 | 0.11997 | 0.11999 | 119.775 | 0 | 0 | -0.21806 | -1 | -0.002627 | -0.1817 | -100 | -100
pf_sweep | 9 | 1 | 14 | 3 | 120 | 1 | -0.022 | 60 | 119.9931 | 1 | -2.1896 | 0.0026274 | 0.11997 | 0.11999 | 119.775 | 0 | 0 | -0.21806 | -1 | -0.002627 | -0.1817 | -100 | -100
pf_sweep | 9 | 1 | 15 | 3 | 120 | 1 | -0.022 | 60 | 119.9931 | 1 | -2.1896 | 0.0026274 | 0.11997 | 0.11999 | 119.775 | 0 | 0 | -0.21806 | -1 | -0.002627 | -0.1817 | -100 | -100
pf_sweep | 9 | 1 | 16 | 3 | 120 | 1 | -0.022 | 60 | 119.9931 | 1 | -2.1896 | 0.0026274 | 0.11997 | 0.11999 | 119.775 | 0 | 0 | -0.21806 | -1 | -0.002627 | -0.1817 | -100 | -100
pf_sweep | 9 | 1 | 17 | 3 | 120 | 1 | -0.022 | 60 | 119.9931 | 1 | -2.1896 | 0.0026274 | 0.11997 | 0.11999 | 119.775 | 0 | 0 | -0.21806 | -1 | -0.002627 | -0.1817 | -100 | -100
pf_sweep | 9 | 1 | 18 | 3 | 120 | 1 | -0.022 | 60 | 119.9931 | 1 | -2.1896 | 0.0026274 | 0.11997 | 0.11999 | 119.775 | 0 | 0 | -0.21806 | -1 | -0.002627 | -0.1817 | -100 | -100
pf_sweep | 9 | 2 | 1 | 1 | 120 | 1 | -0.022 | 60 | 120.0117 | 1 | -2.1872 | 0.0026249 | 0.11998 | 0.12001 | 119.672 | 0 | 0 | -0.3397 | -1 | -0.002625 | -0.2831 | -100 | -100
pf_sweep | 9 | 2 | 2 | 1 | 120 | 1 | -0.022 | 60 | 120.0117 | 1 | -2.1872 | 0.0026249 | 0.11998 | 0.12001 | 119.672 | 0 | 0 | -0.3397 | -1 | -0.002625 | -0.2831 | -100 | -100
pf_sweep | 9 | 2 | 3 | 1 | 120 | 1 | -0.022 | 60 | 120.0117 | 1 | -2.1872 | 0.0026249 | 0.11998 | 0.12001 | 119.672 | 0 | 0 | -0.3397 | -1 | -0.002625 | -0.2831 | -100 | -100
pf_sweep | 9 | 2 | 4 | 1 | 120 | 1 | -0.022 | 60 | 120.0117 | 1 | -2.1872 | 0.0026249 | 0.11998 | 0.12001 | 119.672 | 0 | 0 | -0.3397 | -1 | -0.002625 | -0.2831 | -100 | -100
pf_sweep | 9 | 2 | 5 | 1 | 120 | 1 | -0.022 | 60 | 120.0117 | 1 | -2.1872 | 0.0026249 | 0.11998 | 0.12001 | 119.672 | 0 | 0 | -0.3397 | -1 | -0.002625 | -0.2831 | -100 | -100
pf_sweep | 9 | 2 | 6 | 1 | 120 | 1 | -0.022 | 60 | 120.0117 | 1 | -2.1872 | 0.0026249 | 0.11998 | 0.12001 | 119.672 | 0 | 0 | -0.3397 | -1 | -0.002625 | -0.2831 | -100 | -100
pf_sweep | 9 | 2 | 7 | 2 | 120 | 1 | -0.022 | 60 | 119.9886 | 1.0002 | -2.1456 | 0.0025749 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20963 | -1.0002 | -0.002575 | -0.1747 | -100 | -100
pf_sweep | 9 | 2 | 8 | 2 | 120 | 1 | -0.022 | 60 | 119.9886 | 1.0002 | -2.1456 | 0.0025749 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20963 | -1.0002 | -0.002575 | -0.1747 | -100 | -100
pf_sweep | 9 | 2 | 9 | 2 | 120 | 1 | -0.022 | 60 | 119.9886 | 1.0002 | -2.1456 | 0.0025749 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20963 | -1.0002 | -0.002575 | -0.1747 | -100 | -100
pf_sweep | 9 | 2 | 10 | 2 | 120 | 1 | -0.022 | 60 | 119.9886 | 1.0002 | -2.1456 | 0.0025749 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20963 | -1.0002 | -0.002575 | -0.1747 | -100 | -100
pf_sweep | 9 | 2 | 11 | 2 | 120 | 1 | -0.022 | 60 | 119.9886 | 1.0002 | -2.1456 | 0.0025749 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20963 | -1.0002 | -0.002575 | -0.1747 | -100 | -100
pf_sweep | 9 | 2 | 12 | 2 | 120 | 1 | -0.022 | 60 | 119.9886 | 1.0002 | -2.1456 | 0.0025749 | 0.11998 | 0.12001 | 119.779 | 0 | 0 | -0.20963 | -1.0002 | -0.002575 | -0.1747 | -100 | -100
pf_sweep | 9 | 2 | 13 | 3 | 120 | 1 | -0.022 | 60 | 119.9908 | 1 | -2.1878 | 0.0026252 | 0.11996 | 0.11999 | 119.783 | 0 | 0 | -0.20776 | -1 | -0.002625 | -0.1731 | -100 | -100
pf_sweep | 9 | 2 | 14 | 3 | 120 | 1 | -0.022 | 60 | 119.9908 | 1 | -2.1878 | 0.0026252 | 0.11996 | 0.11999 | 119.783 | 0 | 0 | -0.20776 | -1 | -0.002625 | -0.1731 | -100 | -100
pf_sweep | 9 | 2 | 15 | 3 | 120 | 1 | -0.022 | 60 | 119.9908 | 1 | -2.1878 | 0.0026252 | 0.11996 | 0.11999 | 119.783 | 0 | 0 | -0.20776 | -1 | -0.002625 | -0.1731 | -100 | -100
pf_sweep | 9 | 2 | 16 | 3 | 120 | 1 | -0.022 | 60 | 119.9908 | 1 | -2.1878 | 0.0026252 | 0.11996 | 0.11999 | 119.783 | 0 | 0 | -0.20776 | -1 | -0.002625 | -0.1731 | -100 | -100
pf_sweep | 9 | 2 | 17 | 3 | 120 | 1 | -0.022 | 60 | 119.9908 | 1 | -2.1878 | 0.0026252 | 0.11996 | 0.11999 | 119.783 | 0 | 0 | -0.20776 | -1 | -0.002625 | -0.1731 | -100 | -100
pf_sweep | 9 | 2 | 18 | 3 | 120 | 1 | -0.022 | 60 | 119.9908 | 1 | -2.1878 | 0.0026252 | 0.11996 | 0.11999 | 119.783 | 0 | 0 | -0.20776 | -1 | -0.002625 | -0.1731 | -100 | -100
pf_sweep | 9 | 3 | 1 | 1 | 120 | 1 | -0.022 | 60 | 120.0035 | 1 | -2.187 | 0.0026244 | 0.11997 | 0.12 | 119.717 | 0 | 0 | -0.28647 | -1 | -0.002624 | -0.2387 | -100 | -100
pf_sweep | 9 | 3 | 2 | 1 | 120 | 1 | -0.022 | 60 | 120.0035 | 1 | -2.187 | 0.0026244 | 0.11997 | 0.12 | 119.717 | 0 | 0 | -0.28647 | -1 | -0.002624 | -0.2387 | -100 | -100
pf_sweep | 9 | 3 | 3 | 1 | 120 | 1 | -0.022 | 60 | 120.0035 | 1 | -2.187 | 0.0026244 | 0.11997 | 0.12 | 119.717 | 0 | 0 | -0.28647 | -1 | -0.002624 | -0.2387 | -100 | -100
pf_sweep | 9 | 3 | 4 | 1 | 120 | 1 | -0.022 | 60 | 120.0035 | 1 | -2.187 | 0.0026244 | 0.11997 | 0.12 | 119.717 | 0 | 0 | -0.28647 | -1 | -0.002624 | -0.2387 | -100 | -100
pf_sweep | 9 | 3 | 5 | 1 | 120 | 1 | -0.022 | 60 | 120.0035 | 1 | -2.187 | 0.0026244 | 0.11997 | 0.12 | 119.717 | 0 | 0 | -0.28647 | -1 | -0.002624 | -0.2387 | -100 | -100
pf_sweep | 9 | 3 | 6 | 1 | 120 | 1 | -0.022 | 60 | 120.0035 | 1 | -2.187 | 0.0026244 | 0.11997 | 0.12 | 119.717 | 0 | 0 | -0.28647 | -1 | -0.002624 | -0.2387 | -100 | -100
pf_sweep | 9 | 3 | 7 | 2 | 120 | 1 | -0.022 | 60 | 119.9891 | 1.0002 | -2.1464 | 0.0025759 | 0.11998 | 0.12001 | 119.772 | 0 | 0 | -0.21714 | -1.0002 | -0.002576 | -0.181 | -100 | -100
pf_sweep | 9 | 3 | 8 | 2 | 120 | 1 | -0.022 | 60 | 119.9891 | 1.0002 | -2.1464 | 0.0025759 | 0.11998 | 0.12001 | 119.772 | 0 | 0 | -0.21714 | -1.0002 | -0.002576 | -0.181 | -100 | -100
pf_sweep | 9 | 3 | 9 | 2 | 120 | 1 | -0.022 | 60 | 119.9891 | 1.0002 | -2.1464 | 0.0025759 | 0.11998 | 0.12001 | 119.772 | 0 | 0 | -0.21714 | -1.0002 | -0.002576 | -0.181 | -100 | -100
pf_sweep | 9 | 3 | 10 | 2 | 120 | 1 | -0.022 | 60 | 119.9891 | 1.0002 | -2.1464 | 0.0025759 | 0.11998 | 0.12001 | 119.772 | 0 | 0 | -0.21714 | -1.0002 | -0.002576 | -0.181 | -100 | -100
pf_sweep | 9 | 3 | 11 | 2 | 120 | 1 | -0.022 | 60 | 119.9891 | 1.0002 | -2.1464 | 0.0025759 | 0.11998 | 0.12001 | 119.772 | 0 | 0 | -0.21714 | -1.0002 | -0.002576 | -0.181 | -100 | -100
pf_sweep | 9 | 3 | 12 | 2 | 120 | 1 | -0.022 | 60 | 119.9891 | 1.0002 | -2.1464 | 0.0025759 | 0.11998 | 0.12001 | 119.772 | 0 | 0 | -0.21714 | -1.0002 | -0.002576 | -0.181 | -100 | -100
pf_sweep | 9 | 3 | 13 | 3 | 120 | 1 | -0.022 | 60 | 119.9902 | 1 | -2.1881 | 0.0026255 | 0.11996 | 0.11999 | 119.771 | 0 | 0 | -0.21922 | -1 | -0.002626 | -0.1827 | -100 | -100
pf_sweep | 9 | 3 | 14 | 3 | 120 | 1 | -0.022 | 60 | 119.9902 | 1 | -2.1881 | 0.0026255 | 0.11996 | 0.11999 | 119.771 | 0 | 0 | -0.21922 | -1 | -0.002626 | -0.1827 | -100 | -100
pf_sweep | 9 | 3 | 15 | 3 | 120 | 1 | -0.022 | 60 | 119.9902 | 1 | -2.1881 | 0.0026255 | 0.11996 | 0.11999 | 119.771 | 0 | 0 | -0.21922 | -1 | -0.002626 | -0.1827 | -100 | -100
pf_sweep | 9 | 3 | 16 | 3 | 120 | 1 | -0.022 | 60 | 119.9902 | 1 | -2.1881 | 0.0026255 | 0.11996 | 0.11999 | 119.771 | 0 | 0 | -0.21922 | -1 | -0.002626 | -0.1827 | -100 | -100
pf_sweep | 9 | 3 | 17 | 3 | 120 | 1 | -0.022 | 60 | 119.9902 | 1 | -2.1881 | 0.0026255 | 0.11996 | 0.11999 | 119.771 | 0 | 0 | -0.21922 | -1 | -0.002626 | -0.1827 | -100 | -100
pf_sweep | 9 | 3 | 18 | 3 | 120 | 1 | -0.022 | 60 | 119.9902 | 1 | -2.1881 | 0.0026255 | 0.11996 | 0.11999 | 119.771 | 0 | 0 | -0.21922 | -1 | -0.002626 | -0.1827 | -100 | -100
NOTE
> I did some manual poking... it seems like if the reported PF of the EM16P drops below 0.1PF it will start reporting current as zero... along with other metric that rely on current